## QC Test Summary-Siemens
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Siemens
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer |  |  | Siemens |  |  |  |  | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [6 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [7 empty] | Annual Survey
Features |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Image Quality
Largest 5 fibers, 4 speck groups and 4 masses visible*
(*largest 4 fibers, 3 speck groups and 3 masses acceptable if spatial resolution and CNR pass)
| Phantom image scores: |  |  |  | Fibers |  | Specks | Masses
2. Artifact Detection
3. Printer Check (if applicable)
4. SNR, CNR and AEC Repeatability
| Measured values: |  |  | SNR |  | CNR
CV for mAs and entrance air kerma <5%
Max deviation of mean pixel values and SNR within ±15% of mean for measurements
5. Radiation Dose
| Average glandular dose for average breast is <3 mGy (300 mrad) | [7 empty] | mrad
6. Spatial Resolution
7. AEC Test
8. Detector Uniformity
9. Mechanical Tests
10. Acquisition Workstation Monitor Check
11. Site Audit/Evaluation of Technologist QC Program
12. Collimation, Dead Space & Compression Paddle Position
13. HVL and Radiation Output
14. Tube Voltage Measurement & Repeatability
15. Average Glandular Dose (DBT)
16. Geometric Accuracy in X and Y Direction and Z-Resolution (DBT) (Optional Revelation)
17. Radiation Field (DBT)
18. System Imaging Quality (DBT)
≥ 4 fibers, ≥ 3 speck groups and ≥ 3 masses must be visible
| Phantom image scores: |  |  |  | Fibers |  | Specks | Masses
19. Artifact Detection (DBT)
20. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on N22: Click in boxes to use drop-down lists
Note on H25: Click in boxes to use drop-down lists
Note on H48: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Siemens
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Siemens, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY |  | PASS/FAIL
1. | Phantom Image Quality |  | Novation & Fusion-Daily; Inspiration-Weekly
2. | Artifact Detection | Weekly
3. | SNR and CNR Measurements | Weekly
4. | Detector Calibration* |  | Novation-Weekly; Fusion-Quarterly
5. | Repeat/Reject Analysis | Quarterly
6. | Compression Force | Semi-annually
7. | System Imaging Quality (DBT) | Weekly
8. | Printer Check (if applicable) | Daily, when images printed
9. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
10. | Mobile Unit Quality Control (if applicable) | After every move
* For Mammomat Revelation and Inspriation, indicate NA-calibration required before QC but does not need to be documented
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for Siemens FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Mammomat Fusion
Mammomat Inspiration
Mammomat Inspiration Tomosynthesis
Mammomat Novation
Mammomat Revelation