## QC Test Summary-Carestream
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Carestream CR
Site Name | [6 empty] | Report Date
Address | [6 empty] | Survey Date
Medical Physicist's Name | [6 empty] | Signature
X-Ray Unit Manufacturer | [6 empty] | Model
Date of Unit Installation | [6 empty] | Room ID
CR Image Receptor Mfr |  |  | Carestream Health Inc. |  |  |  | CR Model
QC Manual Version # | [5 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [6 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Cassette Exposure Response Test
2. Cassette/Phosphor Screen Artifact Test
3. Erased Screen Test
4. Scanner Uniformity Test
5. Scanner Response Linearity Test
6. Spatial Frequency Response Test
7. Geometric Accuracy Test
8. Image Plate Fog Test
9. AEC System Performance/Constancy Test
AEC Function with Breast Thickness
AEC Density Control Function
CNR Tracking with Breast Thickness
10. Mammography Unit Assembly Evaluation
11. Mammographic Unit Collimation Assessment | [6 empty] | Test date if different from above:
X-ray Field to Light Field Alignment
X-ray Field to Image Plate Alignment
Alignment of Chest Wall Edge of Compression Paddle to Image Plate
12. Beam Quality and Half-Value Layer | [6 empty] | Test date if different from above:
13. kVp Accuracy and Reproducibility | [6 empty] | Test date if different from above:
Measured average kVp within +5% of indicated kVp
kVp coefficient of variation <0.02
14. Breast Entrance Exposure, Dose and Radiation Output
Average glandular dose for average breast is below 3 mGy (300 mrad)
Average glandular dose to a 4.2-cm-thick breast on your unit is
Radiation output rate is >800 mR/sec (7.0 mGy/sec) at 28 kVp with Mo/Mo
15. Phantom Image Quality Evaluation |  |  |  | Fibers | Specks | Masses | (at least 4 fibers, 3 speck groups & 3 masses)
Phantom IQ (printed images)
Phantom IQ (softcopy)
16. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
17. Film Printer Test (for all printers, even if located offsite)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on I48: Click in boxes to use drop-down lists
Note on M22: Click in boxes to use drop-down lists
Note on M49: Click NA if not used for final interpretation
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Carestream
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Carestream, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Erase Cassettes | Daily
2. | Phantom Test | Weekly
3. | CNR Check | Weekly
4. | Cassette/Phosphor and CR System Visual Inspection and Cleaning | Monthly
5. | Mammography Unit Visual Checklist | Monthly
6. | Repeat Analysis | Quarterly
7. | Compression | Semi-annually
8. | Image Plate Fog Test | Semi-annually
9. | Film Printer QC | See FDA guidance
10. | Review Workstation QC-Overall | See FDA guidance
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. Carestream specifies that the printer and monitor manufacturers' QC program must be followed; if no manufacturer-provided QC is available, the Careteam manual includes testing information. 4. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above.
Note on I10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5