## QC Test Summary-Fischer
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Fischer
Site Name | [6 empty] | Report Date
Address | [6 empty] | Survey Date
Medical Physicist's Name | [6 empty] | Signature
X-Ray Unit Manufacturer |  | Fischer |  |  |  |  | Model
Date of Installation | [6 empty] | Room ID
QC Manual Version #: |  |  |  |  | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [6 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. X-Ray Field, Chest Wall Missed Tissue and Light Field Checks
2. Compression Paddle Alignment
3. kVp Accuracy Test
4. Linearity, Reproducibility and Accuracy
5. Half-Value Layer and Output
6. Dosimetry – Average Glandular Dose and Output
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [6 empty] | mrad
7. Phantom Image Acquisition Test (values required for all tests)
No obvious artifacts
| Background StDev | [6 empty] | within +50/-0 ADU counts of baseline (NA for Equipment Evaluations)
| Background mean | [6 empty] | within ±100 ADU counts of baseline (NA for Equipment Evaluations)
| ADU level difference | [6 empty] | within ±300 ADU counts of baseline (NA for Equipment Evaluations)
8. Image Quality
| Largest 4 fibers, 3 speck groups and 3 masses |  |  |  | Fibers | Specks | Masses
Phantom IQ Test on Review Work Station
9. System Resolution/Scan Speed Uniformity
10. Flat Field Test
11. Geometric Distortion and Resolution Uniformity
12. Automatic Decompression Control
13. System Artifacts
14. Image Viewing Room Illuminance Test (≤50 lux)
15. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on J35: Click in boxes to use drop-down lists
Note on L21: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Fischer
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Fischer, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Laser Image Quality Test | Daily
2. | Phantom Image Acquisition Test | Weekly
3. | Phantom Image Quality Test | Weekly
4. | Detector Calibration and Flat-Field Test | Weekly
5. | System Resolution (Detector Alignment/Scan Speed Uniformity) | Monthly
6. | System Operation | Monthly
7. | Reject/Repeat Analysis | Quarterly
8. | Compression Force Test | Semi-annually
9. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the manufacturers for the QC manual and clearance status of their products.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above.
Note on J10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5