## QC Test Summary-Fuji CR
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Fuji CR
Site Name | [6 empty] | Report Date
Address | [6 empty] | Survey Date
Medical Physicist's Name | [6 empty] | Signature
X-Ray Unit Manufacturer | [6 empty] | Model
Date of Unit Installation | [6 empty] | Room ID
CR Image Receptor Mfr |  |  | Fuji |  |  |  | CR Model
QC Manual Version #: | [5 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [7 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. S Value Confirmation (<120 +20% [96 < corrected S value < 144])
2. System Resolution (8 lp/mm +2 lp/mm in both directions)
3. CR Reader Scanner Performance
4. Mammography Unit Assembly Evaluation
5. Collimation Assessment | [6 empty] | Test date if different from above:
Chest wall edge of X-ray field extends to edge of IR
Deviation between X-ray field and light field <2% of SID
X-ray field does not extend beyond any side of the IR by more than 2% of SID
Paddle chest wall edge not beyond IR by more than 1% of SID or appear on the image
6. Automatic Exposure Control (AEC) System Performance Assessment
AEC density control function meets Fuji performance criteria
Reproducibility (CV) for either exposure or mAs is <0.05
Image mode tracking meets Fuji performance criteria
CNR per object thickness meets Fuji performance criteria
7. System Artifact Evaluation
8. Phantom Image Quality Evaluation |  |  |  | Fibers | Specks | Masses | (at least 4 fibers, 3 speck groups & 3 masses)
Phantom IQ (printed images)
Phantom IQ (softcopy)
Other tests meet Fuji performance criteria (mAs, OD & DD for hardcopy, S value for soft copy)
9. Dynamic Range
10. Primary Erasure (Additive and Multiplicative Lag Effects)
11. Inter-Plate Consistency (variation of mAs within +10%; SNR within +15%)
12. kVp Accuracy and Reproducibility | [6 empty] | Test date if different from above:
Measured average kVp within +5% of indicated kVp
kVp coefficient of variation <0.02
13. Dose (average glandular dose for average breast is <3 mGy [300 mrad]) | [8 empty] | mrad
14. Beam Quality Assessment & HVL Measurement | [6 empty] | Test date if different from above:
HVL >kVp/100 mm Al
15. Radiation Output | [6 empty] | Test date if different from above:
Radiation output rate is >800 mR/sec (7.0 mGy/sec) at 28 kVp with Mo/Mo
16. Viewing and Viewing Conditions | [6 empty] | Test date if different from above:
Room illuminance <20 lux or as recommended by the monitor manufacturer
17. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
18. Film Printer Tests (for all printers used for mammography, even if located offsite)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on I38: Click in boxes to use drop-down lists
Note on M22: Click in boxes to use drop-down lists
Note on M39: Click NA if not used for final interpretation
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Fuji CR
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Fuji CR, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | CNR Weekly Check | Weekly
2. | Phantom Image | Weekly
3. | Visual Checklist | Monthly
4. | Repeat Analysis | Quarterly
5. | Compression | Semi-annually
6. | Imaging Plate (IP) Fog | Semi-annually
7. | Film Printer QC | See FDA guidance
8. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. Fuji specifies that the printer and monitor manufacturers' QC program must be followed; if no manufacturer-provided QC is available, the Fuji manual includes testing information. 4. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 5. See the FDA-approved alternative standard for Fuji FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on I10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5