## QC Test Summary-Giotto
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Giotto
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer |  | Giotto | [5 empty] | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [6 empty] | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  | Manufacturer |  |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Collimation Assessment
2. Artifact Evaluation
3. Automatic Exposure Control (AEC), Signal-To-Noise Ratio (SNR) & Contrast-To-Noise Ratio (CNR)
| SNR (value) |  | ≥60
| CNR (value) |  | (Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
CNR should not vary by more than ±20% (NA for Equipment Evaluation)
4. AEC Reproducibility
5. ACR Phantom Image Quality
| Phantom image scores: |  |  | Fibers |  | Specks |  | Masses
6. Ghost Factor
7. Inactive Border at Chest Wall Edge
8. Flat Field Homogeneity
9. Detector Response Function and Noise Evaluation
10. Spatial Resolution
11. kVp Accuracy and Reproducibility
12. Tube Output
13. Exposure Time
14. Beam Quality (HVL)
15. Mean Glandular Dose (MGD)
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
16. Viewbox Luminance
17. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
18. Film Printer QC
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on G29: Click in boxes to use drop-down lists
Note on M21: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Giotto
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Giotto, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Daily Check | Daily
2. | Artifact Evaluation | Weekly
3. | ACR Phantom Image Quality | Weekly
4. | Flat-Field Homogeneity | Weekly
5. | Signal-To-Noise (SNR) and Contrast-To-Noise (CNR) Ratio | Weekly
6. | Repeat/Reject Analysis | Quarterly
7. | Automatic Exposure Control System (AEC) | Semi-annually
8. | Compression Force | Semi-annually
9. | Mobile Unit Quality Control (if applicable) | After every move
10. | Diagnostic Review Workstation QC (NA if only hardcopy read) | See FDA Guidance
11. | Film Printer QC | See FDA Guidance
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the manufacturers for the QC manual and clearance status of their products.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5