## A
2. Collimation Assessment
Source to image receptor distance (SID): |  |  |  | cm
Deviation between X-ray field and light field:
Target material |  | Mo | Mo
Collimator (cm) |  | 18x24 | 24x30
Left edge deviation
Right edge deviation
Sum of left and right edge deviations
Sum as % of SID
Anterior edge deviation
Chest edge deviation
Sum of anterior and chest edge deviations
Sum as % of SID
ACTION LIMIT: | ACR/MQSA - If sum of left plus right edge deviations or anterior plus chest
edge deviations exceeds 2% of SID, seek service adjustment.
Deviation between X-ray field and edges of the image receptor:
Left edge deviation
% of SID (retain sign)
Right edge deviation
% of SID (retain sign)
Anterior edge deviation
% of SID (retain sign)
Chest edge deviation
% of SID (retain sign)
ACTION LIMIT: | ACR/MQSA - If X-ray field exceeds image receptor at any side by more
than 2% of SID or if X-ray field falls within image receptor on the chest wall
side, seek service adjustment.
ACR - If X-ray field falls within image receptor by more than -2% on the
left and right sides, or by more than -4% on the anterior side, seek service
adjustment.
Alignment of chest-wall edges of compression paddle and film:
Difference between paddle edge and film
Difference as % of SID
ACTION LIMIT: | ACR/MQSA - If chest-wall edge of compression paddle is within the
image receptor or projects beyond the chest-wall edge of the image
receptor by more than 1% of SID, seek service correction.