## A
4. Automatic Exposure Control (AEC) System Performance
| AEC position: |  |  |  | Density control:
| Small cassette ID: |  |  |  | Large cassette ID:
Performance Capability:
Thickness-kVp Tracking
| Imaging mode: | small image receptor with grid
| Focal spot: | large focal spot
mA:
| Phantom thickness | Image ID# | AEC Mode | Target-Filter | kVp | Density Control Setting | mAs | Film Optical Density
2 cm
4 cm
6 cm
8 cm
| Mean Optical Density (2-6 cm) |  | Optical Density Range |  |  | MQSA Allowed Optical Density Range
|  |  | to |  |  |  | to
Image Mode Tracking
mA:
Phantom thickness:
| Target-Filter: | Mo/Mo
| Image Mode | Image ID# | AEC Mode | Focal Spot | kVp | Density Control Setting | mAs | Film Optical Density
small grid
large grid
Magnification/no grid
Overall AEC Performance
| Mean Optical Density | Optical Density Range |  |  | Recommended Optical Density Range
|  |  | to |  |  | to
| ACTION LIMIT: | ACR - The AEC system should be able to maintain constant film optical density
to within +0.30 of the average over the phantom thicknesses and imaging modes
tested.
ACR/MQSA - The AEC system must be capable of maintaining film optical
density within +0.30 of the mean (+0.15 after 10/28/2002) when the thickness of
the phantom is varied over 2 - 6 cm and the kVp is varied over the range of
those used clinically for those thicknesses. The optical density in the center
of the phantom image must not be less than 1.20. If these standards are not met,
seek service adjustment.
Density Control Function:
| Imaging mode: | small grid |  |  | Focal spot: |  | large
| mA: |  |  |  | kVp:
| Phantom thickness: |  |  |  | Cassette ID:
Relative to Normal
| Density Control Setting | Image ID# | mAs | Measured Optical Density |  |  | % mAs change | Optical Density Change
| -4 | [5 empty] | * | *
-3
-2
-1
0 (normal)
+1
+2
+3
+4
| ACTION LIMIT: | ACR - Each step should result in a 12 to 15% change in mAs, or approximately a
0.15 change in film optical density. If not, seek service.