## A
7. Image Quality Evaluation
Phantom used:
AEC detector position:
Cassette size:
Cassette #:
|  | Previous Film | Current Film | Comments
Date
kVp setting
Density control
Phototimed mAs
mAs change
% mAs change = mAs change/mAsx100
Background density
Background density change
Density outside disc
Density inside disc
Density diff = outside-inside
Density difference change
Number of fibers seen
Fibers seen after deduction
Fiber change
Number of speck groups seen
Speck groups after deduction
Speck group change
Number of masses seen
Masses seen after deduction
Mass change
ACTION LIMIT: | ACR/MQSA - The largest 4 fibers, 3 speck groups, and 3 masses
must be visible. Background optical density must be at least 1.20.
Corrective action must be taken before any further examinations
are performed if the results of this test fail any MQSA regulations.
ACR - The density difference should be at least 0.40 for a 4-mm
thick acrylic disk. Background optical density should be at least
1.40 and must be at least 1.20. If % mAs change exceeds +15%,
if background density change exceeds +0.20, if density difference
change exceeds +0.05, or if fiber, speck group or mass score
decreases by more than 0.5, the source of change should be
identified and corrected.