## TOC
Digital Mammography Quality Control Tests
Radiologic Technologist's Tests (2D and DBT)
Important: Before a facility may start using the procedures in the ACR Digital Mammography QC Manual for the first time on a unit, the medical physicist must first conduct an annual survey of the digital mammography unit and display devices using the manual and the ACR Digital Mammography Phantom.
Note: Complete Facility, Unit and Test Equipment Data tab first to populate facility information into forms
| Test* | Minimum Frequency** | Corrective Action Timeframe***
1. | ACR Digital Mammography Phantom Image Quality | Weekly | Before clinical use
2. | CR Cassette Erasure (if applicable) | Weekly | Before clinical use
3. | Compression Thickness Indicator | Monthly | Within 30 days
4. | Visual Checklist | Monthly | Critical: before clinical use; less critical: w/in 30 days
5. | Acquisition Workstation (AW) Monitor QC | Monthly | W/in 30 days; before clinical use for severe defects
6. | Radiologist Workstation (RW) Monitor QC | Monthly | W/in 30 days; before clinical use for severe defects
7. | Film Printer QC (if applicable) | Monthly | Before clinical use
8. | Viewbox Cleanliness (if applicable) | Monthly | Before clinical use
9. | Facility QC Review | Quarterly | Not applicable
10. | Compression Force | Semiannual | Before clinical use
11. | Manufacturer Calibration (if applicable) | Mfr. Recommendation | Before clinical use
| Optional - Repeat Analysis | As Needed | Within 30 days after analysis
| Optional - System QC for Radiologist | As Needed | W/in 30 days; before clinical use for severe artifacts
| Optional - Radiologist Image Quality Feedback | As Needed | Not applicable
* All required tests (except Facility QC Review) must be performed upon installation of new equipment and before clinical use.
** This is a minimum frequency; tests may be performed more often if problems are noted. Also, weekly tests do not need to be performed if mammography is not performed during that week. However, the test must be performed prior to examining patients once mammography resumes. In these cases, be sure to note in the QC charts that mammography was not performed during this time period.
*** Corrective action for MEEs must be performed before clinical use.
Management Forms
ACR Technique and Procedure Summaries
Corrective Action Log
Facility Offsite Display Locations
Digital Mammography Unit QC Summary Checklist
Facility Display Device QC Summary Checklist
Mobile Systems
In addition to meeting the minimum frequencies outlined in the table above, the following tests must be performed, evaluated, and pass after each move of the mobile system to a new location:
• ACR Digital Mammography Phantom Image Quality - after each move and prior to examining patients
• Compression Thickness Indicator - after each move and prior to interpretation
• Radiologist Workstation (RW) Monitor QC (mobile RW only) - after each move and prior to interpretation
• Film Printer QC (mobile film printers only) - after each move and prior to printing patient images
QC Equipment List - Technologist
| ACR Digital Mammography Phantom | Scale | Appropriate monitor cleaning materials
| Densitometer | Towels

## Instr-Rev
Radiologic Technologist Test Form Spreadsheet
Instructions
• | The forms may be completed on paper or on computer by using the Excel spreadsheets.
• | If completing via computer, be sure to verify all entries so they are correct.
• | Before completing the forms, either on paper or by computer, enter facility information in Fac Info tab first; information auto-populates in other pages.
• | If completing the forms on computer, drop down lists are available within some cells to complete the required information; if the list does not contain the necessary information; complete the cell yourself.
• | Your system may have multiple, clinically-used Image Modes that must be tested (e.g., 2D, 2D w/Add-on DBT Device, DBT). See next tab for the required tests.
If completing via paper, write in the Image Mode that is being tested on the top of all applicable forms. Make additional copies of the forms for each image mode tested.
If completing via computer, make a copy of the tab with each form for all applicable image modes tested. Select the Image Mode tested on each form using the drop-down list of Image Modes.
• | Enter data into white cells.
• | Light yellow cells are for pass/fail results of subtests within the main test; depending on the page, you may enter the pass/fail results using the dropdown list.
• | Bright yellow cells are for pass/fail results of the overall test; depending on the page, you may enter the pass/fail results using the dropdown list.
• | Light blue cells contain formulae - DO NOT enter data in these cells; if you do, the formulae and results will be lost.
• | Do not complete gray cells.
Revisions
| 2016-09-08 | New
| 2017-01-13 | Misc typo and formatting corrections
| 2017-08-11 | Instructions-Revisions tab moved from end to beginning of the spreadsheet
Misc typo and formatting corrections
| 2018-11-19 | Updated for DBT

## Modes
Required Tests for Imaging Modes Used on 2D and DBT Systems

## Fac Info
Facility, Unit and Test Equipment Data
Radiologic Technologist's Tests
Facility Information
Facility Name
Address
Address
City, State, Zip
MAP ID# (00000)
MAP ID Unit# (00)
Lead Interpreting Radiologist
Quality Control Technologist
Unit Information
Digital Radiography (DR)
Digital Breast Tomosynthesis (DBT)
Computed Radiography (CR)
Room ID
X-Ray Unit Manufacturer
X-Ray Unit Model
X-Ray Unit Gantry Serial #
X-Ray Unit Date of Manufacture
X-Ray Unit Date of Installation
CR Reader Manufacturer
CR Reader Model
CR Reader Serial #
CR Date of Manufacture
CR Date of Installation
Test Equipment Info
ACR DM Phantom Manufacturer and S/N
Medical Physicist Info
Medical Physicist Name
Telephone Number
email
Note on B14: Select dot or check from dropdown, or type capital "P"
Note on B15: Select dot or check from dropdown, or type capital "P"
Note on B16: Select dot or check from dropdown, or type capital "P"

## 1.
1. ACR DM Phantom Image Quality | [14 empty] | Weekly
Image Mode (2D, 2D w/Add-on DBT, DBT)
|  |  |  | Facility | 0 | [6 empty] | Room ID | 0
|  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0 |  |  | Unit Mfr & Model
Year
Date (month & day)
Tech Initials
Resulting Techniques | Image receptor size
View or selected image
Slice or slab # (DBT only)
| AEC mode | [7 empty] | 0 |  | 0 | 0 |  | 0
Target/filter
kVp
mAs
ACR DM Phantom | Artifacts P/F
Fiber score
Speck group score
Mass score
Overall Pass/Fail
[12 empty] | P = Pass |  | F = Fail
Analyses: | Scoring |  | Full Point |  |  |  |  | Half Point
|  | Fibers | ≥ 8 mm long |  |  |  |  | ≥ 5 and < 8 mm long
|  | Specks | 4 - 6 specks |  |  |  |  | 2 - 3 specks
|  | Masses | ≥ ¾ border |  |  |  |  | ≥ ½ & < ¾ border
Action Limits | Required: |  | ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥ 3.0; mass score must be ≥ 2.0.
| Timeframe: |  | Required items must be corrected before clinical use.

## 2.
2. Computed Radiography (CR) Cassette Erasure | [15 empty] | Weekly
(if applicable) | [13 empty] | Image Mode | 2D
| Facility | 0 | [10 empty] | CR Room
|  |  |  | MAP ID-Unit# (00000-00) | 0 |  |  | - | 0 |  |  | CR Reader Mfr & Model
Year
Month | Jan |  |  |  |  | Feb |  |  |  | Mar | Apr |  | May |  | Jun
Date
Tech Initials
Month | Jul |  |  |  |  | Aug |  |  |  | Sep | Oct |  | Nov |  | Dec
Date
Tech Initials
[9 empty] | Date and initial each test - | date
initial
Action Taken on Cassette (include date & cassette #):
Action Limits |  |  | Required: |  |  |  | Each cassette should successfully pass the CR cassette erasure process
|  |  | Timeframe: |  |  |  | Failing cassettes must be corrected before clinical use.

## 3.
3. Compression Thickness Indicator | [17 empty] | Monthly
Image Mode (2D, 2D w/Add-on DBT, DBT)
| Facility | 0 | [10 empty] | Room ID | 0
| MAP ID-Unit# (00000-00) | 0 |  | - | 0 |  |  |  |  | Unit Mfr & Model
Year
Month |  | Jan | Feb |  | Mar | Apr |  | May | Jun |  | Jul | Aug | Sep |  | Oct | Nov | Dec
Date
Tech Initials
Description of compression thickness indicator phantom
Actual thickness of phantom
[5 empty] | cm |  | mm |  | (Use the same unit displayed on the indicator)
Indicated thickness
Difference between indicated and actual thicknesses (Indicated - Actual) |  | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0
Overall Pass/Fail
[13 empty] | P = Pass |  |  | F = Fail
Action Limits |  | Required: |  |  | Compression thickness indicator must be accurate to within ±0.5 cm (±5 mm) of the actual thickness.
|  | Timeframe: |  |  | Failures must be corrected within 30 days.
Note on F13: Select dot or check from dropdown, or type capital "P"
Note on I13: Select dot or check from dropdown, or type capital "P"

## 4.
4. Visual Checklist | [13 empty] | Monthly
Image Mode (2D, 2D w/Add-on DBT, DBT, All)
|  | Facility | 0 | [6 empty] | Room ID | 0
|  | MAP ID-Unit# (00000-00) | 0 | - | 0 |  |  | Unit Mfr & Model
Procedure |  | Inspect the unit and evaluate the functionality according to the checklist below.
Year
Month |  |  | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Tech Initials
Room Cleanliness | Mag stands and paddles free from dust
Room and countertops free from dust
Cleaning solution available*
X-Ray Unit | Indicators working
Locks (all)*
Collimator light working
Are cables safely positioned
Smoothness of C-arm motion
Smoothness of compression paddle motion
Paddles/face shields not cracked*
Breast support not cracked*
CR (if app) | Cassette holder and lock (small and large)*
Condition of imaging plates and cassettes*
Scanning Detector System (if app)
DBT (if app) | DBT assembly moves as designed*
Other
[7 empty] | P = Pass |  | F = Fail |  | NA = Not Applicable
Action Limits |  | Required: | Jan: All items, both critical (*) and noncritical, must pass.
|  | Timeframe: | Jan: Failures of critical items (*) must be corrected before clinical use; less critical items must be corrected within 30 days.

## 5.
5. Acquisition Workstation (AW) Monitor QC | [15 empty] | Monthly
Image Mode (2D, DBT)
| Facility | 0 | [8 empty] | Room ID | 0
| MAP ID-Unit# (00000-00) | 0 |  | - | 0
Year
Month |  | Jan | Feb |  | Mar | Apr | May | Jun | Jul | Aug | Sep |  | Oct | Nov | Dec
Date
Tech Initials
Monitor Condition P/F (significant findings)
Test Pattern Image Quality (if available) | 0%-5% contrast boxes visible?
95%-100% contrast boxes visible?
Line-pair images distinct (center)?
Line-pair images distinct (corners)?
Test pattern P/F
Monthly Check - Mfr Automated Test P/F (if avail)
Overall Pass/Fail
[11 empty] | P = Pass |  |  | F = Fail
Action Limits |  | Required: |  |  | Any identified screen blemish that could interfere with clinical information must be removed.
Test pattern image quality must pass all visual tests.
Manufacturer's automated tests, if available, must pass mfr specifications (if 1 test fails, indicate F).
|  | Timeframe: |  |  | Significant monitor cleanliness defects must be corrected before clinical use; all other required tests must be corrected within 30 days.

## 6.
6. Radiologist Workstation (RW) Monitor QC | [24 empty] | Monthly
RW Location and ID
|  | MAP ID# (00000) | 0 |  |  |  |  | Monitor Mfr |  |  |  |  | Model |  |  |  |  | SN: | Right |  |  |  | Left
Year
Month |  | Jan |  | Feb |  | Mar |  | Apr |  | May |  | Jun |  | Jul |  | Aug |  | Sep |  | Oct |  | Nov |  | Dec
Date
Tech Initials
Monitor |  | R* | L* | R | L | R | L | R | L | R | L | R | L | R | L | R | L | R | L | R | L | R | L | R | L
Monitor Condition P/F (significant findings)
ACR DM Phantom | Artifacts P/F
Fiber score
Speck group score
Mass score
Phantom P/F
Test Pattern Image Quality | 0%-5% contrast boxes visible
95%-100% contrast boxes visible
Line-pair images distinct (center)
Line-pair images distinct (corners)
Test pattern P/F
Monthly Check - Mfr Automated Test P/F (if avail)
Overall Pass/Fail
[22 empty] | P = Pass |  | F = Fail
Action Limits |  | R*: Required:; R: Any identified monitor blemish that could interfere with clinical information must be removed.; R: * R and L - right and left monitors; if only 1 monitor, use "R" column
ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥ 3.0; mass score must be ≥ 2.0.
Test pattern image quality must pass all visual tests.
Manufacturer's automated tests, if available, must pass mfr specifications (if 1 test fails, indicate F).
|  | Timeframe: |  | Phantom must pass and significant monitor cleanliness defects must be corrected before clinical use; all other tests must be corrected within 30 days.

## 7.
7. Film Printer QC (if applicable) | [17 empty] | Monthly
Film Printer Location and ID
Film Printer and Model
Workstation for printing | [10 empty] | Film size
Procedure |  | Applicability: If film printer is used clinically for mammography (i.e., for interpretation and to provide images to referring physicians and patients)
Equipment: Densitometer
Print an ACR DM Phantom image acquired from any DM unit within facility network.
Do not change window/level settings from acquired image prior to printing.
Print the phantom image from the workstation/computer typically used to print clinical films.
Dmax should be measured either at extreme left or right edge of film or at extreme non-chest wall edge.
Year
Month |  |  |  | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep |  | Oct | Nov |  | Dec
Date
Tech Initials
ACR DM Phantom | Artifacts P/F
Fiber score
Speck group score
Mass score
Phantom P/F
Back-ground | Bkgd OD (Outside cavity)
Bkgd OD ≥ 1.6 (P/F)
Contrast | Cavity OD
| Bkgd OD (use value from above) |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0
| Contrast = Cavity OD - Bkgd OD |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0
Contrast ≥ 0.1 (P/F)
Dmax | Dmax OD
Dmax OD ≥ 3.1 (P/F)
Overall Pass/Fail
[13 empty] | P = Pass |  |  | F = Fail
Action Limits |  | Required: | The ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥ 3.0; mass score must be ≥ 2.0.
Background OD must be ≥ 1.6 (1.7 to 2.2 is recommended; approx 2.0 is optimal).
Contrast (Cavity OD - Background OD) must be ≥ 0.1.
Dmax must be ≥ 3.1 (≥ 3.5 is recommended).
|  | Timeframe: | Failures of required items must be corrected before printing clinical images.

## 8.
8. Viewbox Cleanliness (if applicable) | [11 empty] | Monthly
Viewbox Location and ID
Procedure | Required Equipment: Viewbox manufacturer-recommended cleaner; soft paper or cotton towels
Clean viewbox surfaces and assure all marks have been removed.
Visually inspect the viewboxes for uniformity of luminance and assure masking equipment is functioning.
Year
Month | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Tech Initials
Viewbox Designation
[9 empty] | P = Pass |  | F = Fail
Action Limits | Required: |  | Viewboxes must be clean and free of marks and uniform in brightness.
| Timeframe: |  | Failures must be corrected before clinical images are viewed on the viewbox.

## 9.
9. Facility QC Review | [18 empty] | Quarterly
Image Mode (2D, 2D w/Add-on DBT, DBT)
|  | Facility | 0 | [11 empty] | Date of QC Mtg
Reviewed
1. | Review Medical Physics Surveys and Results
[7 empty] | Room 1 |  |  | Room 2 |  | Room 3 |  | Room 4 |  |  | Room 5
Room ID
Date of last Medical Physicist (MP) survey
MP DM QC Test Summary reviewed by radiologist?
All MP corrective actions completed?
ACR DM Phantom Average Glandular Dose (mGy)
Fiber Score
Speck Score
Mass Score
2. | Review Tech QC
|  | Test |  | Frequency |  |  |  | Summary Comments from Last Quarter
| 1. | ACR DM Phantom Image Quality |  |  | Weekly
[6 empty] | Room 1 |  |  | Room 2 |  | Room 3 |  | Room 4 |  |  | Room 5
|  | Scores of most recent phantom image: |  |  | Date
Fiber score
Speck group score
Mass score
| 2. | CR Cassette Erasure (if app) |  |  | Weekly
| 3. | Compression Thickness Indicator |  |  | Monthly
| 4. | Visual Checklist |  |  | Monthly
| 5. | AW Monitor QC |  |  | Monthly
| 6. | RW Monitor QC |  |  | Monthly
| 7. | Film Printer QC |  |  | Monthly
| 8. | Viewbox Cleanliness (if app) |  |  | Monthly
| 9. | Facility QC Review |  |  | Quarterly
| 10. | Compression Force |  |  | Semiannual
| 11. | Manufacturer Calibrations (if app)
|  | Optional - Repeat Analysis |  |  | As Needed |  |  |  | % Repeats
3. | Review and verify completion of all "Corrective Action"
4. | Technique Chart review for each room (see MP report for recommendations) - (Annually)
5. | Infection Control procedures followed
6. | Offsite RW(s) & Film Printer(s) QC reviewed
7. | Past and future service or service upgrades discussed (if app)
8. | Past and future State and/or MQSA inspections discussed (if app)
9. | Past and future ACR Accreditation issues discussed (if app)
Note on AB7: Select dot or check from dropdown, or type capital "P"

## 9.cont
9. Facility QC Review (cont) | [10 empty] | Quarterly
|  | Facility | 0 |  |  |  |  | Date of QC Mtg
Follow-up Confirmed (If App.)
10. | Notable findings during QC meeting
11. | Items for quality improvement from QC Meeting
12. | Other QC Notes
Overall Pass/Fail
Lead Interpreting Radiologist | [5 empty] | Facility Manager (If App) | QC Technologist
signature | [5 empty] | signature | signature
Action Limit: |  |  |  | Required: | Lead interpreting radiologist and facility manager must review QC quarterly.
The test passes if meeting held.
|  |  |  | Recommended: | Technologist and lead interpreting radiologist should review technique charts at least annually for each DM system.
|  |  |  | Timeframe: | Not applicable.

## 10.
10. Compression Force | [7 empty] | Semiannual
Image Mode (2D, 2D w/Add-on DBT, DBT)
Facility | 0 |  |  |  | Room ID | 0
Unit Mfr & Model |  |  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0
Procedure |  | Required Equipment: Bathroom scale; towels
Place towel on detector, place bathroom scale on towel with dial or read-out positioned for reading.
Place another towel on top of scale.
Using manual fine-adjustment mode, activate compression until at least 25 pounds reached.
Read and record the compression force.
Using initial power-drive mode, activate compression until it stops automatically.
Read and record the compression force.
Check that force is maintained.
Note: Ensure the scale used produces accurate readings as pressure is increased.
Year
Date (month & day)
Tech Initials
|  | Compression Force |  | Units | Compression Force |  | Units
Manual fine-adjustment compression force | [6 empty] | 0
Force is at least 25 lbs (11.1 daN) P/F
Initial power-drive compression force |  |  |  | 0 |  |  | 0
Force is at least 25 lbs (11.1 daN) but no more than 45 lbs (20.0 daN) P/F
Compression remains at least 25 lbs (11.1 daN) throughout typical exposure P/F
Overall Pass/Fail
|  | Enter number where appropriate. |  |  |  |  | P = Pass | F = Fail
|  | Legend: | lbs = pounds
daN = decanewton
Action Limit | Required: | Manual fine-adjustment compression force must be least 25 lbs (11.1 daN).
Initial power-drive compression force must be at least 25 lbs (11.1 daN) but no greater than 45 lbs (20 daN).
Compression remains at least 25 lbs (11.1 daN) throughout typical exposure.
| Timeframe: | Failures must be corrected before further examinations are performed.

## 11.
11. Manufacturer Calibrations (if applicable)
Image Mode (2D, 2D w/Add-on DBT, DBT)
Facility | 0 |  |  | Room ID | 0
MAP ID-Unit# (00000-00) | 0 | - | 0 | Unit Mfr & Model
Manufacturer Procedure | Frequency:
Frequency: Per manufacturer recommendations (if app).
Note: See medical physicist and manufacturer for instructions (if any).
Year
Date (month & day)
Tech Initials
Name of Calibration
[5 empty] | P = Pass | F = Fail
Action Limits | Required: | Unit must pass manufacturer's prescribed periodic calibrations.
| Timeframe: | Failures must be corrected before further examinations are performed.

## O-Rpt
Optional - Repeat Analysis - Summary Form | [7 empty] | As Needed
| Facility | 0 |  |  |  | Year
| MAP ID (00000) | 0
Procedure | Required Equipment: All repeated mammograms and means to count and sort them
Record the total number of exposures for the collection period (month or quarter).
Record the total number of repeat exposures for that time period.
Calculate by hand, or use formulas in spreadsheet, to calculate repeat rate.
Note: Some units may automatically calculate % Repeats. If so, enter this number into "% Repeats".
| Monthly Analysis |  |  |  | Quarterly Analysis
| Total # of Exposures | # of Repeat Exposures | % Repeats | Pass or Fail | Total # of Exposures | # of Repeat Exposures | % Repeats | Pass or Fail
Jan |  |  | #DIV/0! |  |  |  | #DIV/0!
Feb |  |  | #DIV/0!
Mar |  |  | #DIV/0!
Apr |  |  | #DIV/0! |  |  |  | #DIV/0!
May |  |  | #DIV/0!
Jun |  |  | #DIV/0!
Jul |  |  | #DIV/0! |  |  |  | #DIV/0!
Aug |  |  | #DIV/0!
Sep |  |  | #DIV/0!
Oct |  |  | #DIV/0! |  |  |  | #DIV/0!
Nov |  |  | #DIV/0!
Dec |  |  | #DIV/0!
| % Repeats = (# of Repeat Exposures / Total # of Exposures) * 100 | [5 empty] | P = Pass | F = Fail
Action Limits | Recommended: |  | If repeat rate changes from the previously determined rate by more than 2.0% of the total images included in the analysis, the reason(s) for the change must be determined.
| Timeframe: |  | Failures must be corrected within 30 days after analysis.

## O-RptA
Optional - Repeat Analysis A - Tally Sheet
Facility | 0 |  | Date Start (month/day/yr)
MAP ID (00000) | 0 |  | Date End (month/day/yr)
Procedure |  | Record all repeat exposures on the form below.
Transfer the Total Number of Repeats from below to the "Repeat Analysis -Summary Form"
for final calculation of Repeat Analysis.
Total # of images for time period |  |  |  | 0
Reason |  | Comments/Notes | Total # of Repeat Exposures | % Repeats
Patient-Related Repeats:
Poor positioning |  |  |  | #DIV/0!
Patient motion |  |  |  | #DIV/0!
Patient-caused artifacts |  |  |  | #DIV/0!
Incorrect patient ID |  |  |  | #DIV/0!
Technical Repeats:
Exposure too low (excessive noise) |  |  |  | #DIV/0!
Exposure too high (image saturation) |  |  |  | #DIV/0!
Equipment-caused artifacts |  |  |  | #DIV/0!
X-ray equipment failure |  |  |  | #DIV/0!
Software failure |  |  |  | #DIV/0!
Aborted AEC exposure |  |  |  | #DIV/0!
Miscellaneous Repeats:
Blank images |  |  |  | #DIV/0!
Good images (no apparent reason) |  |  |  | #DIV/0!
Other - miscellaneous |  |  |  | #DIV/0!
Do Not Count as Repeats:
Wire localization images |  |  |  | Not Included in Repeat Analysis
I-125 seed localization images
Additional views to image entire breast
Quality control
Total: |  |  | 0 | #DIV/0!
Note: Some equipment manufacturers provide an automated system to collect, record and analyze repeated clinical images. These systems may be used instead of these forms as long as the system includes the following 2 key elements:
1. Count of the total # of exposures made during the evaluation period
2. % Repeats during the same period: (# Repeat Exposures/Total # Exposures)*100

## O-RptB
Optional - Repeat Analysis B - Daily Counting Sheet
Facility | 0 | [13 empty] | Date Start (Month/Day/Yr)
MAP ID (00000) | 0 | [13 empty] | Date End (Month/Day/Yr)
Procedure |  | Record these counts daily.
For each patient, record all images including repeat exposures.
Transfer the Total Number of Repeats from below to the "Repeat Analysis - Tally Sheet" for final calculation of Repeat Analysis.
Patient Name and/or ID |  | Total Number of Images | # of Repeat Expo-sures | Positioning | Patient motion | Patient-caused artifacts | Incorrect patient ID | Exposure too low | Exposure too high | Equipment-caused artifacts | X-ray equipment failure | Software failure | Aborted AEC exposure | Blank images | Good images (no apparent reason) | Other - miscellaneous | Localization images (wire and I-125) | Additional views for entire breast | Quality control
Total |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Note: Some equipment manufacturers provide an automated system to collect, record and analyze repeated clinical images. These systems may be used instead of these forms as long as the system includes the following 2 key elements:
1. Count of the total # of exposures made during the evaluation period
2. % Repeats during the same period: (# Repeat Exposures/Total # Exposures)*100

## O-Rad Sys QC
Optional - System QC for Radiologist | [12 empty] | As Needed
(For Quality Improvement)
[7 empty] | Facility | 0
[7 empty] | MAP ID-Unit# (00000-00) |  |  | -
Procedure |  |  | This test is to be performed or supervised by the lead interpreting radiologist.
The technologist should deliver this form to the radiologist, ensure correct completion, follow-up on any failures, and place the form
into QC notebook.
This test can be performed on the same workstation for multiple DM or CR units.
Example: A radiologist can sit at one workstation and view the images for all the DM and/or CR units within a facility.
The radiologist does not need to evaluate every monitor at every workstation.
| Objective |  | This test is for the Radiologist to perform an evaluation of the entire mammographic imaging chain with the focus being primarily on the detector and secondarily on the monitors. This test is not intended to evaluate technologist issues such as positioning, compression, etc.
| Procedure for Radiologist | [7 empty] | Room ID
DM or CR Unit Mfr & Model
| Step 1: | Complete the demographics: | [6 empty] | Monitor ID
Radiologist Name
Date of Evaluation
| Step 2. | Pull up the recent mammographic study from the above listed DM unit and record ID & Study Date. | [6 empty] | Image ID:
Study Date:
| Step 3. | Place the same MLO image on each monitor.
[11 empty] | Yes | No
| Step 4. | Evaluate the images for artifacts and check the appropriate boxes. |  |  |  | Comparing the monitors (or sides), do the background areas (outside of breast) appear different (darker or lighter, etc.)?
Is there a difference in contrast between monitors/sides?
Does the image contain excessive noise (not patient motion)?
|  | For examples and more detailed descriptions, please see the Guide on Identifying Artifacts. |  |  |  | Do you see ghosting?
Do you see "bad pixels" (singular or clusters) (white or black)?
Do you see white dots that could be from excessive dust?
Do you see any image distortion (not architectural distortion)?
Do you see gridlines?
| Step 5. | If necessary, document any failures on the "Corrective Action Log" form and ensure items are resolved. |  |  |  | Do you see artifacts that could be due to image processing?
Do you see "line artifacts" (single or multiple pixels that form lines extending across image - horizontally or vertically)?
Are there any other artifacts that are present and clinically significant (impeding interpretation)?
Action Limits |  |  |  | Recommended: | If any box is checked "Yes", then seek service.
|  |  |  | Timeframe: | If an image quality problem or artifact impedes clinical interpretation, seek service before further imaging or interpretations are performed.
If the artifact does not impede clinical interpretation, seek service within 30 days.
Left Monitor
Right Monitor

## O-Rad Fdbck
Optional - Radiologist Image Quality Feedback | [17 empty] | As Needed
(For Quality Improvement)
Radiologist's Name
Date
Procedure |  |  | This report is to be completed by the Interpreting Radiologist when asked to interpret sub-optimal cases
requiring the patient to be called back.
The form may also be used to provide feedback on excellent quality.
The radiologists should complete this form as needed for each case.
A system should be in place for analyzing feedback and taking measures for improvement as necessary.
| Objective |  | For the Radiologist to provide routine feedback to the technologists and manager on the quality of images.
Patient Identifier:
Technologist's Name:
Date of Exam:
Overall Assessment
|  | Excellent |  | Good |  |  | Needs improvement, but do not repeat | [7 empty] | Sub-Optimal, and should be repeated
Image Evaluation
[7 empty] | RCC |  | LCC |  | RMLO |  |  | LMLO |  | Other View | Other View
Positioning
Missing tissue
Laterally
Posteriorly
Medially
Inferiorly
Nipple not in profile
Skin fold
Pectoralis not down to PNL
Tissue droopy (camel nose)
Narrow/concave pectoralis
Inframammary fold
Not open
Not shown
Centering not correct
Technical Issues
Not enough compression
Exposure Too Low (Excessive Noise)
Exposure Too High (Image Saturation)
Patient Motion
Artifacts
Incorrect Patient ID
Other
Additional Images Needed for Complete Breast Evaluation
|  |  | Requested views |  | RCC |  |  | LCC |  | RMLO |  | LMLO |  |  | Other View
Action Limits |  |  | Recommended: |  |  | Patients should be called back for additional images if the quality is suboptimal according to the interpreting radiologist's request.
|  |  | Timeframe: |  |  | Not applicable.

## TP
ACR Technique and Procedure Summaries
Image Mode (2D, 2D w/Add-on DBT, DBT)
| Facility | 0 | [6 empty] | Room ID | 0
|  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0 | Unit Mfr & Model
ACR DM Phantom Image Quality
Procedure: |  |  | Use the technique typically used for a 4.2 cm thick compressed breast of 50% glandular and 50% adipose tissue.
Obtain exposure mode, phantom setup, and operating levels from medical physicist.
Use largest image receptor size available and matching clinically used paddle.
Use 5 daN or 12 lbs compression force.
Score and analyze on acquisition workstation (radiologist workstation for DBT).
For DBT, record the slice or slab where the test objects are best visualized.
Adjust W/L to optimize test objects and record; do not subtract for artifacts.
Zoom and pan across entire image to evaluate for artifacts.
Phantom Setup |  | Date
AEC mode
Paddle size (IR size)
Paddle type (reg or flex)
View or selected image
Slice or slab # (DBT only)
Compression force
AEC cell position (if avail)
Target/filter (if app)
kVp (if app)
Density setting (if app)
Window width & window level (approx)
Image name (if app)
Image ID (if app)
SN-1234

## TPcont1
ACR Technique and Procedure Summaries (cont)
| Facility | 0 | [7 empty] | Room ID | 0
|  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0 | Unit Mfr & Model
AW & RW Monitor QC
Procedure: |  |  | Display the AAPM TG 18-QC or SMPTE test pattern.
(Another pattern that allows relevant measurements may also be used.)
Check the monitor screens to verify they are clean.
Verify that 0%-5% and 95%-100% boxes are visible (red circles).
Verify that line-pair images are sharp and distinguishable (red rectangles).
Score the ACR DM Phantom & evaluate for artifacts (RW only).
If applicable, perform monitor calibration tests per manufacturer's recommendations.
Monitor manufacturer instructions for finding & viewing TG 18-QC test pattern:
|  |  | SMPTE Test Pattern | [5 empty] | TG 18-QC Test Pattern

## TPcont2
ACR Technique and Procedure Summaries (cont)
| Facility | 0 | [9 empty] | Room ID | 0
[5 empty] | MAP ID-Unit# (00000-00) | 0 | - | 0 |  | Unit Mfr & Model
Film Printer QC
Procedure: |  |  |  | Required equipment - ACR DM Phantom image and densitometer
Print image acquired in "ACR DM Phantom Image Quality" test.
Print from workstation/computer/PACS used for printing most clinical films.
Do not adjust window and level settings prior to printing.
Evaluate the phantom for artifacts & score.
Note: If possible, print phantom image from same x-ray unit each time for this test.
Note: Use film size most commonly used for mammography.
Note: For each printer, only print a single image from a single x-ray unit or workstation.
Film Printer Setup | [7 empty] | Date:
Film Size (8 x 10 or 10 x 12):
ID of x-ray unit or workstation for printed phantom:
Dmax Note: Dmax to be measured on perimeter of film.
If not available, print a clinical image and measure
in the non-breast area.
|  |  | Background OD |  |  |  |  | Cavity OD |  |  | Dmax
SN-1234

## CA
Corrective Action Log
| Facility | 0 | [6 empty] | MAP ID# (00000-00) | 0 | - | 0
Room or Equipment ID | [7 empty] | Date
QC Test Name and # (if app):
Description:
Relevant Personnel Notified: |  |  | Personnel Name: | [5 empty] | Date/Time of Call/Notification:
(Radiologist, MP, tech, manager, service engineer)
Describe Actions Taken:
To Be Monitored
Confirmation of Resolution: |  |  |  |  | Yes |  | NA
|  |  |  | Event resolved? | [5 empty] | Tech Signature
|  |  |  | Documentation from service engineer obtained? | [5 empty] | Date

## Offsite Disp
Facility Offsite Display Locations
Facility | 0 |  | MAP ID# (00000) | 0
Address | 0
Offsite Locations or Facilities Where Images are Interpreted for this Facility (list facility name, address, and MAP ID)
Location or Facility Name |  | Address |  | MAP ID

## Unit Chklst Jan-Jun
Digital Mammography Unit QC Summary Checklist
Image Mode (2D, 2D w/Add-on DBT, DBT, All)
| Facility | 0 | [6 empty] | Room ID | 0
| MAP ID# (00000-00) | 0 | - | 0 |  |  |  | Unit Mfr & Model
Year
Month | Jan |  | Feb |  |  | Mar | Apr |  |  | May | Jun
ACR DM Phantom Image Quality (weekly)
CR Cassette Erasure, if app (weekly)
Compression Thickness Indicator (monthly)
Visual Checklist (monthly)
AW Monitor QC (monthly)
Compression (semiannual)
Mfr Detector Calibration, if app
Overall (only need to complete once for the facility)
Facility QC Review (quarterly)
Repeat Analysis (optional - as needed)
Detector Calibration Freq: |  |  |  |  | Date and initial each test: | date | [5 empty] | Cross out boxes where mfr calibration test is not required: | X
[6 empty] | initial | [6 empty] | X

## Unit Chklst Jul-Dec
Digital Mammography Unit QC Summary Checklist
Image Mode (2D, 2D w/Add-on DBT, DBT, All)
| Facility | 0 | [6 empty] | Room ID | 0
| MAP ID# (00000-00) | 0 | - | 0 |  |  |  | Unit Mfr & Model
Year
Month | Jul |  | Aug |  |  | Sep | Oct |  |  | Nov | Dec
ACR DM Phantom Image Quality (weekly)
CR Cassette Erasure, if app (weekly)
Compression Thickness Indicator (monthly)
Visual Checklist (monthly)
AW Monitor QC (monthly)
Compression (semiannual)
Mfr Detector Calibration, if app
Overall (only need to complete once for the facility)
Facility QC Review (quarterly)
Repeat Analysis (optional - as needed)
Detector Calibration Freq: |  |  |  |  | Date and initial each test: | date | [5 empty] | Cross out boxes where mfr calibration test is not required: | X
[6 empty] | initial | [6 empty] | X

## Disp Chklst
Facility Display Device QC Summary Checklist
Facility |  | MAP ID# (00000) | 0
Address
Address
QC Summary information for display devices at this MAP ID
Physical Location at Facility/ ID Designation
Device (RW, film printer, viewbox)
Manufacturer
Model
Jan | Date
Tech Initials
Feb | Date
Tech Initials
Mar | Date
Tech Initials
Apr | Date
Tech Initials
May | Date
Tech Initials
Jun | Date
Tech Initials
Jul | Date
Tech Initials
Aug | Date
Tech Initials
Sep | Date
Tech Initials
Oct | Date
Tech Initials
Nov | Date
Tech Initials
Dec | Date
Tech Initials
Medical Physicist Survey Date
Medical Physicist Name(s)

## DD
PF/PFNA | PASSFAIL
P | Pass
F | Fail
NA
TF
Mo/Mo
Mo/Rh
Rh/Al
Rh/Ag
Rh/Rh
W/Ag
W/Al
W/Rh
YN
Yes
No
NA
Check

Force
lbs
daN
lbs daN
DM-DBT | DM-DBT Only
2D | 2D
2D w/Add-on DBT | DBT
DBT
All
Device
RW
film printer
viewbox