## Instructions
Medical Physicist Test Form Spreadsheet
Instructions
• | IMPORTANT NOTE: This Excel Workbook contains some formulae to analyze the data and logic to report pass/fail results in order to help the medical physicist. It is intended to be used with both Windows and Mac operating systems. Please note, ANY alteration of the formulae or logic will result in errors. It is critical that you review the calculated results for accuracy and pass/fail designations for applicability on each sheet as well as the summary. It is the responsibility of the medical physicist to check all automated calculations and pass/fail results in the spreadsheet for correctness. Please notify the ACR if you find any errors.
• | Complete facility information in Facility Information tab first; information auto-populates in other pages.
• | Drop down lists are available within cells to complete the required information; if the list does not contain the necessary information; complete the cell yourself.
• | Your system must be tested using all clinically-used guidance modes (e.g., Stereotactic, Tomographic). See next tab for the required tests.
Make sure that the guidance mode that is being tested is identified in all applicable forms. Make additional copies of the forms for each guidance mode tested.
If you are using electronic forms, make appropriage copies of the forms (tabs) for all applicable guidance modes tested.
• | Enter data into white cells.
• | Light yellow cells are for pass/fail results of subtests within the main test; depending on the page, you may enter the pass/fail results using the dropdown list OR the results may populate via a formula.
• | Bright yellow cells are for pass/fail results of the overall test; depending on the page, you may enter the pass/fail results using the dropdown list OR the results may populate via a formula.
• | Do not enter data into gray cells. Some grey cells might contain formulae that will be lost.
Revisions
New

## TOC
Stereotactic Breast Biopsy
Medical Physicist's Quality Control Tests
Note: Complete Facility, Unit and Test Equipment Data tab first to populate facility information into forms
| Test | Minimum Frequency | Corrective Action Timeframe | No idea
1. | Phantom Image Quality | Acceptance, Annual, Post Repair, Troubleshooting | Before clinical use
2. | Artifacts | Acceptance, Annual, Post Repair, Troubleshooting | Before clinical use
3. | Spatial Resolution | Acceptance, Annual, Post Repair, Troubleshooting | Within 30 days of the survey
4. | DBT Volume Coverage | Acceptance, Annual, Post Repair, Troubleshooting | Before clinical use
5. | Automatic Exposure Control System Performance | Acceptance, Annual, Post Repair, Troubleshooting | Within 30 days of the survey
6. | Average Glandular Dose | Acceptance, Annual, Post Repair, Troubleshooting | Before clinical use
7. | Unit Checklist | Acceptance, Annual | Critical: before clinical use; Less Critical: within 30 days of the survey
8. | Localization Accuracy | Acceptance, Annual, Post Repair, Troubleshooting | Acceptance - before clinical use; Post Repair/Troubleshooting - within 30 days of the survey
9 | Acquisition Workstation (AW) Monitor Performance | Acceptance, Annual | Before clinical use
10 | Evaluation of Site's Technologist QC Program | Acceptance, Annual | Within 30 days of the survey
11 | Manufacturer Calibrations (if applicable) | Per Manufacturer's Requirements | Per Manufacturer's Requirements
12 | Collimation Accuracy | Acceptance, Post Repair, Troubleshooting | Within 30 days of the survey
13 | Compression Force | Acceptance, Post Repair, Troubleshooting | Acceptance - before clinical use; Post Repair/Troubleshooting - within 30 days of the survey
14 | kVp Accuracy and Reproducibility | Acceptance, Post Repair, Troubleshooting | Acceptance - before clinical use; Post Repair/Troubleshooting - within 30 days of the survey
15 | Beam Quality (Half-Value Layer) Assessment | Acceptance, Post Repair, Troubleshooting | Acceptance - before clinical use; Post Repair/Troubleshooting - within 30 days of the survey
16 | Ghost Image | Troubleshooting | Troubleshooting - within 30 days of the survey
Summary Report Forms
Medical Physicist's QC Test Summary
Mammography Technique Chart
Medical Physicist QC Letter for the Radiologist
Supplemental Forms
Facility, Unit and Test Equipment Data
QC Equipment List - Medical Physicist
| ACR Approved Phantom | Dosimeter
| (may use facility's phantom) | Lead sheet or equivalent
| 0.1 mm aluminum sheets | Coins, ready-pack film, electronic collimation test tools, or equivalent
2, 4, 6 and 8 cm thick PMMA, BR-12 or BR-50 sheets
kV meter
Note on D8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I added “of the survey” to clarify. Does not let MP or facility fudge if it is 30 days of the report. I modified all the other tests with 30 deadlines.
Note on D10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I added “of the survey” to clarify. Does not let MP or facility fudge if it is 30 days of the report. I modified all the other tests with 30 deadlines.
Note on D15: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I added “of the survey” to clarify. Does not let MP or facility fudge if it is 30 days of the report. I modified all the other tests with 30 deadlines.
Note on D17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I added “of the survey” to clarify. Does not let MP or facility fudge if it is 30 days of the report. I modified all the other tests with 30 deadlines.

## MP-Sum
Medical Physicist's QC Test Summary
Facility Name | 0 | [7 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
Address | 0 | [7 empty] | Room ID | 0
| 0 | [7 empty] | Report Date | 0
| 0 | [7 empty] | Survey Date | 0
|  | X-Ray Unit Manufacturer | 0 |  |  |  |  | Model | 0
|  | Control Panel Serial # | 0 |  |  |  |  | Manufacture Date | 0 |  |  | Installation Date | 0
Unit Type: |  |  | Stereotactic | Tomographic
Survey Type: |  |  | Initial Survey | Annual survey |  | Post-Repair Testing |  |  |  |  | Repair Desciption:
Oversight Level: |  |  | Medical physicist on-site |  |  | Medical physicist oversight
Medical Physicist: |  | 0 |  |  |  | Signature
Trainee(s): | [5 empty] | Signature
Signature
QC Test Results
Test |  |  |  |  | Pass/Fail* |  |  |  |  | CA
[5 empty] | Stereotactic |  | Tomographic
Medical Physicist Tests
1. Phantom Image Quality |  |  |  |  | #N/A |  | #N/A
2. Artifacts |  |  |  |  | 0 |  | 0
3. Spatial Resolution |  |  |  |  | 0 | [6 empty] | MP Phantom Results - Stereotactic
4. DBT Volume Coverage | [6 empty] | 0
5. Automatic Exposure Control System Performance |  |  |  |  | 0 |  | 0
6. Average Glandular Dose |  |  |  |  | #DIV/0! |  | #DIV/0! | [6 empty] | Fiber (≥ 2.0) | 0
7. Unit Checklist |  |  |  |  | 0 | [8 empty] | Speck grp (≥ 3.0) | 0
8. Localization Accuracy |  |  |  |  | 0 |  | 0 | [6 empty] | Mass (≥ 2.0) | 0
9. Acquisition Workstation Monitor Performance |  |  |  |  | 0 | [8 empty] | AGD (≤ 3.0 mGy) | #DIV/0!
10. Evaluation of Site's Technologist QC Program |  |  |  |  | 0
Acceptance / Post-Repair / Troubleshooting Only | [11 empty] | MP Phantom Results - Tomographic
11. Manufacturer Calibrations (if applicable) |  |  |  |  | 0 |  | 0
12. Collimation Assessment |  |  |  |  | 0
13. Compression Force |  |  |  |  | 0 | [8 empty] | Fiber (≥ 2.0) | 0
14. kVp Accuracy and Reproducibility |  |  |  |  | 0 | [8 empty] | Speck grp (≥ 3.0) | 0
15. Beam Quality (HVL) |  |  |  |  | 0 | [8 empty] | Mass (≥ 2.0) | 0
16. Ghost Image Evaluation |  |  |  |  | 0 | [8 empty] | AGD (≤ 3.0 mGy) | #DIV/0!
Technologist QC Evaluation
1. Localization Accuracy
2. Phantom Image Quality and Artifacts | [11 empty] | * "Pass" means all components of test passes; "Fail" means any or all components fail; if "CA" checked, see Corrective Action Summary
3. Visual Checklist
4. Compression Thickness Indicator
5. Acquisition Workstation (AW) Monitor QC
6. Compression Force
7. Facility QC Review
8. Manufacturer Calibration (if applicable)
9. Physician Feedback (optional)
10. Repeat Analysis (optional)

## Facility Information
Facility, Unit and Test Equipment Data
Medical Physicist's Tests
Facility Information
Facility Name
Address
Address
City, State, Zip
SBBAP ID# (00000)
SBBAP Unit# (00)
Lead Physician
Quality Control Technologist
Survey Information
Room ID
Survey Date
MP Report Date
Date of Previous Survey
X-ray Guided Breast Biopsy Unit Information
X-Ray Unit | Manufacturer
Model
Gantry Serial #
Date of Manufacture
Date of Installation
Biospy Attachment (if applicable) | Model
Serial #
Test Equipment Info
Phantom | Manufacturer
SN
Type
Dosimeter | Manufacturer
Model
Calibration Date
Dosimeter | Manufacturer
Model
Calibration Date
Photometer | Manufacturer
Model
Calibration Date
Medical Physicist Info
Medical Physicist | Name
Telephone Number
Email
Signature

## 1. Phantom IQ
1. Phantom Image Quality
| Facility Name | 0 | [8 empty] | SBBAP ID-Unit# | 0 |  | - | 0
| Mfr & Model | [9 empty] | Room ID | 0
[11 empty] | Survey Date | 0
Procedure |  | Equipment: ACR Phantom (Orginal, DM, or Mini) | [9 empty] | Pantom Type: | 0
|  | • Use clinical technique for typical screening exam of 4.2 cm 50/50 breast | [9 empty] | Phantom Model: | 0
|  | • Adjust W/L to optimize test objects, zoom & pan entire image | [9 empty] | Phantom SN: | 0
[5 empty] | Stereotactic: | WL | WW
[5 empty] | DBT: | WL | WW |  |  |  |  | Phantom Info (PACS)
|  | • For DBT, scroll to best slice/slab to visualize test objects | [9 empty] | Name (Last, First):
[5 empty] | Slice/slab #: | [6 empty] | ID: | phantom
[17 empty] | Stereotactic | DBT
Image Mode (Stereotactic, DBT) | [8 empty] | Stereotactic | DBT | [6 empty] | Mo/Mo | Mo/Mo
Resulting Techniques |  | View | [6 empty] | CC (Scout) | CC (Scout) | [6 empty] | Mo/Rh | Mo/Rh
|  | Paddle | [14 empty] | Rh/Ag | Rh/Ag
[8 empty] | AEC Mode | [8 empty] | Rh/Rh | Rh/Rh
|  | AEC Position | [14 empty] | W/Ag | W/Al
|  | Exposure Compensation Step | [14 empty] | W/Al | W/Rh
|  | Target/Filter | [14 empty] | W/Rh
Tube Voltage - kV
[8 empty] | Tube current exposure time product - mAs | [9 empty] | Object
|  | Unit-indicated AGD (mGy) | [13 empty] | Min Score | Type | Fibers | Specks | Masses
Phantom Evaluation |  | Fiber score | [13 empty] | #N/A | ACR DM | 2 | 3 | 2
|  | Speck group score | [13 empty] | #N/A | ACR Mini | 2 | 2 | 2
|  | Mass score | [13 empty] | #N/A | ACR Original | 4 | 3 | 3
Overall Pass/Fail | [8 empty] | #N/A | #N/A
Comments:
Minimum Scores: | [5 empty] | Original ACR Phantom |  |  | Mini Phantom |  | ACR DM Phantom
|  |  | Fibers |  |  | 4 |  |  | 2 |  | 2
|  |  | Speck Groups |  |  | 3 |  |  | 2 |  | 3
|  |  | Masses |  |  | 3 |  |  | 2 |  | 2
Action Limits |  | Required: |  | The phantom image must score at least the minimum score (above), appropriate to the phantom used.
|  | Timeframe: |  | Failure must be corrected before clinical use.
Note for upright units with biopsy attachments: For upright units with biopsy attachments, evaluations made during other mammography unit QC may not be appropriate for this test because the AEC settings and detector calibrations may be different. In this case, phantom image quality needs to be repeated using clinical biopsy procedure settings.

## 2. Artifacts
2. Artifacts
| Facility Name | 0 | [5 empty] | SBBAP ID-Unit# | 0 |  | - | 0
| Mfr & Model | [6 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Phantom Setup |  | Equipment: Apprxoimately 4 cm of a uniform tissue-simulating attenuator (e.g., PMMA, BR-12, ACR DM phantom)
Attenuator:
• Use clinical technique for a breast of the thickness of the absorber used.
• Adjust W/L to optimize the visualizaton of artifacts.
• For DBT, scroll through all slices and evaluate each slice/slab for artifacts.
Image Mode (Stereotactic, DBT) | [5 empty] | Stereotactic | DBT | [5 empty] | Stereotactic | DBT
Techniques |  | Target/Filter |  |  |  | W/Rh | W/Al | [5 empty] | Mo/Mo | Mo/Mo
[5 empty] | Tube Voltage - kV | [7 empty] | Mo/Rh | Mo/Rh
[5 empty] | Tube current exposure time product - mAs | [7 empty] | Rh/Ag | Rh/Ag
[5 empty] | Pass/Fail: | Pass | [6 empty] | Rh/Rh | Rh/Rh
[13 empty] | W/Ag | W/Al
[8 empty] | Overall Pass/Fail |  | Stereotactic |  |  | W/Al | W/Rh
[10 empty] | DBT |  |  | W/Rh
Comments:
Action Limits |  | Required: | Phantom image must be free of clinically significant artifacts.
|  | Timeframe: | Failure must be corrected before clinical use.
Note for upright units with biopsy attachments: For upright units with biopsy attachments, evaluations made during other mammography unit QC may not be appropriate for this test because the detector calibrations may be different. In this case, artifact evaluation needs to be repeated.

## 3. Spatial Resolution
3. Spatial Resolution
|  | Facility Name | 0 | [6 empty] | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model | [7 empty] | Room ID | 0
[10 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  | A line-pair resolution test tool, and 4 cm of a uniform tissue-simulating attenuator (e.g., PMMA, BR-12, ACR DM phantom)
Attenuator:
Place attenuator material on the breast support.
Place bar pattern on top of attenuator and under paddle at ~45o and 2 cm from the chest wall.
Lightly compress paddle to touch bar pattern.
Acquire "raw" images using manual mode closest to Phantom technique
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
| Image Mode | [6 empty] | Stereotactic
| Setup Technique |  |  | Target/filter
kVp
mAs
Line-pair score (lp/mm)
Pass/Fail
[9 empty] | Overall Pass/Fail |  |  | Stereotactic
Comments:
Action Limits |  |  | Required: |  | For 2D, spatial resolution must be ≥ 4.0 lp/mm
|  |  | Timeframe: |  | Failures must be corrected within 30 days. For acceptance testing, failures must be corrected before clinical use.
Note on AC7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Removed paddle type (reg or flex), as I don’t' think that it applies.

## 4. DBT Volume Coverage
4. DBT Volume Coverage
|  | Facility Name | 0 |  |  |  |  | SBBAP ID-Unit# | 0 | - | 0
|  | Mfr & Model | [5 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  | Tissue-simulating attenuator (4 cm PMMA or BR-12, ACR DM phantom), 2 sheets of 0.1 mm Al
Place attenuator material on the breast support.
Place aluminum (Al) sheets on top and bottom of the attenuator, diagonally across chest wall.
Acquire DBT image of phantom.
View reconstructed images and verify that both Al sheets are appropriately in focus within the volume.
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
| Setup |  |  |  |  | Compression Paddle
Attenuator
| Results (Yes/No) |  |  |  |  | Upper Al sheet in focus in a slice within the volume
Lower Al sheet in focus in a slice within the volume
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: | Both Al sheets must be seen in focus in volume.
|  |  | Timeframe: | Failures must be corrected before clinical use.

## 5. AEC (Stereo)
5. Automatic Exposure Control System Performance
|  | Facility Name | 0 | [12 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
|  | Mfr & Model | [13 empty] | Room ID | 0
[16 empty] | Survey Date | 0
Procedure |  |  | Equipment: | 2, 4, 6, 8 cm of BR-12, BR-50 or PMMA | [9 empty] | Setup: |  |  |  | Paddle size (IR Size):
[5 empty] | Attenuator: | [12 empty] | AEC cell position:
|  |  | Set thickness at actual thickness of material (2, 4, 6, or 8 cm) | [14 empty] | Mfr DC offset: | 50
|  |  | or 5 decanewtons or 12 lbs of compression force | [14 empty] | Other settings:
Acquire images using clinical techniques
[18 empty] | SNR Baselines - MEE Date: | 2025-10-11
AEC Thickness Tracking | [12 empty] | Image Mode | Stereotactic
Mode | Thick - ness (cm) |  | Setup Techniques |  |  | Resultant Techniques | [6 empty] | Signal and Noise Measurements
|  |  | AEC Mode | Density setting |  | Target/ Filter | kVp |  | mAs |  | EI/AGD/Other (optional) |  | Mean Bkgd Signal |  | Std Dev of Bkgd |  | DC Offset (if app) |  |  | SNR
Contact | 2 | [15 empty] | 50 |  |  | #DIV/0!
Contact | 4 | [15 empty] | 50 |  |  | #DIV/0!
Contact | 6 | [15 empty] | 50 |  |  | #DIV/0!
Contact | 8 | [15 empty] | 50 |  |  | #DIV/0!
Analysis
Mode | Thick - ness (cm) |  | SNR |  | Annual
[5 empty] | Baseline SNR |  |  | Lower Limit |  | Upper Limit |  | SNR within ±15% of Initial (P/F)
Contact | 2 |  | #DIV/0! |  |  |  |  | 0 |  | 0 |  | #DIV/0!
Contact | 4 |  | #DIV/0! |  |  |  |  | 0 |  | 0 |  | #DIV/0!
Contact | 6 |  | #DIV/0! |  |  |  |  | 0 |  | 0 |  | #DIV/0!
Contact | 8 |  | #DIV/0! |  |  |  |  | 0 |  | 0 |  | #DIV/0!
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: | Annual: SNR must be within ±15% of Baseline over the clinically used phantom thickness and imaging modes.
|  |  | Timeframe: | Failures must be corrected within 30 days; for Initial, before clinical use.

## 5. AEC (Tomo)
5. Automatic Exposure Control System Performance
|  | Facility Name | 0 | [13 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
|  | Mfr & Model | [14 empty] | Room ID | 0
[17 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  | 2, 4, 6, 8 cm of BR-12, BR-50 or PMMA | [9 empty] | Setup: |  |  |  | Paddle size (IR Size):
[6 empty] | Attenuator: | [12 empty] | AEC cell position:
|  |  | Set thickness at actual thickness of material (2, 4, 6, or 8 cm) | [15 empty] | Mfr DC offset: | 50
|  |  |  | or 5 decanewtons or 12 lbs of compression force | [14 empty] | Other settings:
Acquire images using clinical techniques
[19 empty] | SNR Baselines - MEE Date: | 2025-10-11
AEC Thickness Tracking | [13 empty] | Image Mode | DBT
Mode | Thick - ness (cm) |  | Setup Techniques |  |  |  | Resultant Techniques | [6 empty] | Signal and Noise Measurements
|  |  | AEC Mode |  | Density setting |  | Target/ Filter | kVp |  | mAs |  | EI/AGD/Other (optional) |  | Mean Bkgd Signal |  | Std Dev of Bkgd |  | DC Offset (if app) |  |  | SNR
Contact | 2 | [16 empty] | 50 |  |  | #DIV/0!
Contact | 4 | [16 empty] | 50 |  |  | #DIV/0!
Contact | 6 | [16 empty] | 50 |  |  | #DIV/0!
Contact | 8 | [16 empty] | 50 |  |  | #DIV/0!
Analysis
Mode | Thick - ness (cm) |  | SNR |  |  | Annual
[6 empty] | Baseline SNR |  |  | Lower Limit |  | Upper Limit |  | SNR within ±15% of Initial (P/F)
Contact | 2 |  | #DIV/0! | [5 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 4 |  | #DIV/0! | [5 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 6 |  | #DIV/0! | [5 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 8 |  | #DIV/0! | [5 empty] | 0 |  | 0 |  | #DIV/0!
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: |  | Annual: SNR must be within ±15% of Baseline over the clinically used phantom thickness and imaging modes.
|  |  | Timeframe: |  | Failures must be corrected within 30 days; for Initial, before clinical use.

## 6. AGD
6. Average Glandular Dose
| Facility Name | 0 |  |  |  | SBBAP ID-Unit# | 0 |  |  | - | 0
| Mfr & Model |  |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
Procedure |  | Equipment: | Dosimeter |  |  |  | Dosimetry system:
|  | • Use manual technique closest to Test 2 phantom technique. |  |  |  |  | Calibration date:
|  | • Protect image receptor. |  |  |  |  | Correction factor, if app:
|  | • Place dosimeter on breast support, 5cm from chest wall edge. | [8 empty] | Source-to-Object Table Distance lt - (mm): | 648 | <----- | NOTE: Units in mm
|  | • Measure air kerma at measurement point. | [8 empty] | Source-to-Measurement Distance lm - (mm): | 648
|  | • Correct for mAs and distance to obtain reference air kerma. | [8 empty] | DBT Acquisition Geometry (# Projections / Angular Range): | 25 projections, 50 degrees | <----- | DBT Geometry
[14 empty] | 15 projections, 15 degrees | 15_15
Imaging Mode (Stereotactic, DBT) | [8 empty] | Stereotactic |  | DBT |  |  | 25 projections, 50 degrees | 25_50
[14 empty] | 30 projections, 15 degrees | 30_15
Technique Factors Resulting From ACR Phantom Acquisition | [7 empty] | View | CC (Scout) |  | CC (Scout) |  |  | 9 projections, 25 degrees | 09_25
[8 empty] | Paddle | 0 |  | 0
[8 empty] | AEC mode | 0 |  | 0
[8 empty] | Target/Filter | 0 |  | 0
[8 empty] | Tube Potential (kVp) | 0 |  | 0
[8 empty] | Tube current - exposure time product - mAsphantom | 0 |  | 0
Air Kerma at Measurement Point | [7 empty] | Measured HVL (mm Al)
mAs for manual air kerma measurement - mAsmanual
Measurement #1 (mGy)
Measurement #2 (mGy)
Measurement #3 (mGy)
[8 empty] | Average Air Kerma for Manual Technique - Km,manual (mGy) | #DIV/0! |  | #DIV/0!
[8 empty] | Air Kerma for Phantom Technique - Km,phantom (mGy) | #DIV/0! |  | #DIV/0!
Reference AK | [7 empty] | Reference Air Kerma - Kref (mGy) | #DIV/0! |  | #DIV/0!
AGD Calculations Dg = Kref Γ |  |  |  | AGD to a "standard breast" as defined by the FDA (4.2 cm, 50% glandularity by mass or 35% VBD)
[8 empty] | Conversion Coefficient under Monte Carlo Geometry - Γsim | #VALUE! |  | #VALUE! |  | <-- densityMode = "FDA" for function GET_GAMMA_DM and/or GET_GAMMA_DBT
[8 empty] | Conversion Coefficient under System Geometry - Γ | #VALUE! |  | #VALUE!
[8 empty] | Dg,FDA (mGy) | #DIV/0! |  | #DIV/0!
[8 empty] | Overall Pass/Fail | #DIV/0! |  | #DIV/0!
AGD to a breast with average glandularity (50th percentile %VBD for 4.2cm) - Informational Only
[8 empty] | Conversion Coefficient under Monte Carlo Geometry - Γsim | #VALUE! |  | #VALUE! |  | <-- densityMode = "US50" for function GET_GAMMA_DM and/or GET_GAMMA_DBT
[8 empty] | Conversion Coefficient under System Geometry - Γ | #VALUE! |  | #VALUE!
[8 empty] | Dg,average (mGy) | #DIV/0! |  | #DIV/0!
Comments:
Action Limits |  | Required: |  | AGD for a single cranio-caudal scout view of the phantom in either Stereotactic or Tomographic mode must not exceed 3.0 mGy for a "standard breast" (as defined by the FDA).
|  | Timeframe: |  | Doses > 3 mGy must be corrected before clinical use;
"GET_GAMMA" Function Inputs
[7 empty] | mode |  | DM |  | DBT
[7 empty] | targetFilter |  | 0 |  | 0
[7 empty] | HVL |  | 0 |  | 0
[7 empty] | kVp |  | 0 |  | 0
[7 empty] | proj_ang |  |  |  | 25_50
Γsim- FDA --> "standard breast" as defined by FDA
[7 empty] | denistyMode |  | FDA |  | FDA
[7 empty] | Γsim |  | #VALUE! |  | #VALUE!
Γsim- Typical Breast --> Glandularity set to 50th percentile %VBD for 4.2 cm
[7 empty] | denistyMode |  | US50 |  | US50
[7 empty] | Γsim |  | #VALUE! |  | #VALUE!
To lookup Γsim for stereo/DM accuiqisitions, use function GET_GAMMA_DM:
= GET_GAMMA_DM(densityMode,TargetFilter,HVL,kVp)
To lookup Γsim for DBT acquisitions, use function GET_GAMMA_DBT:
= GET_GAMMA_DBT(densityMode,TargetFilter,ProjAngle,HVL,kVp)
Note:
Cells calculating gamma coefficients are not protected. User is cautioned not to override these cells.
Input variables:
DensityMode = FDA, US50
TargetFilter (DM): MoMo, MoRh, RhAg, RhRh, WAg, WAl, WRh
TargetFilter (DBT): MoMo, MoRh, Rh,Ag, RhRh, WAl, WRh
HVL: number value in mm
kVp: integer value between 20-49
proj_ang (DBT): format #Projections_Angular Range
15_15, 25_50, 30_15, 09_25

## 7. Unit Checklist
7. Unit Checklist
|  | Facility Name | 0 |  |  |  |  | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model | [5 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  |  | Equipment: | None
Inspect the unit and evaluate the functionality according to the checklist below.
Critical items (*) are in bold typeface.
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
Item | [10 empty] | Pass/Fail/NA
1. | Unit is mechanically stable under normal operating conditions.*
2. | All moving parts move smoothly without obstructions to motion.*
3. | All locks and detents work properly.*
4. | Image receptor assembly and compression paddle is free from vibrations during normal operation.*
5. | Patient or operator is not exposed to sharp or rough edges, or other hazards.*
6. | Paddles are all intact with no cracks or sharp edges.*
7. | All indicators are working properly.
8. | Indicated breast thickness accurate to 5 mm over thicknesses of 2 to 8 cm, reproducible to 2 mm.
[6 empty] | Actual Thickness (cm): | Measured Thickness (cm): | [5 empty] | Accuracy | Reproducibility
[6 empty] | 2 | [6 empty] | 2
[6 empty] | 4 | [6 empty] | 4 | 0
[6 empty] | 6 | [6 empty] | 6
[6 empty] | 8 | [6 empty] | 8
9. | Needle holder and guides are firmly attached and needle is straight to within 1 mm.* | [11 empty] | FAIL | PASS
10. | Compression can be manually released with absence of power.*
11. | Clinical imaging area is clean and free from significant dust and debris that may cause artifacts.
12. | Operator and other personnel are protected during exposure by adequate radiation shielding.*
13. | Technique Charts posted and acceptable.
14. | Audible exposure indicator at an appropriate volume level.
15. | Verify that the DBT assembly moves as designed through its range of motion. (If applicable)*
16. | Unit is clean and free of blood or other fluids*
17. | Other:
18. | Other:
19. | Other:
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: |  | All items must pass. Critical items are marked with an *.
|  |  | Timeframe: |  | Failures of critical items (*) must be corrected before clinical use; less critical items must be corrected within 30 days.
Note on E10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I added “diagnostic” to clarify

## 8. Localization
8. Localization Accuracy
| Facility Name | 0 | [5 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
| Mfr & Model | [6 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  |  | Equipment: Targeting phantom with appropriately small (<5mm diameter) targets, a core biopsy system, a core biopsy
needle (a clinically used type and gauge) or alternative tissue sampling device, and a Manufacturer’s supplied or
self-manufactured phantom localization test device.
Localization in a Phantom: Test can be completed with a clincal case review or a core biopsy procedure using a
targeting phantom. Evaluate the most common clinically used mode and approach.
Localization in Air: Test can be completed with a Manufacturer's supplied localization test device or a self-manufactured
phantom localization test device and localization procedure. Evaluate all other clinically used mode and approach combinations.
Refer to section II.A.8. in the Medical Physics Section of the Quality Control Manual for detailed procedure instructions.
Localization in a Phantom or Medium
Most used clinical mode:
Most used clinical approach:
Technologist (if applicable):
Verification Method:
The test results can be considered passing for localization in a target if any of the following is achieved: | [9 empty] | Criteria Met
The needle tip is within the target in pre-fire images and beyond the target in the post-fire images with the trough area within the target.
The target is visible in the specimen radiograph.
The visual inspection of the needle trough confirms the target was captured.
Pass/Fail
Localization in Air
Mode |  |  |  |  | Methodology for testing | Were the calculated coordinates within +/-1 mm of target |  |  |  |  | Pass / Fail / N/A
Stereotactic |  | CC
Stereotactic |  | Lateral Left
Stereotactic |  | Lateral Right
DBT |  | CC
DBT |  | Lateral Left
DBT |  | Lateral Right
[6 empty] | Overall Pass/Fail | Stereotactic
DBT
Comments:
Action Limits |  | Required: |  | If passing criteria for localization in a phantom and/or localization in air are not met, the reasons for these errors should be determined and the procedure repeated. If localizations continue to be inaccurate, a service engineer should be contacted by facility personnel and the problem corrected.
|  | Timeframe: |  | Failures must be corrected before clinical use.

## 9. AW
9. Acquisition Workstation Monitor Performance
|  | Facility Name | 0 | [5 empty] | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model | [6 empty] | Room ID | 0
[9 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  | Luminance meter
|  |  | Note: Some of these QC tests may or may not be possible to perform depending on the monitor QC capabilities | [8 empty] | Pattern Used:
|  |  | Test Pattern Image Quality: Use TG270-sQC, TG18-OIQ, TG18-QC, SMPTE or other relevant pattern (if available) | [8 empty] | TG-270 QC
Luminance Check: TG270-sQC or TG18 LN8-01 & LN8-18 test patterns, or others that provide Lmin & Lmax (if available)
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
|  |  |  | Monitor manufacturer: |  |  | Model: |  |  |  | Significant findings indicated on figure below
Monitor serial number
Monitor date of manufacture
Monitor Condition: Significant findings P/F
Test Pattern Image Quality: (Pattern) |  |  |  | Test pattern centered appropriately?
0%-5% contrast boxes visible?
95%-100% contrast boxes visible?
Alphanumerics sharp and legible?
3 "Quality Control" patches visible?
Line-pair images distinct (center)?
Line-pair images distinct (corners)?
Grayscale ramps smooth ?
|  |  |  | Test pattern P/F |  |  |  | Pass |  |  | Luminance Uniformity
Luminance Check (if available) |  |  |  | Estimated Ambient Luminance Lamb (cd/m2) | [6 empty] | Center
|  |  |  | Measured Luminance minimum - Lmin (cd/m2) | [6 empty] | Upper Left
|  |  |  | Measured Luminance maximum - Lmax (cd/m2) | [6 empty] | Upper Right
| Mfr recommendation for Lmax or Baseline Lmax | [9 empty] | Lower Left
|  | Lmax within ±10% of mfr recommendation or baseline? | [5 empty] | #DIV/0! |  |  | Lower Right
|  |  |  | Luminance Ratio: L'max/L'min |  |  |  | #DIV/0! |  |  | Max |  | 0
|  |  |  | Luminance Ratio ≥ 350 |  |  |  | #DIV/0! |  |  | Min |  | 0
|  |  |  | Luminance check P/F | [6 empty] | % Diff |  | #DIV/0!
DICOM GSDF |  |  |  | W/in ±10% of targeted contrast response P/F (if avail) | [6 empty] | P/F |  | #DIV/0!
Mfr Automated Test |  |  |  | Most recent set of mfr automated tests P/F
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: |  |  | Any identified screen blemish that could interfere with clinical information must be removed.
Test pattern image quality must pass all visual tests.
Lmax must be within ±10% of mfr specification or baseline established by physicist.
Luminance uniformity must be ≤ 30%
Luminance ratio must be LR’ ≥ 350.
GSDF measured contrast response must be within ±10% of targeted contrast response.
Mfr's automated tests must pass mfr specifications (if 1 test fails, indicate "F").
|  |  | Timeframe: |  |  | Significant monitor cleanliness defects must be corrected before clinical use; all other required tests must be corrected within 30 days.

## 10. Tech QC
10. Evaluation of Technologist's QC Program
|  | Facility Name | 0 | [8 empty] | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model | [9 empty] | Room ID | 0
[12 empty] | Survey Date | 0
Deficiencies Noted
Radiologic Technologist's Quality Control Tests | [5 empty] | Frequency | Review of Test by MP | Test Performed, Analyzed & Documented Incorrectly | Missing Data | Incorrect Scoring or Calculations | Missing Corrective Action Documentation | Other |  | Comments
1. | Localization Accuracy |  |  |  |  | Daily
Monthly
2. | Phantom Image Quality |  |  |  |  | Weekly
3. | Visual Checklist |  |  |  |  | Monthly
4. | Compression Thickness Indicator |  |  |  |  | Monthly
5. | AW Monitor QC |  |  |  |  | Monthly
6. | Compression Force |  |  |  |  | Semiannually
7. | Repeat Analysis (Optional) |  |  |  |  | Semiannually
8. | Facility QC Review |  |  |  |  | Annually
9. | Mfr Calibrations (if applicable)
10. | Physician Feedback (Optional) |  |  |  |  | As Needed
Corrective Action Log documentation adequate?
Overall Pass/Fail for Performance of Technologist QC Program
Additional Comments:
Action Limits |  |  | Required: |  | In order for the overall evaluation to pass, there must be a) no significant missing data, b) the tests must be analyzed without gross errors, and c) appropriate corrective action for failures must be taken (and documented). See test procedures for more information.
|  |  | Timeframe: |  | Failures must be corrected within 30 days.
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Reviewed: 1/3/2025

## 11. Man Cal
11. Manufacturer Calibrations (if applicable)
|  | Facility Name | 0 |  |  |  |  | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model | [5 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  |  | Follow manufacturer's instructions.
|  |  | Notes: | Acceptance Testing: Perform required manufacturers calibrations/tests. Annual testing: verify that manufacturer
calibrations have been performed at the required frequency.
| Image Mode (Stereotactic, DBT) | [5 empty] | Stereotactic |  |  | DBT
Name of Calibration
Overall Pass/Fail
Action Limits |  |  | Required: |  | Unit must pass all manufacturer's calibrations to pass overall.
|  |  | Timeframe: |  | Failures must be corrected per manufacturer's instructions.

## 12. Collimation
12. Collimation Assessment
| Facility Name | 0 |  |  |  | SBBAP ID-Unit# | 0 |  | - | 0
| Mfr & Model |  |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
Procedure |  | Equipment: | Coins, film, electronic collimation test tool(s), etc.
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
Technique |  |  |  | Target material
Paddle
Collimator size (cm)
|  |  |  | SID (mm) | #REF! |  |  | #REF!
Deviation Between X-ray Field and Edges of the Image Receptor |  |  |  | Left edge deviation
|  |  |  | % of SID (retain sign) | #REF! |  |  | #REF!
Right edge deviation
|  |  |  | % of SID (retain sign) | #REF! |  |  | #REF!
Anterior edge deviation
|  |  |  | % of SID (retain sign) | #REF! |  |  | #REF!
Chest edge deviation
|  |  |  | % of SID (retain sign) | #REF! |  |  | #REF!
Pass/Fail |  |  |  |  | #REF! |  |  | #REF!
Overall Pass/Fail
Comments:
Action Limits |  | Required: | If X-ray field exceeds image receptor at any side by more than +2% of SID or if X-ray field falls within image receptor on any side, seek service adjustment.
|  | Timeframe: | Failures must be corrected within 30 days; for Initial, before clinical use.

## 13. Comp Force
13. Compression Force
| Facility Name | 0 |  |  |  | SBBAP ID-Unit# | 00000 - 00
| Mfr & Model |  |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
Procedure |  | Required Equipment: Bathroom scale; towels Place towel on detector, place bathroom scale on towel with dial or read-out positioned for reading. Place another towel on top of scale. Using initial power-drive mode, activate compression until it stops automatically. Read and record the initial compression force. Continue to compress using power-drive mode and record maximum compression force achieved. Using manual fine-adjustment mode, activate compression until maximum compression is reached. Read and record the manual compression force. Set the compression force to a clincally used force, approximately 10 lbs Check that force is maintained within 1lb for approximately 2 minutes. Note: Ensure that the scale that is used produces accurate readings as pressure is increased.
Date: |  |  |  |  | Measured Force |  |  | Units
Initial power-drive compression force
Maximum power-drive compression force
Power-drive force is at least 10 lbs (4.4 daN) but no more than 45 lbs (20.0 daN) Pass/Fail
Maximum manual fine-adjustment compression force
Manual compression force is at least 10 lb (4.4 daN)
Compression Force maintains at least 10 lbs (4.4 daN) for at least 2 minutes Pass/Fail
Overall Pass/Fail
Comments:
Action Limits |  | Required: | The compression force using the power mode must be at least 10 pounds, sufficient to stabilize the breast before final positioning, and must not exceed 45 pounds. The unit must be able to hold a compression of at least 10 pounds for at least 2 minutes and should not vary by more than one pound. The manually applied compression force must be at least 10 pounds but may exceed 45 pounds.
|  | Timeframe: | Initial Survey: Before clinical use; Post Repair/Troubleshooting: Before clinical use.

## 14 kVp
14. kVp Accuracy and Reproducibility
|  | Facility Name | 0 | [6 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
|  | Mfr & Model | [7 empty] | Room ID | 0
[10 empty] | Survey Date | 0
Procedure |  |  | Equipment: | kVp meter, protective device
Cover the entire detector with a steel plate, protective apron, or other protective device.
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
|  |  |  | kVp meter | 0 |  |  | Setting
|  |  |  | Calibration Date: | 0
[6 empty] | Low Clinical kVp* | Phantom Clinical kVp (Primary Clinical Mode) |  | Phantom Clinical kVp (Other Clinical Mode) |  |  |  | High Clinical kVp*
Mode
Technique | Nominal kVp setting
Target/filter
Focal spot
mAs
Data | Measured kVp value 1
| Measured kVp value 2 |  |  |  |  | -- |  |  | -- |  |  |  | --
| Measured kVp value 3 |  |  |  |  | -- |  |  | -- |  |  |  | --
Analysis | Mean kVp |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0!
| Standard deviation (SD) |  |  |  |  | -- | #DIV/0! |  | -- |  |  |  | --
| Mean kVp - nominal kVp |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0!
| 0.05 x nominal kVp |  |  |  |  | 0 | 0 |  | 0 |  |  |  | 0
| % error |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0!
| % error P/F |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0!
| Coefficient of variation (CV) |  |  |  |  | -- | #DIV/0! |  | -- |  |  |  | --
| CV P/F |  |  |  |  | -- | #DIV/0! |  | -- |  |  |  | --
Pass/Fail | [5 empty] | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0!
* The low and high kVps may be either those used for 2D or DBT imaging, whichever are lowest or highest
Overall Pass/Fail
Comments:
Action Limits |  |  | Required: | Mean kVp must not differ from the nominal by more than ±5% of the nominal kVp.
Coefficient of variation must be ≤ 0.02.
|  |  | Timeframe: | Failures must be corrected within 30 days; for Initial Survey, before clinical use.

## 15. HVL
15. Beam Quality (Half-Value Layer)
| Facility Name | 0 | [12 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
| Mfr & Model | [13 empty] | Room ID | 0
[15 empty] | Survey Date | 0
Procedure |  | Equipment: |  |  | Dosimeter, 0.1 mm Al sheets, protective device | [10 empty] | Dosimetry system: | 0
|  | Cover the entire detector with a steel plate, protective apron, or other protective device. | [13 empty] | Cal.date: | 0
Make at least 1 measurement for each available target-filter combination used.
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
Imaging Mode (Stereotactic, DBT) | [5 empty] | Stereotactic |  | DBT
[6 empty] | Clinical |  | Clinical |  |  | Target/Filter 3 |  |  | Target/Filter 4 |  |  | Target/Filter 5 |  |  | Target/Filter 6
Target/filter
Nominal kVp setting
mAs
Paddle
Air Kerma Measurements
[6 empty] | mm AL | K (mGy) | mm AL |  | K (mGy) | mm AL | K (mGy) |  | mm AL | K (mGy) |  | mm AL | K (mGy) |  | mm AL | K (mGy)
No Aluminum |  |  |  | E0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Al Thickness (mm) ta |  |  |  | Ea
Al Thickness (mm) tb |  |  |  | Eb
Calculated or measured HVL (mm Al) | [5 empty] | #DIV/0! |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0!
Minimum allowed HVL | [5 empty] | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Overall Pass/Fail | [5 empty] | #DIV/0! |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0!
[8 empty] | HVL = | tb ln[2Ea/E0] - ta ln[2Eb/E0]
ln [Ea/Eb]
Overall Pass/Fail
Comments:
Action Limits |  | Required: |  |  | The HVL must meet the specifications of FDA's Performance Standards for Ionizing Radiation Emitting Products (Part 1020.30) as shown below.
|  | Timeframe: |  |  | Failures must be corrected within 30 days; for Initial Survey, before clinical use.
FDA X-ray Tube Voltage (kilovolt peak) and Minimum HVL
|  |  | Designed Operating Range (kV) | [5 empty] | Measured Operating Voltage (kV) |  |  |  | Minimum HVL (mm of Al)
|  |  | Below 50 | [5 empty] | 20 |  |  |  | 0.2
[9 empty] | 25 |  |  |  | 0.25
[9 empty] | 30 |  |  |  | 0.3
Imaging Mode (Stereotactic, DBT)
Note for upright units with biopsy attachments: For upright units with biopsy attachments, evaluations made during other mammography unit QC may not be appropriate for this test because the paddle material is in the biopsy field during the measurement. In this case, HVL measurements need to be repeated using the compression paddle used for clinical biopsy procedures, which is typically an open-window paddle.

## 17. (hidden)
17. Ghost Image
|  | Facility Name | 0 |  | SBBAP ID-Unit# | 0 | - | 0
|  | Mfr & Model |  |  | Room ID | 0
Procedure
Results (P/F)
Overall Pass/Fail
Action Limits |  |  | Required:
Timeframe:

## 16. Ghost
16. Ghost Image
| Facility Name | 0 |  |  |  |  | SBBAP ID-Unit# | 0 |  |  |  | - | 0
| Mfr & Model | [5 empty] | Room ID | 0
[7 empty] | Survey Date | 0
Procedure |  | Equipment: |  | ACR DM Phantom, 0.1 mm Al sheet (10 cm x 10 cm) | [6 empty] | Phantom Setup
|  | Largest image receptor size | [8 empty] | Paddle Size (IR size):
|  | Apply 5 daN or 12 lbs comp force | [8 empty] | Exposure Mode:
Compression Force:
|  | Position ACR DM Phantom with wax insert opposite from chest wall edge. | [8 empty] | AEC Cell Position (if avail):
|  | AEC cell position to position "3". | [8 empty] | Density Setting:
Use clinical (AEC) technique for both images
Image 1: Position phantom like Setup Image #1 (edge extends 1" beyond midline).
Image 2: Position phantom like Setup Image #2 with Al placed on top.
Signal data must be obtained from raw image
[6 empty] | Image 1 |  |  |  | Image 2
Resulting Techniques from Image Acquisition |  |  |  | Target/filter
kVp
mAs
Ghosting Analysis (see images below) |  |  |  | S1
S2
S3
|  |  |  | Ghosting Index | [5 empty] | #DIV/0!
Overall Pass/Fail
|  |  |  | Ghosting Index = | (S3 - S2)
(S1 - S2)
Action Limits |  | Required: | The ghosting index must be within 0±0.3.
|  | Timeframe: | Failures must be corrected before clinical use.
Setup Image 1
Setup Image 2
Evaluation Image

## MP-Sum cont
Medical Physicist's QC Test Summary (cont)
|  | Facility Name | 0 |  | SBBAP ID-Unit# | 0 |  | - | 0
|  | Mfr & Model |  |  | Room ID | 0
[5 empty] | Survey Date | 0
Corrective Action Summary*
*Note: This is only a summary page, the Corrective Action Log Form may contain further details.
Utilize Corrective Action Log Form
Required/ Recommended |  |  | Time Frame | Description |  |  | Date Completed |  |  | Initials

## Tech Chart (Stereotactic)
Technique Chart
|  | Facility Name | 0 |  |  | SBBAP ID-Unit# | 0 | - | 0
|  | Mfr & Model |  |  |  | Room ID | 0
Guidance Mode |  |  | Stereotactic |  |  | Survey Date | 0
AEC Mode
| Compressed Breast Thickness |  |  | 50% Fatty - 50% Dense Breast
|  |  |  | AEC Mode | Target/Filter |  | kVp
< 3 cm
3 to 5 cm
5 to 7 cm
> 7 cm
.
Manual Mode
| Compressed Breast Thickness |  |  | 50% Fatty - 50% Dense Breast
|  |  |  | AEC Mode | Target/Filter |  | kVp
< 3 cm
3 to 5 cm
5 to 7 cm
> 7 cm
QC Phantom Technique
Stereotactic
AEC mode
Paddle size (IR Size)
Paddle type (reg or flex)
View/selected image type
Slice or Slab # (DBT only)
Compression force
AEC cell position (if avail)
Target/filter (if app)
kVp (if app)
Density setting (if app)

## Tech Chart (Tomosynthesis)
Technique Chart
|  | Facility Name | 0 |  |  | SBBAP ID-Unit# | 0 | - | 0
|  | Mfr & Model |  |  |  | Room ID | 0
Guidance Mode |  |  | Tomographic |  |  | Survey Date | 0
AEC Mode
| Compressed Breast Thickness |  |  | 50% Fatty - 50% Dense Breast
|  |  |  | AEC Mode | Target/Filter |  | kVp
< 3 cm
3 to 5 cm
5 to 7 cm
> 7 cm
.
Manual Mode
| Compressed Breast Thickness |  |  | 50% Fatty - 50% Dense Breast
|  |  |  | AEC Mode | Target/Filter |  | kVp
< 3 cm
3 to 5 cm
5 to 7 cm
> 7 cm
QC Phantom Technique
Tomographic
AEC mode
Paddle size (IR Size)
Paddle type (reg or flex)
View/selected image type
Slice or Slab # (DBT only)
Compression force
AEC cell position (if avail)
Target/filter (if app)
kVp (if app)
Density setting (if app)

## Physician Ltr
Medical Physicist QC Letter for the Physician
0
Dr.
0
0
0
0
Re: Medical Physicist Survey of | [5 empty] | on | 0
Dear Lead Radiologist,
The above X-ray guided breast biopsy unit at your facility was surveyed recently. Below is the relevant summary information as a result of this survey. Please note that your facility must follow-up on the Action Items below and obtain relevant documentation from the service engineer. Please evaluate the Phantom image acquired during the medical physicist testing (Image ID information listed below) and see my comments. If you have any questions please don't hesitate to call.
● Image Quality
Patient Name (Phantom):
|  | Patient ID (Phantom): | 0
|  | Date: | 0
|  | Room: | 0
|  | Phantom Type: | 0
Phantom Scores
|  |  | Stereotactic | Tomographic | Passing Criteria | Pass /Fail
|  | Fiber score | [5 empty] | ≥ 2.0
|  | Speck group score | [5 empty] | ≥ 3.0
|  | Mass score | [5 empty] | ≥ 2.0
Artifacts
Comments on phantom image:
● Phantom Average Glandular Dose
Phantom Average Glandular Dose
|  |  | Stereotactic | Tomographic | Passing Criteria | Pass /Fail
|  | Phantom Dose (mGy) |  |  | ≤ 3.0
|  | Note: | The above dose is an estimate determined with a phantom representing the FDA-defined 4.2 cm thick, 50% glandular/50% adipose standard breast. Actual patient doses will vary with patient size and density. Specific patient doses can be estimated by your medical physicist.
Comments on dose findings:

## Physician Ltr cont
Medical Physicist QC Summary Letter for the Physician (cont)
● Required Action Items
|  | Time Frame | Description
● Recommended Action Items
|  | Time Frame | Description
● Comments on Technologist's QC
|  | Time Frame | Description
If you have any questions, please do not hesitate to call.
Sincerely,
|  |  |  | Phone | 0
0 |  |  |  | Email | 0

## DD
PF/PFNA | Pass/Fail
P | Pass
F | Fail
NA | NA
TF | Paddle
Mo/Mo | reg
Mo/Rh | flex
Rh/Al | other
Rh/Ag
Rh/Rh
W/Ag | None
W/Al | None
W/Rh
Target | Phantom Type
Molybdenum | ACR Original
Rhodium | ACR Mini
Tungsten | ACR DM
YN | Test Pattern
Yes | SMPTE
No | TG-18 QC
TG-270 QC
NAYN
NA
Yes
No
Check

Y/N/NA
Yes
No
NA
Timeframe
Immediately
W/in 30 days
NA
Req
Required
Recommended
NA
Display Device
AW
Viewbox
DM-DBT
Stereotactic
DBT
All

## GAMMA_DM_42mm_CC
densityMode | TargetFilter | HVL | kVp | Gamma
FDA | MoMo | 0.2 | 20 | 0.046
FDA | MoMo | 0.2 | 21 | 0.049
FDA | MoMo | 0.2 | 22 | 0.052
FDA | MoMo | 0.2 | 23 | 0.054
FDA | MoMo | 0.2 | 24 | 0.055
FDA | MoMo | 0.2 | 25 | 0.057
FDA | MoMo | 0.21 | 20 | 0.05
FDA | MoMo | 0.21 | 21 | 0.053
FDA | MoMo | 0.21 | 22 | 0.056
FDA | MoMo | 0.21 | 23 | 0.058
FDA | MoMo | 0.21 | 24 | 0.059
FDA | MoMo | 0.21 | 25 | 0.061
FDA | MoMo | 0.22 | 20 | 0.054
FDA | MoMo | 0.22 | 21 | 0.057
FDA | MoMo | 0.22 | 22 | 0.059
FDA | MoMo | 0.22 | 23 | 0.061
FDA | MoMo | 0.22 | 24 | 0.063
FDA | MoMo | 0.22 | 25 | 0.064
FDA | MoMo | 0.22 | 26 | 0.066
FDA | MoMo | 0.23 | 20 | 0.057
FDA | MoMo | 0.23 | 21 | 0.061
FDA | MoMo | 0.23 | 22 | 0.063
FDA | MoMo | 0.23 | 23 | 0.065
FDA | MoMo | 0.23 | 24 | 0.067
FDA | MoMo | 0.23 | 25 | 0.068
FDA | MoMo | 0.23 | 26 | 0.069
FDA | MoMo | 0.23 | 27 | 0.071
FDA | MoMo | 0.23 | 28 | 0.072
FDA | MoMo | 0.24 | 20 | 0.061
FDA | MoMo | 0.24 | 21 | 0.064
FDA | MoMo | 0.24 | 22 | 0.067
FDA | MoMo | 0.24 | 23 | 0.069
FDA | MoMo | 0.24 | 24 | 0.07
FDA | MoMo | 0.24 | 25 | 0.072
FDA | MoMo | 0.24 | 26 | 0.073
FDA | MoMo | 0.24 | 27 | 0.074
FDA | MoMo | 0.24 | 28 | 0.075
FDA | MoMo | 0.24 | 29 | 0.077
FDA | MoMo | 0.25 | 20 | 0.064
FDA | MoMo | 0.25 | 21 | 0.068
FDA | MoMo | 0.25 | 22 | 0.071
FDA | MoMo | 0.25 | 23 | 0.072
FDA | MoMo | 0.25 | 24 | 0.074
FDA | MoMo | 0.25 | 25 | 0.075
FDA | MoMo | 0.25 | 26 | 0.076
FDA | MoMo | 0.25 | 27 | 0.077
FDA | MoMo | 0.25 | 28 | 0.079
FDA | MoMo | 0.25 | 29 | 0.08
FDA | MoMo | 0.25 | 30 | 0.081
FDA | MoMo | 0.25 | 31 | 0.082
FDA | MoMo | 0.26 | 20 | 0.068
FDA | MoMo | 0.26 | 21 | 0.072
FDA | MoMo | 0.26 | 22 | 0.074
FDA | MoMo | 0.26 | 23 | 0.076
FDA | MoMo | 0.26 | 24 | 0.077
FDA | MoMo | 0.26 | 25 | 0.078
FDA | MoMo | 0.26 | 26 | 0.08
FDA | MoMo | 0.26 | 27 | 0.081
FDA | MoMo | 0.26 | 28 | 0.082
FDA | MoMo | 0.26 | 29 | 0.083
FDA | MoMo | 0.26 | 30 | 0.084
FDA | MoMo | 0.26 | 31 | 0.085
FDA | MoMo | 0.26 | 32 | 0.086
FDA | MoMo | 0.27 | 20 | 0.071
FDA | MoMo | 0.27 | 21 | 0.075
FDA | MoMo | 0.27 | 22 | 0.078
FDA | MoMo | 0.27 | 23 | 0.079
FDA | MoMo | 0.27 | 24 | 0.081
FDA | MoMo | 0.27 | 25 | 0.082
FDA | MoMo | 0.27 | 26 | 0.083
FDA | MoMo | 0.27 | 27 | 0.084
FDA | MoMo | 0.27 | 28 | 0.085
FDA | MoMo | 0.27 | 29 | 0.086
FDA | MoMo | 0.27 | 30 | 0.087
FDA | MoMo | 0.27 | 31 | 0.088
FDA | MoMo | 0.27 | 32 | 0.089
FDA | MoMo | 0.27 | 33 | 0.09
FDA | MoMo | 0.27 | 34 | 0.091
FDA | MoMo | 0.28 | 20 | 0.075
FDA | MoMo | 0.28 | 21 | 0.079
FDA | MoMo | 0.28 | 22 | 0.081
FDA | MoMo | 0.28 | 23 | 0.083
FDA | MoMo | 0.28 | 24 | 0.084
FDA | MoMo | 0.28 | 25 | 0.085
FDA | MoMo | 0.28 | 26 | 0.086
FDA | MoMo | 0.28 | 27 | 0.087
FDA | MoMo | 0.28 | 28 | 0.088
FDA | MoMo | 0.28 | 29 | 0.089
FDA | MoMo | 0.28 | 30 | 0.09
FDA | MoMo | 0.28 | 31 | 0.091
FDA | MoMo | 0.28 | 32 | 0.092
FDA | MoMo | 0.28 | 33 | 0.093
FDA | MoMo | 0.28 | 34 | 0.094
FDA | MoMo | 0.28 | 35 | 0.095
FDA | MoMo | 0.28 | 36 | 0.096
FDA | MoMo | 0.29 | 20 | 0.079
FDA | MoMo | 0.29 | 21 | 0.083
FDA | MoMo | 0.29 | 22 | 0.085
FDA | MoMo | 0.29 | 23 | 0.086
FDA | MoMo | 0.29 | 24 | 0.088
FDA | MoMo | 0.29 | 25 | 0.089
FDA | MoMo | 0.29 | 26 | 0.089
FDA | MoMo | 0.29 | 27 | 0.09
FDA | MoMo | 0.29 | 28 | 0.091
FDA | MoMo | 0.29 | 29 | 0.092
FDA | MoMo | 0.29 | 30 | 0.093
FDA | MoMo | 0.29 | 31 | 0.094
FDA | MoMo | 0.29 | 32 | 0.095
FDA | MoMo | 0.29 | 33 | 0.096
FDA | MoMo | 0.29 | 34 | 0.097
FDA | MoMo | 0.29 | 35 | 0.098
FDA | MoMo | 0.29 | 36 | 0.099
FDA | MoMo | 0.29 | 37 | 0.099
FDA | MoMo | 0.29 | 38 | 0.1
FDA | MoMo | 0.3 | 20 | 0.082
FDA | MoMo | 0.3 | 21 | 0.086
FDA | MoMo | 0.3 | 22 | 0.088
FDA | MoMo | 0.3 | 23 | 0.09
FDA | MoMo | 0.3 | 24 | 0.091
FDA | MoMo | 0.3 | 25 | 0.092
FDA | MoMo | 0.3 | 26 | 0.093
FDA | MoMo | 0.3 | 27 | 0.093
FDA | MoMo | 0.3 | 28 | 0.094
FDA | MoMo | 0.3 | 29 | 0.095
FDA | MoMo | 0.3 | 30 | 0.096
FDA | MoMo | 0.3 | 31 | 0.097
FDA | MoMo | 0.3 | 32 | 0.098
FDA | MoMo | 0.3 | 33 | 0.099
FDA | MoMo | 0.3 | 34 | 0.1
FDA | MoMo | 0.3 | 35 | 0.1
FDA | MoMo | 0.3 | 36 | 0.101
FDA | MoMo | 0.3 | 37 | 0.102
FDA | MoMo | 0.3 | 38 | 0.103
FDA | MoMo | 0.3 | 39 | 0.104
FDA | MoMo | 0.3 | 40 | 0.104
FDA | MoMo | 0.3 | 41 | 0.105
FDA | MoMo | 0.3 | 42 | 0.106
FDA | MoMo | 0.31 | 20 | 0.086
FDA | MoMo | 0.31 | 21 | 0.09
FDA | MoMo | 0.31 | 22 | 0.092
FDA | MoMo | 0.31 | 23 | 0.093
FDA | MoMo | 0.31 | 24 | 0.094
FDA | MoMo | 0.31 | 25 | 0.095
FDA | MoMo | 0.31 | 26 | 0.096
FDA | MoMo | 0.31 | 27 | 0.097
FDA | MoMo | 0.31 | 28 | 0.097
FDA | MoMo | 0.31 | 29 | 0.098
FDA | MoMo | 0.31 | 30 | 0.099
FDA | MoMo | 0.31 | 31 | 0.1
FDA | MoMo | 0.31 | 32 | 0.101
FDA | MoMo | 0.31 | 33 | 0.102
FDA | MoMo | 0.31 | 34 | 0.102
FDA | MoMo | 0.31 | 35 | 0.103
FDA | MoMo | 0.31 | 36 | 0.104
FDA | MoMo | 0.31 | 37 | 0.105
FDA | MoMo | 0.31 | 38 | 0.106
FDA | MoMo | 0.31 | 39 | 0.106
FDA | MoMo | 0.31 | 40 | 0.107
FDA | MoMo | 0.31 | 41 | 0.108
FDA | MoMo | 0.31 | 42 | 0.108
FDA | MoMo | 0.31 | 43 | 0.109
FDA | MoMo | 0.32 | 20 | 0.089
FDA | MoMo | 0.32 | 21 | 0.093
FDA | MoMo | 0.32 | 22 | 0.095
FDA | MoMo | 0.32 | 23 | 0.097
FDA | MoMo | 0.32 | 24 | 0.098
FDA | MoMo | 0.32 | 25 | 0.098
FDA | MoMo | 0.32 | 26 | 0.099
FDA | MoMo | 0.32 | 27 | 0.1
FDA | MoMo | 0.32 | 28 | 0.1
FDA | MoMo | 0.32 | 29 | 0.101
FDA | MoMo | 0.32 | 30 | 0.102
FDA | MoMo | 0.32 | 31 | 0.103
FDA | MoMo | 0.32 | 32 | 0.103
FDA | MoMo | 0.32 | 33 | 0.104
FDA | MoMo | 0.32 | 34 | 0.105
FDA | MoMo | 0.32 | 35 | 0.106
FDA | MoMo | 0.32 | 36 | 0.107
FDA | MoMo | 0.32 | 37 | 0.108
FDA | MoMo | 0.32 | 38 | 0.108
FDA | MoMo | 0.32 | 39 | 0.109
FDA | MoMo | 0.32 | 40 | 0.11
FDA | MoMo | 0.32 | 41 | 0.11
FDA | MoMo | 0.32 | 42 | 0.111
FDA | MoMo | 0.32 | 43 | 0.112
FDA | MoMo | 0.32 | 44 | 0.112
FDA | MoMo | 0.32 | 45 | 0.113
FDA | MoMo | 0.32 | 46 | 0.113
FDA | MoMo | 0.33 | 22 | 0.099
FDA | MoMo | 0.33 | 23 | 0.1
FDA | MoMo | 0.33 | 24 | 0.101
FDA | MoMo | 0.33 | 25 | 0.102
FDA | MoMo | 0.33 | 26 | 0.102
FDA | MoMo | 0.33 | 27 | 0.103
FDA | MoMo | 0.33 | 28 | 0.103
FDA | MoMo | 0.33 | 29 | 0.104
FDA | MoMo | 0.33 | 30 | 0.105
FDA | MoMo | 0.33 | 31 | 0.105
FDA | MoMo | 0.33 | 32 | 0.106
FDA | MoMo | 0.33 | 33 | 0.107
FDA | MoMo | 0.33 | 34 | 0.108
FDA | MoMo | 0.33 | 35 | 0.109
FDA | MoMo | 0.33 | 36 | 0.109
FDA | MoMo | 0.33 | 37 | 0.11
FDA | MoMo | 0.33 | 38 | 0.111
FDA | MoMo | 0.33 | 39 | 0.112
FDA | MoMo | 0.33 | 40 | 0.112
FDA | MoMo | 0.33 | 41 | 0.113
FDA | MoMo | 0.33 | 42 | 0.114
FDA | MoMo | 0.33 | 43 | 0.114
FDA | MoMo | 0.33 | 44 | 0.115
FDA | MoMo | 0.33 | 45 | 0.115
FDA | MoMo | 0.33 | 46 | 0.116
FDA | MoMo | 0.33 | 47 | 0.116
FDA | MoMo | 0.33 | 48 | 0.117
FDA | MoMo | 0.33 | 49 | 0.117
FDA | MoMo | 0.34 | 22 | 0.102
FDA | MoMo | 0.34 | 23 | 0.104
FDA | MoMo | 0.34 | 24 | 0.104
FDA | MoMo | 0.34 | 25 | 0.105
FDA | MoMo | 0.34 | 26 | 0.105
FDA | MoMo | 0.34 | 27 | 0.106
FDA | MoMo | 0.34 | 28 | 0.106
FDA | MoMo | 0.34 | 29 | 0.107
FDA | MoMo | 0.34 | 30 | 0.108
FDA | MoMo | 0.34 | 31 | 0.108
FDA | MoMo | 0.34 | 32 | 0.109
FDA | MoMo | 0.34 | 33 | 0.11
FDA | MoMo | 0.34 | 34 | 0.111
FDA | MoMo | 0.34 | 35 | 0.111
FDA | MoMo | 0.34 | 36 | 0.112
FDA | MoMo | 0.34 | 37 | 0.113
FDA | MoMo | 0.34 | 38 | 0.114
FDA | MoMo | 0.34 | 39 | 0.114
FDA | MoMo | 0.34 | 40 | 0.115
FDA | MoMo | 0.34 | 41 | 0.116
FDA | MoMo | 0.34 | 42 | 0.116
FDA | MoMo | 0.34 | 43 | 0.117
FDA | MoMo | 0.34 | 44 | 0.118
FDA | MoMo | 0.34 | 45 | 0.118
FDA | MoMo | 0.34 | 46 | 0.119
FDA | MoMo | 0.34 | 47 | 0.119
FDA | MoMo | 0.34 | 48 | 0.12
FDA | MoMo | 0.34 | 49 | 0.12
FDA | MoMo | 0.35 | 22 | 0.106
FDA | MoMo | 0.35 | 23 | 0.107
FDA | MoMo | 0.35 | 24 | 0.108
FDA | MoMo | 0.35 | 25 | 0.108
FDA | MoMo | 0.35 | 26 | 0.108
FDA | MoMo | 0.35 | 27 | 0.109
FDA | MoMo | 0.35 | 28 | 0.109
FDA | MoMo | 0.35 | 29 | 0.11
FDA | MoMo | 0.35 | 30 | 0.111
FDA | MoMo | 0.35 | 31 | 0.111
FDA | MoMo | 0.35 | 32 | 0.112
FDA | MoMo | 0.35 | 33 | 0.113
FDA | MoMo | 0.35 | 34 | 0.113
FDA | MoMo | 0.35 | 35 | 0.114
FDA | MoMo | 0.35 | 36 | 0.115
FDA | MoMo | 0.35 | 37 | 0.116
FDA | MoMo | 0.35 | 38 | 0.116
FDA | MoMo | 0.35 | 39 | 0.117
FDA | MoMo | 0.35 | 40 | 0.118
FDA | MoMo | 0.35 | 41 | 0.118
FDA | MoMo | 0.35 | 42 | 0.119
FDA | MoMo | 0.35 | 43 | 0.12
FDA | MoMo | 0.35 | 44 | 0.12
FDA | MoMo | 0.35 | 45 | 0.121
FDA | MoMo | 0.35 | 46 | 0.121
FDA | MoMo | 0.35 | 47 | 0.122
FDA | MoMo | 0.35 | 48 | 0.122
FDA | MoMo | 0.35 | 49 | 0.123
FDA | MoMo | 0.36 | 23 | 0.11
FDA | MoMo | 0.36 | 24 | 0.111
FDA | MoMo | 0.36 | 25 | 0.111
FDA | MoMo | 0.36 | 26 | 0.112
FDA | MoMo | 0.36 | 27 | 0.112
FDA | MoMo | 0.36 | 28 | 0.112
FDA | MoMo | 0.36 | 29 | 0.113
FDA | MoMo | 0.36 | 30 | 0.113
FDA | MoMo | 0.36 | 31 | 0.114
FDA | MoMo | 0.36 | 32 | 0.115
FDA | MoMo | 0.36 | 33 | 0.115
FDA | MoMo | 0.36 | 34 | 0.116
FDA | MoMo | 0.36 | 35 | 0.117
FDA | MoMo | 0.36 | 36 | 0.117
FDA | MoMo | 0.36 | 37 | 0.118
FDA | MoMo | 0.36 | 38 | 0.119
FDA | MoMo | 0.36 | 39 | 0.12
FDA | MoMo | 0.36 | 40 | 0.12
FDA | MoMo | 0.36 | 41 | 0.121
FDA | MoMo | 0.36 | 42 | 0.122
FDA | MoMo | 0.36 | 43 | 0.122
FDA | MoMo | 0.36 | 44 | 0.123
FDA | MoMo | 0.36 | 45 | 0.123
FDA | MoMo | 0.36 | 46 | 0.124
FDA | MoMo | 0.36 | 47 | 0.124
FDA | MoMo | 0.36 | 48 | 0.125
FDA | MoMo | 0.36 | 49 | 0.125
FDA | MoMo | 0.37 | 24 | 0.114
FDA | MoMo | 0.37 | 25 | 0.115
FDA | MoMo | 0.37 | 26 | 0.115
FDA | MoMo | 0.37 | 27 | 0.115
FDA | MoMo | 0.37 | 28 | 0.116
FDA | MoMo | 0.37 | 29 | 0.116
FDA | MoMo | 0.37 | 30 | 0.116
FDA | MoMo | 0.37 | 31 | 0.117
FDA | MoMo | 0.37 | 32 | 0.117
FDA | MoMo | 0.37 | 33 | 0.118
FDA | MoMo | 0.37 | 34 | 0.119
FDA | MoMo | 0.37 | 35 | 0.119
FDA | MoMo | 0.37 | 36 | 0.12
FDA | MoMo | 0.37 | 37 | 0.121
FDA | MoMo | 0.37 | 38 | 0.122
FDA | MoMo | 0.37 | 39 | 0.122
FDA | MoMo | 0.37 | 40 | 0.123
FDA | MoMo | 0.37 | 41 | 0.124
FDA | MoMo | 0.37 | 42 | 0.124
FDA | MoMo | 0.37 | 43 | 0.125
FDA | MoMo | 0.37 | 44 | 0.125
FDA | MoMo | 0.37 | 45 | 0.126
FDA | MoMo | 0.37 | 46 | 0.126
FDA | MoMo | 0.37 | 47 | 0.127
FDA | MoMo | 0.37 | 48 | 0.127
FDA | MoMo | 0.37 | 49 | 0.128
FDA | MoMo | 0.38 | 26 | 0.118
FDA | MoMo | 0.38 | 27 | 0.118
FDA | MoMo | 0.38 | 28 | 0.119
FDA | MoMo | 0.38 | 29 | 0.119
FDA | MoMo | 0.38 | 30 | 0.119
FDA | MoMo | 0.38 | 31 | 0.12
FDA | MoMo | 0.38 | 32 | 0.12
FDA | MoMo | 0.38 | 33 | 0.121
FDA | MoMo | 0.38 | 34 | 0.122
FDA | MoMo | 0.38 | 35 | 0.122
FDA | MoMo | 0.38 | 36 | 0.123
FDA | MoMo | 0.38 | 37 | 0.123
FDA | MoMo | 0.38 | 38 | 0.124
FDA | MoMo | 0.38 | 39 | 0.125
FDA | MoMo | 0.38 | 40 | 0.125
FDA | MoMo | 0.38 | 41 | 0.126
FDA | MoMo | 0.38 | 42 | 0.127
FDA | MoMo | 0.38 | 43 | 0.127
FDA | MoMo | 0.38 | 44 | 0.128
FDA | MoMo | 0.38 | 45 | 0.129
FDA | MoMo | 0.38 | 46 | 0.129
FDA | MoMo | 0.38 | 47 | 0.13
FDA | MoMo | 0.38 | 48 | 0.13
FDA | MoMo | 0.38 | 49 | 0.131
FDA | MoMo | 0.39 | 26 | 0.121
FDA | MoMo | 0.39 | 27 | 0.122
FDA | MoMo | 0.39 | 28 | 0.122
FDA | MoMo | 0.39 | 29 | 0.122
FDA | MoMo | 0.39 | 30 | 0.122
FDA | MoMo | 0.39 | 31 | 0.123
FDA | MoMo | 0.39 | 32 | 0.123
FDA | MoMo | 0.39 | 33 | 0.124
FDA | MoMo | 0.39 | 34 | 0.124
FDA | MoMo | 0.39 | 35 | 0.125
FDA | MoMo | 0.39 | 36 | 0.125
FDA | MoMo | 0.39 | 37 | 0.126
FDA | MoMo | 0.39 | 38 | 0.127
FDA | MoMo | 0.39 | 39 | 0.127
FDA | MoMo | 0.39 | 40 | 0.128
FDA | MoMo | 0.39 | 41 | 0.129
FDA | MoMo | 0.39 | 42 | 0.129
FDA | MoMo | 0.39 | 43 | 0.13
FDA | MoMo | 0.39 | 44 | 0.131
FDA | MoMo | 0.39 | 45 | 0.131
FDA | MoMo | 0.39 | 46 | 0.132
FDA | MoMo | 0.39 | 47 | 0.132
FDA | MoMo | 0.39 | 48 | 0.133
FDA | MoMo | 0.39 | 49 | 0.133
FDA | MoMo | 0.4 | 27 | 0.125
FDA | MoMo | 0.4 | 28 | 0.125
FDA | MoMo | 0.4 | 29 | 0.125
FDA | MoMo | 0.4 | 30 | 0.125
FDA | MoMo | 0.4 | 31 | 0.126
FDA | MoMo | 0.4 | 32 | 0.126
FDA | MoMo | 0.4 | 33 | 0.126
FDA | MoMo | 0.4 | 34 | 0.127
FDA | MoMo | 0.4 | 35 | 0.128
FDA | MoMo | 0.4 | 36 | 0.128
FDA | MoMo | 0.4 | 37 | 0.129
FDA | MoMo | 0.4 | 38 | 0.13
FDA | MoMo | 0.4 | 39 | 0.13
FDA | MoMo | 0.4 | 40 | 0.131
FDA | MoMo | 0.4 | 41 | 0.131
FDA | MoMo | 0.4 | 42 | 0.132
FDA | MoMo | 0.4 | 43 | 0.133
FDA | MoMo | 0.4 | 44 | 0.133
FDA | MoMo | 0.4 | 45 | 0.134
FDA | MoMo | 0.4 | 46 | 0.134
FDA | MoMo | 0.4 | 47 | 0.135
FDA | MoMo | 0.4 | 48 | 0.135
FDA | MoMo | 0.4 | 49 | 0.136
FDA | MoMo | 0.41 | 29 | 0.128
FDA | MoMo | 0.41 | 30 | 0.128
FDA | MoMo | 0.41 | 31 | 0.129
FDA | MoMo | 0.41 | 32 | 0.129
FDA | MoMo | 0.41 | 33 | 0.129
FDA | MoMo | 0.41 | 34 | 0.13
FDA | MoMo | 0.41 | 35 | 0.13
FDA | MoMo | 0.41 | 36 | 0.131
FDA | MoMo | 0.41 | 37 | 0.132
FDA | MoMo | 0.41 | 38 | 0.132
FDA | MoMo | 0.41 | 39 | 0.133
FDA | MoMo | 0.41 | 40 | 0.133
FDA | MoMo | 0.41 | 41 | 0.134
FDA | MoMo | 0.41 | 42 | 0.135
FDA | MoMo | 0.41 | 43 | 0.135
FDA | MoMo | 0.41 | 44 | 0.136
FDA | MoMo | 0.41 | 45 | 0.136
FDA | MoMo | 0.41 | 46 | 0.137
FDA | MoMo | 0.41 | 47 | 0.138
FDA | MoMo | 0.41 | 48 | 0.138
FDA | MoMo | 0.41 | 49 | 0.139
FDA | MoMo | 0.42 | 30 | 0.131
FDA | MoMo | 0.42 | 31 | 0.132
FDA | MoMo | 0.42 | 32 | 0.132
FDA | MoMo | 0.42 | 33 | 0.132
FDA | MoMo | 0.42 | 34 | 0.133
FDA | MoMo | 0.42 | 35 | 0.133
FDA | MoMo | 0.42 | 36 | 0.134
FDA | MoMo | 0.42 | 37 | 0.134
FDA | MoMo | 0.42 | 38 | 0.135
FDA | MoMo | 0.42 | 39 | 0.136
FDA | MoMo | 0.42 | 40 | 0.136
FDA | MoMo | 0.42 | 41 | 0.137
FDA | MoMo | 0.42 | 42 | 0.137
FDA | MoMo | 0.42 | 43 | 0.138
FDA | MoMo | 0.42 | 44 | 0.139
FDA | MoMo | 0.42 | 45 | 0.139
FDA | MoMo | 0.42 | 46 | 0.14
FDA | MoMo | 0.42 | 47 | 0.14
FDA | MoMo | 0.42 | 48 | 0.141
FDA | MoMo | 0.42 | 49 | 0.141
FDA | MoMo | 0.43 | 32 | 0.135
FDA | MoMo | 0.43 | 33 | 0.135
FDA | MoMo | 0.43 | 34 | 0.136
FDA | MoMo | 0.43 | 35 | 0.136
FDA | MoMo | 0.43 | 36 | 0.137
FDA | MoMo | 0.43 | 37 | 0.137
FDA | MoMo | 0.43 | 38 | 0.138
FDA | MoMo | 0.43 | 39 | 0.138
FDA | MoMo | 0.43 | 40 | 0.139
FDA | MoMo | 0.43 | 41 | 0.14
FDA | MoMo | 0.43 | 42 | 0.14
FDA | MoMo | 0.43 | 43 | 0.141
FDA | MoMo | 0.43 | 44 | 0.141
FDA | MoMo | 0.43 | 45 | 0.142
FDA | MoMo | 0.43 | 46 | 0.142
FDA | MoMo | 0.43 | 47 | 0.143
FDA | MoMo | 0.43 | 48 | 0.144
FDA | MoMo | 0.43 | 49 | 0.144
FDA | MoMo | 0.44 | 33 | 0.138
FDA | MoMo | 0.44 | 34 | 0.139
FDA | MoMo | 0.44 | 35 | 0.139
FDA | MoMo | 0.44 | 36 | 0.14
FDA | MoMo | 0.44 | 37 | 0.14
FDA | MoMo | 0.44 | 38 | 0.141
FDA | MoMo | 0.44 | 39 | 0.141
FDA | MoMo | 0.44 | 40 | 0.142
FDA | MoMo | 0.44 | 41 | 0.142
FDA | MoMo | 0.44 | 42 | 0.143
FDA | MoMo | 0.44 | 43 | 0.144
FDA | MoMo | 0.44 | 44 | 0.144
FDA | MoMo | 0.44 | 45 | 0.145
FDA | MoMo | 0.44 | 46 | 0.145
FDA | MoMo | 0.44 | 47 | 0.146
FDA | MoMo | 0.44 | 48 | 0.146
FDA | MoMo | 0.44 | 49 | 0.147
FDA | MoMo | 0.45 | 35 | 0.142
FDA | MoMo | 0.45 | 36 | 0.143
FDA | MoMo | 0.45 | 37 | 0.143
FDA | MoMo | 0.45 | 38 | 0.144
FDA | MoMo | 0.45 | 39 | 0.144
FDA | MoMo | 0.45 | 40 | 0.145
FDA | MoMo | 0.45 | 41 | 0.145
FDA | MoMo | 0.45 | 42 | 0.146
FDA | MoMo | 0.45 | 43 | 0.146
FDA | MoMo | 0.45 | 44 | 0.147
FDA | MoMo | 0.45 | 45 | 0.148
FDA | MoMo | 0.45 | 46 | 0.148
FDA | MoMo | 0.45 | 47 | 0.149
FDA | MoMo | 0.45 | 48 | 0.149
FDA | MoMo | 0.45 | 49 | 0.15
FDA | MoMo | 0.46 | 37 | 0.146
FDA | MoMo | 0.46 | 38 | 0.147
FDA | MoMo | 0.46 | 39 | 0.147
FDA | MoMo | 0.46 | 40 | 0.148
FDA | MoMo | 0.46 | 41 | 0.148
FDA | MoMo | 0.46 | 42 | 0.149
FDA | MoMo | 0.46 | 43 | 0.149
FDA | MoMo | 0.46 | 44 | 0.15
FDA | MoMo | 0.46 | 45 | 0.151
FDA | MoMo | 0.46 | 46 | 0.151
FDA | MoMo | 0.46 | 47 | 0.152
FDA | MoMo | 0.46 | 48 | 0.153
FDA | MoMo | 0.46 | 49 | 0.153
FDA | MoMo | 0.47 | 39 | 0.15
FDA | MoMo | 0.47 | 40 | 0.151
FDA | MoMo | 0.47 | 41 | 0.151
FDA | MoMo | 0.47 | 42 | 0.152
FDA | MoMo | 0.47 | 43 | 0.153
FDA | MoMo | 0.47 | 44 | 0.153
FDA | MoMo | 0.47 | 45 | 0.154
FDA | MoMo | 0.47 | 46 | 0.155
FDA | MoMo | 0.47 | 47 | 0.155
FDA | MoMo | 0.47 | 48 | 0.156
FDA | MoMo | 0.47 | 49 | 0.156
FDA | MoMo | 0.48 | 43 | 0.156
FDA | MoMo | 0.48 | 44 | 0.157
FDA | MoMo | 0.48 | 45 | 0.157
FDA | MoMo | 0.48 | 46 | 0.158
FDA | MoMo | 0.48 | 47 | 0.159
FDA | MoMo | 0.48 | 48 | 0.159
FDA | MoMo | 0.48 | 49 | 0.16
FDA | MoMo | 0.49 | 44 | 0.16
FDA | MoMo | 0.49 | 45 | 0.161
FDA | MoMo | 0.49 | 46 | 0.161
FDA | MoMo | 0.49 | 47 | 0.162
FDA | MoMo | 0.49 | 48 | 0.163
FDA | MoMo | 0.49 | 49 | 0.164
FDA | MoMo | 0.5 | 47 | 0.166
FDA | MoMo | 0.5 | 48 | 0.167
FDA | MoMo | 0.5 | 49 | 0.167
FDA | MoRh | 0.195 | 20 | 0.044
FDA | MoRh | 0.195 | 21 | 0.048
FDA | MoRh | 0.195 | 22 | 0.05
FDA | MoRh | 0.21 | 20 | 0.05
FDA | MoRh | 0.21 | 21 | 0.053
FDA | MoRh | 0.21 | 22 | 0.056
FDA | MoRh | 0.21 | 23 | 0.059
FDA | MoRh | 0.21 | 24 | 0.061
FDA | MoRh | 0.225 | 20 | 0.055
FDA | MoRh | 0.225 | 21 | 0.059
FDA | MoRh | 0.225 | 22 | 0.062
FDA | MoRh | 0.225 | 23 | 0.065
FDA | MoRh | 0.225 | 24 | 0.067
FDA | MoRh | 0.225 | 25 | 0.069
FDA | MoRh | 0.225 | 26 | 0.07
FDA | MoRh | 0.225 | 27 | 0.072
FDA | MoRh | 0.24 | 20 | 0.061
FDA | MoRh | 0.24 | 21 | 0.065
FDA | MoRh | 0.24 | 22 | 0.068
FDA | MoRh | 0.24 | 23 | 0.071
FDA | MoRh | 0.24 | 24 | 0.073
FDA | MoRh | 0.24 | 25 | 0.075
FDA | MoRh | 0.24 | 26 | 0.076
FDA | MoRh | 0.24 | 27 | 0.077
FDA | MoRh | 0.255 | 20 | 0.066
FDA | MoRh | 0.255 | 21 | 0.07
FDA | MoRh | 0.255 | 22 | 0.073
FDA | MoRh | 0.255 | 23 | 0.076
FDA | MoRh | 0.255 | 24 | 0.079
FDA | MoRh | 0.255 | 25 | 0.081
FDA | MoRh | 0.255 | 26 | 0.082
FDA | MoRh | 0.255 | 27 | 0.083
FDA | MoRh | 0.255 | 28 | 0.084
FDA | MoRh | 0.27 | 20 | 0.071
FDA | MoRh | 0.27 | 21 | 0.075
FDA | MoRh | 0.27 | 22 | 0.079
FDA | MoRh | 0.27 | 23 | 0.082
FDA | MoRh | 0.27 | 24 | 0.084
FDA | MoRh | 0.27 | 25 | 0.086
FDA | MoRh | 0.27 | 26 | 0.088
FDA | MoRh | 0.27 | 27 | 0.089
FDA | MoRh | 0.27 | 28 | 0.09
FDA | MoRh | 0.27 | 29 | 0.09
FDA | MoRh | 0.27 | 30 | 0.091
FDA | MoRh | 0.27 | 31 | 0.092
FDA | MoRh | 0.285 | 20 | 0.077
FDA | MoRh | 0.285 | 21 | 0.081
FDA | MoRh | 0.285 | 22 | 0.084
FDA | MoRh | 0.285 | 23 | 0.087
FDA | MoRh | 0.285 | 24 | 0.09
FDA | MoRh | 0.285 | 25 | 0.092
FDA | MoRh | 0.285 | 26 | 0.093
FDA | MoRh | 0.285 | 27 | 0.094
FDA | MoRh | 0.285 | 28 | 0.095
FDA | MoRh | 0.285 | 29 | 0.096
FDA | MoRh | 0.285 | 30 | 0.096
FDA | MoRh | 0.285 | 31 | 0.097
FDA | MoRh | 0.285 | 32 | 0.098
FDA | MoRh | 0.285 | 33 | 0.098
FDA | MoRh | 0.3 | 20 | 0.082
FDA | MoRh | 0.3 | 21 | 0.086
FDA | MoRh | 0.3 | 22 | 0.09
FDA | MoRh | 0.3 | 23 | 0.093
FDA | MoRh | 0.3 | 24 | 0.095
FDA | MoRh | 0.3 | 25 | 0.097
FDA | MoRh | 0.3 | 26 | 0.098
FDA | MoRh | 0.3 | 27 | 0.099
FDA | MoRh | 0.3 | 28 | 0.1
FDA | MoRh | 0.3 | 29 | 0.101
FDA | MoRh | 0.3 | 30 | 0.101
FDA | MoRh | 0.3 | 31 | 0.102
FDA | MoRh | 0.3 | 32 | 0.103
FDA | MoRh | 0.3 | 33 | 0.103
FDA | MoRh | 0.3 | 34 | 0.104
FDA | MoRh | 0.3 | 35 | 0.105
FDA | MoRh | 0.3 | 36 | 0.105
FDA | MoRh | 0.3 | 37 | 0.106
FDA | MoRh | 0.3 | 38 | 0.106
FDA | MoRh | 0.3 | 39 | 0.107
FDA | MoRh | 0.3 | 40 | 0.107
FDA | MoRh | 0.3 | 41 | 0.108
FDA | MoRh | 0.3 | 42 | 0.108
FDA | MoRh | 0.315 | 20 | 0.088
FDA | MoRh | 0.315 | 21 | 0.092
FDA | MoRh | 0.315 | 22 | 0.095
FDA | MoRh | 0.315 | 23 | 0.098
FDA | MoRh | 0.315 | 24 | 0.101
FDA | MoRh | 0.315 | 25 | 0.102
FDA | MoRh | 0.315 | 26 | 0.104
FDA | MoRh | 0.315 | 27 | 0.104
FDA | MoRh | 0.315 | 28 | 0.105
FDA | MoRh | 0.315 | 29 | 0.106
FDA | MoRh | 0.315 | 30 | 0.106
FDA | MoRh | 0.315 | 31 | 0.107
FDA | MoRh | 0.315 | 32 | 0.108
FDA | MoRh | 0.315 | 33 | 0.108
FDA | MoRh | 0.315 | 34 | 0.109
FDA | MoRh | 0.315 | 35 | 0.109
FDA | MoRh | 0.315 | 36 | 0.11
FDA | MoRh | 0.315 | 37 | 0.11
FDA | MoRh | 0.315 | 38 | 0.111
FDA | MoRh | 0.315 | 39 | 0.111
FDA | MoRh | 0.315 | 40 | 0.112
FDA | MoRh | 0.315 | 41 | 0.112
FDA | MoRh | 0.315 | 42 | 0.113
FDA | MoRh | 0.33 | 20 | 0.093
FDA | MoRh | 0.33 | 21 | 0.097
FDA | MoRh | 0.33 | 22 | 0.1
FDA | MoRh | 0.33 | 23 | 0.103
FDA | MoRh | 0.33 | 24 | 0.106
FDA | MoRh | 0.33 | 25 | 0.108
FDA | MoRh | 0.33 | 26 | 0.109
FDA | MoRh | 0.33 | 27 | 0.11
FDA | MoRh | 0.33 | 28 | 0.11
FDA | MoRh | 0.33 | 29 | 0.111
FDA | MoRh | 0.33 | 30 | 0.111
FDA | MoRh | 0.33 | 31 | 0.112
FDA | MoRh | 0.33 | 32 | 0.112
FDA | MoRh | 0.33 | 33 | 0.113
FDA | MoRh | 0.33 | 34 | 0.113
FDA | MoRh | 0.33 | 35 | 0.114
FDA | MoRh | 0.33 | 36 | 0.114
FDA | MoRh | 0.33 | 37 | 0.115
FDA | MoRh | 0.33 | 38 | 0.115
FDA | MoRh | 0.33 | 39 | 0.116
FDA | MoRh | 0.33 | 40 | 0.116
FDA | MoRh | 0.33 | 41 | 0.117
FDA | MoRh | 0.33 | 42 | 0.117
FDA | MoRh | 0.345 | 20 | 0.098
FDA | MoRh | 0.345 | 21 | 0.102
FDA | MoRh | 0.345 | 22 | 0.106
FDA | MoRh | 0.345 | 23 | 0.109
FDA | MoRh | 0.345 | 24 | 0.111
FDA | MoRh | 0.345 | 25 | 0.113
FDA | MoRh | 0.345 | 26 | 0.114
FDA | MoRh | 0.345 | 27 | 0.115
FDA | MoRh | 0.345 | 28 | 0.115
FDA | MoRh | 0.345 | 29 | 0.116
FDA | MoRh | 0.345 | 30 | 0.116
FDA | MoRh | 0.345 | 31 | 0.117
FDA | MoRh | 0.345 | 32 | 0.117
FDA | MoRh | 0.345 | 33 | 0.118
FDA | MoRh | 0.345 | 34 | 0.118
FDA | MoRh | 0.345 | 35 | 0.119
FDA | MoRh | 0.345 | 36 | 0.119
FDA | MoRh | 0.345 | 37 | 0.119
FDA | MoRh | 0.345 | 38 | 0.12
FDA | MoRh | 0.345 | 39 | 0.121
FDA | MoRh | 0.345 | 40 | 0.121
FDA | MoRh | 0.345 | 41 | 0.121
FDA | MoRh | 0.345 | 42 | 0.122
FDA | MoRh | 0.345 | 43 | 0.122
FDA | MoRh | 0.345 | 44 | 0.123
FDA | MoRh | 0.345 | 45 | 0.123
FDA | MoRh | 0.345 | 46 | 0.123
FDA | MoRh | 0.36 | 20 | 0.104
FDA | MoRh | 0.36 | 21 | 0.108
FDA | MoRh | 0.36 | 22 | 0.111
FDA | MoRh | 0.36 | 23 | 0.114
FDA | MoRh | 0.36 | 24 | 0.117
FDA | MoRh | 0.36 | 25 | 0.118
FDA | MoRh | 0.36 | 26 | 0.119
FDA | MoRh | 0.36 | 27 | 0.12
FDA | MoRh | 0.36 | 28 | 0.12
FDA | MoRh | 0.36 | 29 | 0.121
FDA | MoRh | 0.36 | 30 | 0.121
FDA | MoRh | 0.36 | 31 | 0.121
FDA | MoRh | 0.36 | 32 | 0.122
FDA | MoRh | 0.36 | 33 | 0.122
FDA | MoRh | 0.36 | 34 | 0.123
FDA | MoRh | 0.36 | 35 | 0.123
FDA | MoRh | 0.36 | 36 | 0.124
FDA | MoRh | 0.36 | 37 | 0.124
FDA | MoRh | 0.36 | 38 | 0.124
FDA | MoRh | 0.36 | 39 | 0.125
FDA | MoRh | 0.36 | 40 | 0.125
FDA | MoRh | 0.36 | 41 | 0.126
FDA | MoRh | 0.36 | 42 | 0.126
FDA | MoRh | 0.36 | 43 | 0.127
FDA | MoRh | 0.36 | 44 | 0.127
FDA | MoRh | 0.36 | 45 | 0.127
FDA | MoRh | 0.36 | 46 | 0.128
FDA | MoRh | 0.36 | 47 | 0.128
FDA | MoRh | 0.36 | 48 | 0.128
FDA | MoRh | 0.36 | 49 | 0.129
FDA | MoRh | 0.375 | 20 | 0.109
FDA | MoRh | 0.375 | 21 | 0.113
FDA | MoRh | 0.375 | 22 | 0.116
FDA | MoRh | 0.375 | 23 | 0.119
FDA | MoRh | 0.375 | 24 | 0.122
FDA | MoRh | 0.375 | 25 | 0.123
FDA | MoRh | 0.375 | 26 | 0.124
FDA | MoRh | 0.375 | 27 | 0.125
FDA | MoRh | 0.375 | 28 | 0.125
FDA | MoRh | 0.375 | 29 | 0.126
FDA | MoRh | 0.375 | 30 | 0.126
FDA | MoRh | 0.375 | 31 | 0.126
FDA | MoRh | 0.375 | 32 | 0.127
FDA | MoRh | 0.375 | 33 | 0.127
FDA | MoRh | 0.375 | 34 | 0.127
FDA | MoRh | 0.375 | 35 | 0.128
FDA | MoRh | 0.375 | 36 | 0.128
FDA | MoRh | 0.375 | 37 | 0.129
FDA | MoRh | 0.375 | 38 | 0.129
FDA | MoRh | 0.375 | 39 | 0.129
FDA | MoRh | 0.375 | 40 | 0.13
FDA | MoRh | 0.375 | 41 | 0.13
FDA | MoRh | 0.375 | 42 | 0.131
FDA | MoRh | 0.375 | 43 | 0.131
FDA | MoRh | 0.375 | 44 | 0.132
FDA | MoRh | 0.375 | 45 | 0.132
FDA | MoRh | 0.375 | 46 | 0.132
FDA | MoRh | 0.375 | 47 | 0.133
FDA | MoRh | 0.375 | 48 | 0.133
FDA | MoRh | 0.375 | 49 | 0.133
FDA | MoRh | 0.39 | 20 | 0.115
FDA | MoRh | 0.39 | 21 | 0.119
FDA | MoRh | 0.39 | 22 | 0.122
FDA | MoRh | 0.39 | 23 | 0.125
FDA | MoRh | 0.39 | 24 | 0.127
FDA | MoRh | 0.39 | 25 | 0.129
FDA | MoRh | 0.39 | 26 | 0.13
FDA | MoRh | 0.39 | 27 | 0.13
FDA | MoRh | 0.39 | 28 | 0.13
FDA | MoRh | 0.39 | 29 | 0.131
FDA | MoRh | 0.39 | 30 | 0.131
FDA | MoRh | 0.39 | 31 | 0.131
FDA | MoRh | 0.39 | 32 | 0.131
FDA | MoRh | 0.39 | 33 | 0.132
FDA | MoRh | 0.39 | 34 | 0.132
FDA | MoRh | 0.39 | 35 | 0.132
FDA | MoRh | 0.39 | 36 | 0.133
FDA | MoRh | 0.39 | 37 | 0.133
FDA | MoRh | 0.39 | 38 | 0.133
FDA | MoRh | 0.39 | 39 | 0.134
FDA | MoRh | 0.39 | 40 | 0.134
FDA | MoRh | 0.39 | 41 | 0.135
FDA | MoRh | 0.39 | 42 | 0.135
FDA | MoRh | 0.39 | 43 | 0.136
FDA | MoRh | 0.39 | 44 | 0.136
FDA | MoRh | 0.39 | 45 | 0.136
FDA | MoRh | 0.39 | 46 | 0.137
FDA | MoRh | 0.39 | 47 | 0.137
FDA | MoRh | 0.39 | 48 | 0.137
FDA | MoRh | 0.39 | 49 | 0.138
FDA | MoRh | 0.405 | 20 | 0.12
FDA | MoRh | 0.405 | 21 | 0.124
FDA | MoRh | 0.405 | 22 | 0.127
FDA | MoRh | 0.405 | 23 | 0.13
FDA | MoRh | 0.405 | 24 | 0.132
FDA | MoRh | 0.405 | 25 | 0.134
FDA | MoRh | 0.405 | 26 | 0.135
FDA | MoRh | 0.405 | 27 | 0.135
FDA | MoRh | 0.405 | 28 | 0.135
FDA | MoRh | 0.405 | 29 | 0.136
FDA | MoRh | 0.405 | 30 | 0.136
FDA | MoRh | 0.405 | 31 | 0.136
FDA | MoRh | 0.405 | 32 | 0.136
FDA | MoRh | 0.405 | 33 | 0.136
FDA | MoRh | 0.405 | 34 | 0.137
FDA | MoRh | 0.405 | 35 | 0.137
FDA | MoRh | 0.405 | 36 | 0.137
FDA | MoRh | 0.405 | 37 | 0.138
FDA | MoRh | 0.405 | 38 | 0.138
FDA | MoRh | 0.405 | 39 | 0.138
FDA | MoRh | 0.405 | 40 | 0.139
FDA | MoRh | 0.405 | 41 | 0.139
FDA | MoRh | 0.405 | 42 | 0.14
FDA | MoRh | 0.405 | 43 | 0.14
FDA | MoRh | 0.405 | 44 | 0.14
FDA | MoRh | 0.405 | 45 | 0.141
FDA | MoRh | 0.405 | 46 | 0.141
FDA | MoRh | 0.405 | 47 | 0.141
FDA | MoRh | 0.405 | 48 | 0.142
FDA | MoRh | 0.405 | 49 | 0.142
FDA | MoRh | 0.42 | 23 | 0.135
FDA | MoRh | 0.42 | 24 | 0.138
FDA | MoRh | 0.42 | 25 | 0.139
FDA | MoRh | 0.42 | 26 | 0.14
FDA | MoRh | 0.42 | 27 | 0.14
FDA | MoRh | 0.42 | 28 | 0.141
FDA | MoRh | 0.42 | 29 | 0.141
FDA | MoRh | 0.42 | 30 | 0.141
FDA | MoRh | 0.42 | 31 | 0.141
FDA | MoRh | 0.42 | 32 | 0.141
FDA | MoRh | 0.42 | 33 | 0.141
FDA | MoRh | 0.42 | 34 | 0.141
FDA | MoRh | 0.42 | 35 | 0.142
FDA | MoRh | 0.42 | 36 | 0.142
FDA | MoRh | 0.42 | 37 | 0.142
FDA | MoRh | 0.42 | 38 | 0.143
FDA | MoRh | 0.42 | 39 | 0.143
FDA | MoRh | 0.42 | 40 | 0.143
FDA | MoRh | 0.42 | 41 | 0.144
FDA | MoRh | 0.42 | 42 | 0.144
FDA | MoRh | 0.42 | 43 | 0.144
FDA | MoRh | 0.42 | 44 | 0.145
FDA | MoRh | 0.42 | 45 | 0.145
FDA | MoRh | 0.42 | 46 | 0.146
FDA | MoRh | 0.42 | 47 | 0.146
FDA | MoRh | 0.42 | 48 | 0.146
FDA | MoRh | 0.42 | 49 | 0.147
FDA | MoRh | 0.435 | 23 | 0.141
FDA | MoRh | 0.435 | 24 | 0.143
FDA | MoRh | 0.435 | 25 | 0.145
FDA | MoRh | 0.435 | 26 | 0.145
FDA | MoRh | 0.435 | 27 | 0.146
FDA | MoRh | 0.435 | 28 | 0.146
FDA | MoRh | 0.435 | 29 | 0.146
FDA | MoRh | 0.435 | 30 | 0.146
FDA | MoRh | 0.435 | 31 | 0.146
FDA | MoRh | 0.435 | 32 | 0.146
FDA | MoRh | 0.435 | 33 | 0.146
FDA | MoRh | 0.435 | 34 | 0.146
FDA | MoRh | 0.435 | 35 | 0.146
FDA | MoRh | 0.435 | 36 | 0.147
FDA | MoRh | 0.435 | 37 | 0.147
FDA | MoRh | 0.435 | 38 | 0.147
FDA | MoRh | 0.435 | 39 | 0.148
FDA | MoRh | 0.435 | 40 | 0.148
FDA | MoRh | 0.435 | 41 | 0.148
FDA | MoRh | 0.435 | 42 | 0.149
FDA | MoRh | 0.435 | 43 | 0.149
FDA | MoRh | 0.435 | 44 | 0.149
FDA | MoRh | 0.435 | 45 | 0.15
FDA | MoRh | 0.435 | 46 | 0.15
FDA | MoRh | 0.435 | 47 | 0.15
FDA | MoRh | 0.435 | 48 | 0.151
FDA | MoRh | 0.435 | 49 | 0.151
FDA | MoRh | 0.45 | 23 | 0.146
FDA | MoRh | 0.45 | 24 | 0.148
FDA | MoRh | 0.45 | 25 | 0.15
FDA | MoRh | 0.45 | 26 | 0.151
FDA | MoRh | 0.45 | 27 | 0.151
FDA | MoRh | 0.45 | 28 | 0.151
FDA | MoRh | 0.45 | 29 | 0.151
FDA | MoRh | 0.45 | 30 | 0.151
FDA | MoRh | 0.45 | 31 | 0.151
FDA | MoRh | 0.45 | 32 | 0.151
FDA | MoRh | 0.45 | 33 | 0.151
FDA | MoRh | 0.45 | 34 | 0.151
FDA | MoRh | 0.45 | 35 | 0.151
FDA | MoRh | 0.45 | 36 | 0.151
FDA | MoRh | 0.45 | 37 | 0.152
FDA | MoRh | 0.45 | 38 | 0.152
FDA | MoRh | 0.45 | 39 | 0.152
FDA | MoRh | 0.45 | 40 | 0.153
FDA | MoRh | 0.45 | 41 | 0.153
FDA | MoRh | 0.45 | 42 | 0.153
FDA | MoRh | 0.45 | 43 | 0.154
FDA | MoRh | 0.45 | 44 | 0.154
FDA | MoRh | 0.45 | 45 | 0.154
FDA | MoRh | 0.45 | 46 | 0.155
FDA | MoRh | 0.45 | 47 | 0.155
FDA | MoRh | 0.45 | 48 | 0.155
FDA | MoRh | 0.45 | 49 | 0.156
FDA | MoRh | 0.465 | 23 | 0.152
FDA | MoRh | 0.465 | 24 | 0.154
FDA | MoRh | 0.465 | 25 | 0.155
FDA | MoRh | 0.465 | 26 | 0.156
FDA | MoRh | 0.465 | 27 | 0.156
FDA | MoRh | 0.465 | 28 | 0.156
FDA | MoRh | 0.465 | 29 | 0.156
FDA | MoRh | 0.465 | 30 | 0.156
FDA | MoRh | 0.465 | 31 | 0.156
FDA | MoRh | 0.465 | 32 | 0.156
FDA | MoRh | 0.465 | 33 | 0.156
FDA | MoRh | 0.465 | 34 | 0.156
FDA | MoRh | 0.465 | 35 | 0.156
FDA | MoRh | 0.465 | 36 | 0.156
FDA | MoRh | 0.465 | 37 | 0.157
FDA | MoRh | 0.465 | 38 | 0.157
FDA | MoRh | 0.465 | 39 | 0.157
FDA | MoRh | 0.465 | 40 | 0.157
FDA | MoRh | 0.465 | 41 | 0.158
FDA | MoRh | 0.465 | 42 | 0.158
FDA | MoRh | 0.465 | 43 | 0.158
FDA | MoRh | 0.465 | 44 | 0.159
FDA | MoRh | 0.465 | 45 | 0.159
FDA | MoRh | 0.465 | 46 | 0.16
FDA | MoRh | 0.465 | 47 | 0.16
FDA | MoRh | 0.465 | 48 | 0.16
FDA | MoRh | 0.465 | 49 | 0.161
FDA | MoRh | 0.48 | 25 | 0.161
FDA | MoRh | 0.48 | 26 | 0.161
FDA | MoRh | 0.48 | 27 | 0.161
FDA | MoRh | 0.48 | 28 | 0.161
FDA | MoRh | 0.48 | 29 | 0.161
FDA | MoRh | 0.48 | 30 | 0.161
FDA | MoRh | 0.48 | 31 | 0.161
FDA | MoRh | 0.48 | 32 | 0.161
FDA | MoRh | 0.48 | 33 | 0.161
FDA | MoRh | 0.48 | 34 | 0.161
FDA | MoRh | 0.48 | 35 | 0.161
FDA | MoRh | 0.48 | 36 | 0.161
FDA | MoRh | 0.48 | 37 | 0.162
FDA | MoRh | 0.48 | 38 | 0.162
FDA | MoRh | 0.48 | 39 | 0.162
FDA | MoRh | 0.48 | 40 | 0.162
FDA | MoRh | 0.48 | 41 | 0.163
FDA | MoRh | 0.48 | 42 | 0.163
FDA | MoRh | 0.48 | 43 | 0.163
FDA | MoRh | 0.48 | 44 | 0.164
FDA | MoRh | 0.48 | 45 | 0.164
FDA | MoRh | 0.48 | 46 | 0.164
FDA | MoRh | 0.48 | 47 | 0.165
FDA | MoRh | 0.48 | 48 | 0.165
FDA | MoRh | 0.48 | 49 | 0.166
FDA | MoRh | 0.495 | 28 | 0.167
FDA | MoRh | 0.495 | 29 | 0.167
FDA | MoRh | 0.495 | 30 | 0.167
FDA | MoRh | 0.495 | 31 | 0.167
FDA | MoRh | 0.495 | 32 | 0.167
FDA | MoRh | 0.495 | 33 | 0.167
FDA | MoRh | 0.495 | 34 | 0.167
FDA | MoRh | 0.495 | 35 | 0.167
FDA | MoRh | 0.495 | 36 | 0.167
FDA | MoRh | 0.495 | 37 | 0.167
FDA | MoRh | 0.495 | 38 | 0.167
FDA | MoRh | 0.495 | 39 | 0.167
FDA | MoRh | 0.495 | 40 | 0.168
FDA | MoRh | 0.495 | 41 | 0.168
FDA | MoRh | 0.495 | 42 | 0.168
FDA | MoRh | 0.495 | 43 | 0.169
FDA | MoRh | 0.495 | 44 | 0.169
FDA | MoRh | 0.495 | 45 | 0.169
FDA | MoRh | 0.495 | 46 | 0.17
FDA | MoRh | 0.495 | 47 | 0.17
FDA | MoRh | 0.495 | 48 | 0.17
FDA | MoRh | 0.495 | 49 | 0.171
FDA | MoRh | 0.51 | 28 | 0.172
FDA | MoRh | 0.51 | 29 | 0.172
FDA | MoRh | 0.51 | 30 | 0.172
FDA | MoRh | 0.51 | 31 | 0.172
FDA | MoRh | 0.51 | 32 | 0.172
FDA | MoRh | 0.51 | 33 | 0.172
FDA | MoRh | 0.51 | 34 | 0.172
FDA | MoRh | 0.51 | 35 | 0.172
FDA | MoRh | 0.51 | 36 | 0.172
FDA | MoRh | 0.51 | 37 | 0.172
FDA | MoRh | 0.51 | 38 | 0.173
FDA | MoRh | 0.51 | 39 | 0.173
FDA | MoRh | 0.51 | 40 | 0.173
FDA | MoRh | 0.51 | 41 | 0.173
FDA | MoRh | 0.51 | 42 | 0.174
FDA | MoRh | 0.51 | 43 | 0.174
FDA | MoRh | 0.51 | 44 | 0.174
FDA | MoRh | 0.51 | 45 | 0.175
FDA | MoRh | 0.51 | 46 | 0.175
FDA | MoRh | 0.51 | 47 | 0.175
FDA | MoRh | 0.51 | 48 | 0.176
FDA | MoRh | 0.51 | 49 | 0.176
FDA | MoRh | 0.525 | 29 | 0.178
FDA | MoRh | 0.525 | 30 | 0.178
FDA | MoRh | 0.525 | 31 | 0.178
FDA | MoRh | 0.525 | 32 | 0.178
FDA | MoRh | 0.525 | 33 | 0.178
FDA | MoRh | 0.525 | 34 | 0.178
FDA | MoRh | 0.525 | 35 | 0.178
FDA | MoRh | 0.525 | 36 | 0.178
FDA | MoRh | 0.525 | 37 | 0.178
FDA | MoRh | 0.525 | 38 | 0.178
FDA | MoRh | 0.525 | 39 | 0.178
FDA | MoRh | 0.525 | 40 | 0.178
FDA | MoRh | 0.525 | 41 | 0.179
FDA | MoRh | 0.525 | 42 | 0.179
FDA | MoRh | 0.525 | 43 | 0.179
FDA | MoRh | 0.525 | 44 | 0.18
FDA | MoRh | 0.525 | 45 | 0.18
FDA | MoRh | 0.525 | 46 | 0.181
FDA | MoRh | 0.525 | 47 | 0.181
FDA | MoRh | 0.525 | 48 | 0.181
FDA | MoRh | 0.525 | 49 | 0.182
FDA | MoRh | 0.54 | 32 | 0.183
FDA | MoRh | 0.54 | 33 | 0.183
FDA | MoRh | 0.54 | 34 | 0.183
FDA | MoRh | 0.54 | 35 | 0.183
FDA | MoRh | 0.54 | 36 | 0.183
FDA | MoRh | 0.54 | 37 | 0.183
FDA | MoRh | 0.54 | 38 | 0.184
FDA | MoRh | 0.54 | 39 | 0.184
FDA | MoRh | 0.54 | 40 | 0.184
FDA | MoRh | 0.54 | 41 | 0.184
FDA | MoRh | 0.54 | 42 | 0.185
FDA | MoRh | 0.54 | 43 | 0.185
FDA | MoRh | 0.54 | 44 | 0.185
FDA | MoRh | 0.54 | 45 | 0.186
FDA | MoRh | 0.54 | 46 | 0.186
FDA | MoRh | 0.54 | 47 | 0.187
FDA | MoRh | 0.54 | 48 | 0.187
FDA | MoRh | 0.54 | 49 | 0.188
FDA | MoRh | 0.555 | 34 | 0.189
FDA | MoRh | 0.555 | 35 | 0.189
FDA | MoRh | 0.555 | 36 | 0.189
FDA | MoRh | 0.555 | 37 | 0.189
FDA | MoRh | 0.555 | 38 | 0.189
FDA | MoRh | 0.555 | 39 | 0.189
FDA | MoRh | 0.555 | 40 | 0.19
FDA | MoRh | 0.555 | 41 | 0.19
FDA | MoRh | 0.555 | 42 | 0.19
FDA | MoRh | 0.555 | 43 | 0.191
FDA | MoRh | 0.555 | 44 | 0.191
FDA | MoRh | 0.555 | 45 | 0.192
FDA | MoRh | 0.555 | 46 | 0.192
FDA | MoRh | 0.555 | 47 | 0.193
FDA | MoRh | 0.555 | 48 | 0.193
FDA | MoRh | 0.555 | 49 | 0.194
FDA | MoRh | 0.57 | 43 | 0.196
FDA | MoRh | 0.57 | 44 | 0.197
FDA | MoRh | 0.57 | 45 | 0.197
FDA | MoRh | 0.57 | 46 | 0.198
FDA | MoRh | 0.57 | 47 | 0.198
FDA | MoRh | 0.57 | 48 | 0.199
FDA | MoRh | 0.57 | 49 | 0.199
FDA | MoRh | 0.585 | 43 | 0.202
FDA | MoRh | 0.585 | 44 | 0.203
FDA | MoRh | 0.585 | 45 | 0.203
FDA | MoRh | 0.585 | 46 | 0.204
FDA | MoRh | 0.585 | 47 | 0.204
FDA | MoRh | 0.585 | 48 | 0.205
FDA | MoRh | 0.585 | 49 | 0.205
FDA | MoRh | 0.6 | 44 | 0.208
FDA | MoRh | 0.6 | 45 | 0.209
FDA | MoRh | 0.6 | 46 | 0.209
FDA | MoRh | 0.6 | 47 | 0.21
FDA | MoRh | 0.6 | 48 | 0.21
FDA | MoRh | 0.6 | 49 | 0.211
FDA | RhRh | 0.2 | 20 | 0.046
FDA | RhRh | 0.2 | 21 | 0.049
FDA | RhRh | 0.2 | 22 | 0.051
FDA | RhRh | 0.2 | 23 | 0.053
FDA | RhRh | 0.2 | 24 | 0.056
FDA | RhRh | 0.2 | 25 | 0.059
FDA | RhRh | 0.22 | 20 | 0.053
FDA | RhRh | 0.22 | 21 | 0.057
FDA | RhRh | 0.22 | 22 | 0.059
FDA | RhRh | 0.22 | 23 | 0.061
FDA | RhRh | 0.22 | 24 | 0.065
FDA | RhRh | 0.22 | 25 | 0.068
FDA | RhRh | 0.22 | 26 | 0.07
FDA | RhRh | 0.22 | 27 | 0.073
FDA | RhRh | 0.24 | 20 | 0.061
FDA | RhRh | 0.24 | 21 | 0.064
FDA | RhRh | 0.24 | 22 | 0.067
FDA | RhRh | 0.24 | 23 | 0.069
FDA | RhRh | 0.24 | 24 | 0.073
FDA | RhRh | 0.24 | 25 | 0.076
FDA | RhRh | 0.24 | 26 | 0.079
FDA | RhRh | 0.24 | 27 | 0.081
FDA | RhRh | 0.24 | 28 | 0.083
FDA | RhRh | 0.24 | 29 | 0.085
FDA | RhRh | 0.26 | 20 | 0.068
FDA | RhRh | 0.26 | 21 | 0.072
FDA | RhRh | 0.26 | 22 | 0.074
FDA | RhRh | 0.26 | 23 | 0.077
FDA | RhRh | 0.26 | 24 | 0.081
FDA | RhRh | 0.26 | 25 | 0.085
FDA | RhRh | 0.26 | 26 | 0.087
FDA | RhRh | 0.26 | 27 | 0.09
FDA | RhRh | 0.26 | 28 | 0.092
FDA | RhRh | 0.26 | 29 | 0.094
FDA | RhRh | 0.26 | 30 | 0.096
FDA | RhRh | 0.26 | 31 | 0.097
FDA | RhRh | 0.26 | 32 | 0.099
FDA | RhRh | 0.28 | 20 | 0.075
FDA | RhRh | 0.28 | 21 | 0.079
FDA | RhRh | 0.28 | 22 | 0.082
FDA | RhRh | 0.28 | 23 | 0.085
FDA | RhRh | 0.28 | 24 | 0.089
FDA | RhRh | 0.28 | 25 | 0.093
FDA | RhRh | 0.28 | 26 | 0.095
FDA | RhRh | 0.28 | 27 | 0.098
FDA | RhRh | 0.28 | 28 | 0.1
FDA | RhRh | 0.28 | 29 | 0.102
FDA | RhRh | 0.28 | 30 | 0.104
FDA | RhRh | 0.28 | 31 | 0.105
FDA | RhRh | 0.28 | 32 | 0.107
FDA | RhRh | 0.28 | 33 | 0.108
FDA | RhRh | 0.28 | 34 | 0.11
FDA | RhRh | 0.3 | 20 | 0.082
FDA | RhRh | 0.3 | 21 | 0.086
FDA | RhRh | 0.3 | 22 | 0.089
FDA | RhRh | 0.3 | 23 | 0.093
FDA | RhRh | 0.3 | 24 | 0.097
FDA | RhRh | 0.3 | 25 | 0.1
FDA | RhRh | 0.3 | 26 | 0.103
FDA | RhRh | 0.3 | 27 | 0.106
FDA | RhRh | 0.3 | 28 | 0.108
FDA | RhRh | 0.3 | 29 | 0.109
FDA | RhRh | 0.3 | 30 | 0.111
FDA | RhRh | 0.3 | 31 | 0.113
FDA | RhRh | 0.3 | 32 | 0.114
FDA | RhRh | 0.3 | 33 | 0.116
FDA | RhRh | 0.3 | 34 | 0.117
FDA | RhRh | 0.3 | 35 | 0.118
FDA | RhRh | 0.3 | 36 | 0.12
FDA | RhRh | 0.3 | 37 | 0.121
FDA | RhRh | 0.32 | 20 | 0.089
FDA | RhRh | 0.32 | 21 | 0.093
FDA | RhRh | 0.32 | 22 | 0.097
FDA | RhRh | 0.32 | 23 | 0.1
FDA | RhRh | 0.32 | 24 | 0.105
FDA | RhRh | 0.32 | 25 | 0.108
FDA | RhRh | 0.32 | 26 | 0.111
FDA | RhRh | 0.32 | 27 | 0.113
FDA | RhRh | 0.32 | 28 | 0.115
FDA | RhRh | 0.32 | 29 | 0.117
FDA | RhRh | 0.32 | 30 | 0.118
FDA | RhRh | 0.32 | 31 | 0.12
FDA | RhRh | 0.32 | 32 | 0.121
FDA | RhRh | 0.32 | 33 | 0.123
FDA | RhRh | 0.32 | 34 | 0.124
FDA | RhRh | 0.32 | 35 | 0.125
FDA | RhRh | 0.32 | 36 | 0.127
FDA | RhRh | 0.32 | 37 | 0.128
FDA | RhRh | 0.32 | 38 | 0.129
FDA | RhRh | 0.32 | 39 | 0.13
FDA | RhRh | 0.34 | 20 | 0.096
FDA | RhRh | 0.34 | 21 | 0.101
FDA | RhRh | 0.34 | 22 | 0.104
FDA | RhRh | 0.34 | 23 | 0.108
FDA | RhRh | 0.34 | 24 | 0.112
FDA | RhRh | 0.34 | 25 | 0.116
FDA | RhRh | 0.34 | 26 | 0.118
FDA | RhRh | 0.34 | 27 | 0.12
FDA | RhRh | 0.34 | 28 | 0.122
FDA | RhRh | 0.34 | 29 | 0.124
FDA | RhRh | 0.34 | 30 | 0.126
FDA | RhRh | 0.34 | 31 | 0.127
FDA | RhRh | 0.34 | 32 | 0.128
FDA | RhRh | 0.34 | 33 | 0.13
FDA | RhRh | 0.34 | 34 | 0.131
FDA | RhRh | 0.34 | 35 | 0.132
FDA | RhRh | 0.34 | 36 | 0.133
FDA | RhRh | 0.34 | 37 | 0.134
FDA | RhRh | 0.34 | 38 | 0.136
FDA | RhRh | 0.34 | 39 | 0.137
FDA | RhRh | 0.34 | 40 | 0.138
FDA | RhRh | 0.34 | 41 | 0.139
FDA | RhRh | 0.34 | 42 | 0.139
FDA | RhRh | 0.36 | 20 | 0.104
FDA | RhRh | 0.36 | 21 | 0.108
FDA | RhRh | 0.36 | 22 | 0.111
FDA | RhRh | 0.36 | 23 | 0.115
FDA | RhRh | 0.36 | 24 | 0.12
FDA | RhRh | 0.36 | 25 | 0.123
FDA | RhRh | 0.36 | 26 | 0.126
FDA | RhRh | 0.36 | 27 | 0.128
FDA | RhRh | 0.36 | 28 | 0.129
FDA | RhRh | 0.36 | 29 | 0.131
FDA | RhRh | 0.36 | 30 | 0.132
FDA | RhRh | 0.36 | 31 | 0.134
FDA | RhRh | 0.36 | 32 | 0.135
FDA | RhRh | 0.36 | 33 | 0.136
FDA | RhRh | 0.36 | 34 | 0.138
FDA | RhRh | 0.36 | 35 | 0.139
FDA | RhRh | 0.36 | 36 | 0.14
FDA | RhRh | 0.36 | 37 | 0.141
FDA | RhRh | 0.36 | 38 | 0.142
FDA | RhRh | 0.36 | 39 | 0.143
FDA | RhRh | 0.36 | 40 | 0.144
FDA | RhRh | 0.36 | 41 | 0.145
FDA | RhRh | 0.36 | 42 | 0.146
FDA | RhRh | 0.36 | 43 | 0.147
FDA | RhRh | 0.36 | 44 | 0.147
FDA | RhRh | 0.36 | 45 | 0.148
FDA | RhRh | 0.36 | 46 | 0.149
FDA | RhRh | 0.38 | 20 | 0.111
FDA | RhRh | 0.38 | 21 | 0.116
FDA | RhRh | 0.38 | 22 | 0.119
FDA | RhRh | 0.38 | 23 | 0.122
FDA | RhRh | 0.38 | 24 | 0.127
FDA | RhRh | 0.38 | 25 | 0.13
FDA | RhRh | 0.38 | 26 | 0.133
FDA | RhRh | 0.38 | 27 | 0.135
FDA | RhRh | 0.38 | 28 | 0.136
FDA | RhRh | 0.38 | 29 | 0.138
FDA | RhRh | 0.38 | 30 | 0.139
FDA | RhRh | 0.38 | 31 | 0.141
FDA | RhRh | 0.38 | 32 | 0.142
FDA | RhRh | 0.38 | 33 | 0.143
FDA | RhRh | 0.38 | 34 | 0.144
FDA | RhRh | 0.38 | 35 | 0.145
FDA | RhRh | 0.38 | 36 | 0.146
FDA | RhRh | 0.38 | 37 | 0.147
FDA | RhRh | 0.38 | 38 | 0.148
FDA | RhRh | 0.38 | 39 | 0.149
FDA | RhRh | 0.38 | 40 | 0.15
FDA | RhRh | 0.38 | 41 | 0.151
FDA | RhRh | 0.38 | 42 | 0.152
FDA | RhRh | 0.38 | 43 | 0.153
FDA | RhRh | 0.38 | 44 | 0.154
FDA | RhRh | 0.38 | 45 | 0.154
FDA | RhRh | 0.38 | 46 | 0.155
FDA | RhRh | 0.38 | 47 | 0.156
FDA | RhRh | 0.38 | 48 | 0.156
FDA | RhRh | 0.38 | 49 | 0.157
FDA | RhRh | 0.4 | 20 | 0.118
FDA | RhRh | 0.4 | 21 | 0.123
FDA | RhRh | 0.4 | 22 | 0.126
FDA | RhRh | 0.4 | 23 | 0.13
FDA | RhRh | 0.4 | 24 | 0.134
FDA | RhRh | 0.4 | 25 | 0.138
FDA | RhRh | 0.4 | 26 | 0.14
FDA | RhRh | 0.4 | 27 | 0.142
FDA | RhRh | 0.4 | 28 | 0.143
FDA | RhRh | 0.4 | 29 | 0.145
FDA | RhRh | 0.4 | 30 | 0.146
FDA | RhRh | 0.4 | 31 | 0.147
FDA | RhRh | 0.4 | 32 | 0.148
FDA | RhRh | 0.4 | 33 | 0.149
FDA | RhRh | 0.4 | 34 | 0.15
FDA | RhRh | 0.4 | 35 | 0.151
FDA | RhRh | 0.4 | 36 | 0.153
FDA | RhRh | 0.4 | 37 | 0.154
FDA | RhRh | 0.4 | 38 | 0.155
FDA | RhRh | 0.4 | 39 | 0.155
FDA | RhRh | 0.4 | 40 | 0.156
FDA | RhRh | 0.4 | 41 | 0.157
FDA | RhRh | 0.4 | 42 | 0.158
FDA | RhRh | 0.4 | 43 | 0.159
FDA | RhRh | 0.4 | 44 | 0.16
FDA | RhRh | 0.4 | 45 | 0.16
FDA | RhRh | 0.4 | 46 | 0.161
FDA | RhRh | 0.4 | 47 | 0.162
FDA | RhRh | 0.4 | 48 | 0.162
FDA | RhRh | 0.4 | 49 | 0.163
FDA | RhRh | 0.42 | 21 | 0.13
FDA | RhRh | 0.42 | 22 | 0.134
FDA | RhRh | 0.42 | 23 | 0.137
FDA | RhRh | 0.42 | 24 | 0.142
FDA | RhRh | 0.42 | 25 | 0.145
FDA | RhRh | 0.42 | 26 | 0.147
FDA | RhRh | 0.42 | 27 | 0.149
FDA | RhRh | 0.42 | 28 | 0.15
FDA | RhRh | 0.42 | 29 | 0.151
FDA | RhRh | 0.42 | 30 | 0.152
FDA | RhRh | 0.42 | 31 | 0.154
FDA | RhRh | 0.42 | 32 | 0.155
FDA | RhRh | 0.42 | 33 | 0.156
FDA | RhRh | 0.42 | 34 | 0.157
FDA | RhRh | 0.42 | 35 | 0.158
FDA | RhRh | 0.42 | 36 | 0.159
FDA | RhRh | 0.42 | 37 | 0.16
FDA | RhRh | 0.42 | 38 | 0.16
FDA | RhRh | 0.42 | 39 | 0.161
FDA | RhRh | 0.42 | 40 | 0.162
FDA | RhRh | 0.42 | 41 | 0.163
FDA | RhRh | 0.42 | 42 | 0.164
FDA | RhRh | 0.42 | 43 | 0.165
FDA | RhRh | 0.42 | 44 | 0.166
FDA | RhRh | 0.42 | 45 | 0.166
FDA | RhRh | 0.42 | 46 | 0.167
FDA | RhRh | 0.42 | 47 | 0.168
FDA | RhRh | 0.42 | 48 | 0.168
FDA | RhRh | 0.42 | 49 | 0.169
FDA | RhRh | 0.44 | 21 | 0.138
FDA | RhRh | 0.44 | 22 | 0.141
FDA | RhRh | 0.44 | 23 | 0.145
FDA | RhRh | 0.44 | 24 | 0.149
FDA | RhRh | 0.44 | 25 | 0.152
FDA | RhRh | 0.44 | 26 | 0.154
FDA | RhRh | 0.44 | 27 | 0.156
FDA | RhRh | 0.44 | 28 | 0.157
FDA | RhRh | 0.44 | 29 | 0.158
FDA | RhRh | 0.44 | 30 | 0.159
FDA | RhRh | 0.44 | 31 | 0.16
FDA | RhRh | 0.44 | 32 | 0.161
FDA | RhRh | 0.44 | 33 | 0.162
FDA | RhRh | 0.44 | 34 | 0.163
FDA | RhRh | 0.44 | 35 | 0.164
FDA | RhRh | 0.44 | 36 | 0.165
FDA | RhRh | 0.44 | 37 | 0.166
FDA | RhRh | 0.44 | 38 | 0.166
FDA | RhRh | 0.44 | 39 | 0.167
FDA | RhRh | 0.44 | 40 | 0.168
FDA | RhRh | 0.44 | 41 | 0.169
FDA | RhRh | 0.44 | 42 | 0.17
FDA | RhRh | 0.44 | 43 | 0.171
FDA | RhRh | 0.44 | 44 | 0.171
FDA | RhRh | 0.44 | 45 | 0.172
FDA | RhRh | 0.44 | 46 | 0.173
FDA | RhRh | 0.44 | 47 | 0.173
FDA | RhRh | 0.44 | 48 | 0.174
FDA | RhRh | 0.44 | 49 | 0.174
FDA | RhRh | 0.46 | 22 | 0.148
FDA | RhRh | 0.46 | 23 | 0.152
FDA | RhRh | 0.46 | 24 | 0.156
FDA | RhRh | 0.46 | 25 | 0.159
FDA | RhRh | 0.46 | 26 | 0.161
FDA | RhRh | 0.46 | 27 | 0.162
FDA | RhRh | 0.46 | 28 | 0.163
FDA | RhRh | 0.46 | 29 | 0.164
FDA | RhRh | 0.46 | 30 | 0.165
FDA | RhRh | 0.46 | 31 | 0.166
FDA | RhRh | 0.46 | 32 | 0.167
FDA | RhRh | 0.46 | 33 | 0.168
FDA | RhRh | 0.46 | 34 | 0.169
FDA | RhRh | 0.46 | 35 | 0.17
FDA | RhRh | 0.46 | 36 | 0.17
FDA | RhRh | 0.46 | 37 | 0.171
FDA | RhRh | 0.46 | 38 | 0.172
FDA | RhRh | 0.46 | 39 | 0.173
FDA | RhRh | 0.46 | 40 | 0.174
FDA | RhRh | 0.46 | 41 | 0.175
FDA | RhRh | 0.46 | 42 | 0.175
FDA | RhRh | 0.46 | 43 | 0.176
FDA | RhRh | 0.46 | 44 | 0.177
FDA | RhRh | 0.46 | 45 | 0.178
FDA | RhRh | 0.46 | 46 | 0.178
FDA | RhRh | 0.46 | 47 | 0.179
FDA | RhRh | 0.46 | 48 | 0.179
FDA | RhRh | 0.46 | 49 | 0.18
FDA | RhRh | 0.48 | 22 | 0.156
FDA | RhRh | 0.48 | 23 | 0.159
FDA | RhRh | 0.48 | 24 | 0.163
FDA | RhRh | 0.48 | 25 | 0.166
FDA | RhRh | 0.48 | 26 | 0.168
FDA | RhRh | 0.48 | 27 | 0.169
FDA | RhRh | 0.48 | 28 | 0.17
FDA | RhRh | 0.48 | 29 | 0.171
FDA | RhRh | 0.48 | 30 | 0.172
FDA | RhRh | 0.48 | 31 | 0.173
FDA | RhRh | 0.48 | 32 | 0.173
FDA | RhRh | 0.48 | 33 | 0.174
FDA | RhRh | 0.48 | 34 | 0.175
FDA | RhRh | 0.48 | 35 | 0.176
FDA | RhRh | 0.48 | 36 | 0.176
FDA | RhRh | 0.48 | 37 | 0.177
FDA | RhRh | 0.48 | 38 | 0.178
FDA | RhRh | 0.48 | 39 | 0.179
FDA | RhRh | 0.48 | 40 | 0.18
FDA | RhRh | 0.48 | 41 | 0.18
FDA | RhRh | 0.48 | 42 | 0.181
FDA | RhRh | 0.48 | 43 | 0.182
FDA | RhRh | 0.48 | 44 | 0.182
FDA | RhRh | 0.48 | 45 | 0.183
FDA | RhRh | 0.48 | 46 | 0.184
FDA | RhRh | 0.48 | 47 | 0.184
FDA | RhRh | 0.48 | 48 | 0.185
FDA | RhRh | 0.48 | 49 | 0.186
FDA | RhRh | 0.5 | 23 | 0.167
FDA | RhRh | 0.5 | 24 | 0.17
FDA | RhRh | 0.5 | 25 | 0.173
FDA | RhRh | 0.5 | 26 | 0.175
FDA | RhRh | 0.5 | 27 | 0.176
FDA | RhRh | 0.5 | 28 | 0.177
FDA | RhRh | 0.5 | 29 | 0.177
FDA | RhRh | 0.5 | 30 | 0.178
FDA | RhRh | 0.5 | 31 | 0.179
FDA | RhRh | 0.5 | 32 | 0.179
FDA | RhRh | 0.5 | 33 | 0.18
FDA | RhRh | 0.5 | 34 | 0.181
FDA | RhRh | 0.5 | 35 | 0.181
FDA | RhRh | 0.5 | 36 | 0.182
FDA | RhRh | 0.5 | 37 | 0.183
FDA | RhRh | 0.5 | 38 | 0.184
FDA | RhRh | 0.5 | 39 | 0.184
FDA | RhRh | 0.5 | 40 | 0.185
FDA | RhRh | 0.5 | 41 | 0.186
FDA | RhRh | 0.5 | 42 | 0.187
FDA | RhRh | 0.5 | 43 | 0.187
FDA | RhRh | 0.5 | 44 | 0.188
FDA | RhRh | 0.5 | 45 | 0.189
FDA | RhRh | 0.5 | 46 | 0.189
FDA | RhRh | 0.5 | 47 | 0.19
FDA | RhRh | 0.5 | 48 | 0.19
FDA | RhRh | 0.5 | 49 | 0.191
FDA | RhRh | 0.52 | 23 | 0.174
FDA | RhRh | 0.52 | 24 | 0.177
FDA | RhRh | 0.52 | 25 | 0.18
FDA | RhRh | 0.52 | 26 | 0.181
FDA | RhRh | 0.52 | 27 | 0.182
FDA | RhRh | 0.52 | 28 | 0.183
FDA | RhRh | 0.52 | 29 | 0.184
FDA | RhRh | 0.52 | 30 | 0.184
FDA | RhRh | 0.52 | 31 | 0.185
FDA | RhRh | 0.52 | 32 | 0.185
FDA | RhRh | 0.52 | 33 | 0.186
FDA | RhRh | 0.52 | 34 | 0.187
FDA | RhRh | 0.52 | 35 | 0.187
FDA | RhRh | 0.52 | 36 | 0.188
FDA | RhRh | 0.52 | 37 | 0.189
FDA | RhRh | 0.52 | 38 | 0.189
FDA | RhRh | 0.52 | 39 | 0.19
FDA | RhRh | 0.52 | 40 | 0.191
FDA | RhRh | 0.52 | 41 | 0.191
FDA | RhRh | 0.52 | 42 | 0.192
FDA | RhRh | 0.52 | 43 | 0.193
FDA | RhRh | 0.52 | 44 | 0.193
FDA | RhRh | 0.52 | 45 | 0.194
FDA | RhRh | 0.52 | 46 | 0.195
FDA | RhRh | 0.52 | 47 | 0.195
FDA | RhRh | 0.52 | 48 | 0.196
FDA | RhRh | 0.52 | 49 | 0.196
FDA | RhRh | 0.54 | 24 | 0.184
FDA | RhRh | 0.54 | 25 | 0.187
FDA | RhRh | 0.54 | 26 | 0.188
FDA | RhRh | 0.54 | 27 | 0.189
FDA | RhRh | 0.54 | 28 | 0.19
FDA | RhRh | 0.54 | 29 | 0.19
FDA | RhRh | 0.54 | 30 | 0.191
FDA | RhRh | 0.54 | 31 | 0.191
FDA | RhRh | 0.54 | 32 | 0.192
FDA | RhRh | 0.54 | 33 | 0.192
FDA | RhRh | 0.54 | 34 | 0.193
FDA | RhRh | 0.54 | 35 | 0.193
FDA | RhRh | 0.54 | 36 | 0.194
FDA | RhRh | 0.54 | 37 | 0.194
FDA | RhRh | 0.54 | 38 | 0.195
FDA | RhRh | 0.54 | 39 | 0.196
FDA | RhRh | 0.54 | 40 | 0.196
FDA | RhRh | 0.54 | 41 | 0.197
FDA | RhRh | 0.54 | 42 | 0.198
FDA | RhRh | 0.54 | 43 | 0.198
FDA | RhRh | 0.54 | 44 | 0.199
FDA | RhRh | 0.54 | 45 | 0.199
FDA | RhRh | 0.54 | 46 | 0.2
FDA | RhRh | 0.54 | 47 | 0.201
FDA | RhRh | 0.54 | 48 | 0.201
FDA | RhRh | 0.54 | 49 | 0.202
FDA | RhRh | 0.56 | 26 | 0.195
FDA | RhRh | 0.56 | 27 | 0.195
FDA | RhRh | 0.56 | 28 | 0.196
FDA | RhRh | 0.56 | 29 | 0.196
FDA | RhRh | 0.56 | 30 | 0.197
FDA | RhRh | 0.56 | 31 | 0.197
FDA | RhRh | 0.56 | 32 | 0.198
FDA | RhRh | 0.56 | 33 | 0.198
FDA | RhRh | 0.56 | 34 | 0.198
FDA | RhRh | 0.56 | 35 | 0.199
FDA | RhRh | 0.56 | 36 | 0.199
FDA | RhRh | 0.56 | 37 | 0.2
FDA | RhRh | 0.56 | 38 | 0.201
FDA | RhRh | 0.56 | 39 | 0.201
FDA | RhRh | 0.56 | 40 | 0.202
FDA | RhRh | 0.56 | 41 | 0.202
FDA | RhRh | 0.56 | 42 | 0.203
FDA | RhRh | 0.56 | 43 | 0.204
FDA | RhRh | 0.56 | 44 | 0.204
FDA | RhRh | 0.56 | 45 | 0.205
FDA | RhRh | 0.56 | 46 | 0.205
FDA | RhRh | 0.56 | 47 | 0.206
FDA | RhRh | 0.56 | 48 | 0.206
FDA | RhRh | 0.56 | 49 | 0.207
FDA | RhRh | 0.58 | 28 | 0.202
FDA | RhRh | 0.58 | 29 | 0.203
FDA | RhRh | 0.58 | 30 | 0.203
FDA | RhRh | 0.58 | 31 | 0.203
FDA | RhRh | 0.58 | 32 | 0.204
FDA | RhRh | 0.58 | 33 | 0.204
FDA | RhRh | 0.58 | 34 | 0.204
FDA | RhRh | 0.58 | 35 | 0.205
FDA | RhRh | 0.58 | 36 | 0.205
FDA | RhRh | 0.58 | 37 | 0.206
FDA | RhRh | 0.58 | 38 | 0.206
FDA | RhRh | 0.58 | 39 | 0.207
FDA | RhRh | 0.58 | 40 | 0.207
FDA | RhRh | 0.58 | 41 | 0.208
FDA | RhRh | 0.58 | 42 | 0.208
FDA | RhRh | 0.58 | 43 | 0.209
FDA | RhRh | 0.58 | 44 | 0.21
FDA | RhRh | 0.58 | 45 | 0.21
FDA | RhRh | 0.58 | 46 | 0.211
FDA | RhRh | 0.58 | 47 | 0.211
FDA | RhRh | 0.58 | 48 | 0.212
FDA | RhRh | 0.58 | 49 | 0.212
FDA | RhRh | 0.6 | 30 | 0.209
FDA | RhRh | 0.6 | 31 | 0.209
FDA | RhRh | 0.6 | 32 | 0.21
FDA | RhRh | 0.6 | 33 | 0.21
FDA | RhRh | 0.6 | 34 | 0.21
FDA | RhRh | 0.6 | 35 | 0.21
FDA | RhRh | 0.6 | 36 | 0.211
FDA | RhRh | 0.6 | 37 | 0.211
FDA | RhRh | 0.6 | 38 | 0.212
FDA | RhRh | 0.6 | 39 | 0.212
FDA | RhRh | 0.6 | 40 | 0.213
FDA | RhRh | 0.6 | 41 | 0.213
FDA | RhRh | 0.6 | 42 | 0.214
FDA | RhRh | 0.6 | 43 | 0.214
FDA | RhRh | 0.6 | 44 | 0.215
FDA | RhRh | 0.6 | 45 | 0.215
FDA | RhRh | 0.6 | 46 | 0.216
FDA | RhRh | 0.6 | 47 | 0.216
FDA | RhRh | 0.6 | 48 | 0.217
FDA | RhRh | 0.6 | 49 | 0.217
FDA | RhRh | 0.62 | 33 | 0.216
FDA | RhRh | 0.62 | 34 | 0.216
FDA | RhRh | 0.62 | 35 | 0.216
FDA | RhRh | 0.62 | 36 | 0.217
FDA | RhRh | 0.62 | 37 | 0.217
FDA | RhRh | 0.62 | 38 | 0.217
FDA | RhRh | 0.62 | 39 | 0.218
FDA | RhRh | 0.62 | 40 | 0.218
FDA | RhRh | 0.62 | 41 | 0.219
FDA | RhRh | 0.62 | 42 | 0.219
FDA | RhRh | 0.62 | 43 | 0.22
FDA | RhRh | 0.62 | 44 | 0.22
FDA | RhRh | 0.62 | 45 | 0.221
FDA | RhRh | 0.62 | 46 | 0.221
FDA | RhRh | 0.62 | 47 | 0.222
FDA | RhRh | 0.62 | 48 | 0.222
FDA | RhRh | 0.62 | 49 | 0.223
FDA | RhRh | 0.64 | 35 | 0.222
FDA | RhRh | 0.64 | 36 | 0.222
FDA | RhRh | 0.64 | 37 | 0.223
FDA | RhRh | 0.64 | 38 | 0.223
FDA | RhRh | 0.64 | 39 | 0.223
FDA | RhRh | 0.64 | 40 | 0.224
FDA | RhRh | 0.64 | 41 | 0.224
FDA | RhRh | 0.64 | 42 | 0.225
FDA | RhRh | 0.64 | 43 | 0.225
FDA | RhRh | 0.64 | 44 | 0.226
FDA | RhRh | 0.64 | 45 | 0.226
FDA | RhRh | 0.64 | 46 | 0.227
FDA | RhRh | 0.64 | 47 | 0.227
FDA | RhRh | 0.64 | 48 | 0.227
FDA | RhRh | 0.64 | 49 | 0.228
FDA | RhRh | 0.66 | 38 | 0.229
FDA | RhRh | 0.66 | 39 | 0.229
FDA | RhRh | 0.66 | 40 | 0.229
FDA | RhRh | 0.66 | 41 | 0.23
FDA | RhRh | 0.66 | 42 | 0.23
FDA | RhRh | 0.66 | 43 | 0.231
FDA | RhRh | 0.66 | 44 | 0.231
FDA | RhRh | 0.66 | 45 | 0.231
FDA | RhRh | 0.66 | 46 | 0.232
FDA | RhRh | 0.66 | 47 | 0.232
FDA | RhRh | 0.66 | 48 | 0.233
FDA | RhRh | 0.66 | 49 | 0.233
FDA | RhRh | 0.68 | 40 | 0.235
FDA | RhRh | 0.68 | 41 | 0.235
FDA | RhRh | 0.68 | 42 | 0.236
FDA | RhRh | 0.68 | 43 | 0.236
FDA | RhRh | 0.68 | 44 | 0.236
FDA | RhRh | 0.68 | 45 | 0.237
FDA | RhRh | 0.68 | 46 | 0.237
FDA | RhRh | 0.68 | 47 | 0.238
FDA | RhRh | 0.68 | 48 | 0.238
FDA | RhRh | 0.68 | 49 | 0.239
FDA | RhRh | 0.7 | 43 | 0.242
FDA | RhRh | 0.7 | 44 | 0.242
FDA | RhRh | 0.7 | 45 | 0.242
FDA | RhRh | 0.7 | 46 | 0.243
FDA | RhRh | 0.7 | 47 | 0.243
FDA | RhRh | 0.7 | 48 | 0.244
FDA | RhRh | 0.7 | 49 | 0.244
FDA | RhRh | 0.72 | 47 | 0.249
FDA | RhRh | 0.72 | 48 | 0.249
FDA | RhRh | 0.72 | 49 | 0.25
FDA | RhAg | 0.2 | 20 | 0.046
FDA | RhAg | 0.2 | 21 | 0.049
FDA | RhAg | 0.2 | 22 | 0.051
FDA | RhAg | 0.2 | 23 | 0.053
FDA | RhAg | 0.2 | 24 | 0.056
FDA | RhAg | 0.2 | 25 | 0.059
FDA | RhAg | 0.22 | 20 | 0.053
FDA | RhAg | 0.22 | 21 | 0.057
FDA | RhAg | 0.22 | 22 | 0.059
FDA | RhAg | 0.22 | 23 | 0.061
FDA | RhAg | 0.22 | 24 | 0.065
FDA | RhAg | 0.22 | 25 | 0.068
FDA | RhAg | 0.22 | 26 | 0.071
FDA | RhAg | 0.22 | 27 | 0.073
FDA | RhAg | 0.24 | 20 | 0.061
FDA | RhAg | 0.24 | 21 | 0.064
FDA | RhAg | 0.24 | 22 | 0.067
FDA | RhAg | 0.24 | 23 | 0.069
FDA | RhAg | 0.24 | 24 | 0.073
FDA | RhAg | 0.24 | 25 | 0.076
FDA | RhAg | 0.24 | 26 | 0.08
FDA | RhAg | 0.24 | 27 | 0.082
FDA | RhAg | 0.24 | 28 | 0.085
FDA | RhAg | 0.24 | 29 | 0.087
FDA | RhAg | 0.26 | 20 | 0.068
FDA | RhAg | 0.26 | 21 | 0.072
FDA | RhAg | 0.26 | 22 | 0.074
FDA | RhAg | 0.26 | 23 | 0.077
FDA | RhAg | 0.26 | 24 | 0.081
FDA | RhAg | 0.26 | 25 | 0.085
FDA | RhAg | 0.26 | 26 | 0.088
FDA | RhAg | 0.26 | 27 | 0.091
FDA | RhAg | 0.26 | 28 | 0.093
FDA | RhAg | 0.26 | 29 | 0.095
FDA | RhAg | 0.26 | 30 | 0.097
FDA | RhAg | 0.26 | 31 | 0.099
FDA | RhAg | 0.26 | 32 | 0.1
FDA | RhAg | 0.28 | 20 | 0.075
FDA | RhAg | 0.28 | 21 | 0.079
FDA | RhAg | 0.28 | 22 | 0.082
FDA | RhAg | 0.28 | 23 | 0.085
FDA | RhAg | 0.28 | 24 | 0.089
FDA | RhAg | 0.28 | 25 | 0.093
FDA | RhAg | 0.28 | 26 | 0.096
FDA | RhAg | 0.28 | 27 | 0.099
FDA | RhAg | 0.28 | 28 | 0.101
FDA | RhAg | 0.28 | 29 | 0.103
FDA | RhAg | 0.28 | 30 | 0.105
FDA | RhAg | 0.28 | 31 | 0.107
FDA | RhAg | 0.28 | 32 | 0.108
FDA | RhAg | 0.28 | 33 | 0.11
FDA | RhAg | 0.28 | 34 | 0.111
FDA | RhAg | 0.3 | 20 | 0.082
FDA | RhAg | 0.3 | 21 | 0.086
FDA | RhAg | 0.3 | 22 | 0.089
FDA | RhAg | 0.3 | 23 | 0.093
FDA | RhAg | 0.3 | 24 | 0.097
FDA | RhAg | 0.3 | 25 | 0.101
FDA | RhAg | 0.3 | 26 | 0.104
FDA | RhAg | 0.3 | 27 | 0.107
FDA | RhAg | 0.3 | 28 | 0.109
FDA | RhAg | 0.3 | 29 | 0.111
FDA | RhAg | 0.3 | 30 | 0.113
FDA | RhAg | 0.3 | 31 | 0.115
FDA | RhAg | 0.3 | 32 | 0.116
FDA | RhAg | 0.3 | 33 | 0.117
FDA | RhAg | 0.3 | 34 | 0.119
FDA | RhAg | 0.3 | 35 | 0.12
FDA | RhAg | 0.3 | 36 | 0.121
FDA | RhAg | 0.3 | 37 | 0.122
FDA | RhAg | 0.32 | 20 | 0.089
FDA | RhAg | 0.32 | 21 | 0.093
FDA | RhAg | 0.32 | 22 | 0.097
FDA | RhAg | 0.32 | 23 | 0.1
FDA | RhAg | 0.32 | 24 | 0.105
FDA | RhAg | 0.32 | 25 | 0.108
FDA | RhAg | 0.32 | 26 | 0.112
FDA | RhAg | 0.32 | 27 | 0.115
FDA | RhAg | 0.32 | 28 | 0.117
FDA | RhAg | 0.32 | 29 | 0.119
FDA | RhAg | 0.32 | 30 | 0.121
FDA | RhAg | 0.32 | 31 | 0.122
FDA | RhAg | 0.32 | 32 | 0.124
FDA | RhAg | 0.32 | 33 | 0.125
FDA | RhAg | 0.32 | 34 | 0.126
FDA | RhAg | 0.32 | 35 | 0.127
FDA | RhAg | 0.32 | 36 | 0.128
FDA | RhAg | 0.32 | 37 | 0.129
FDA | RhAg | 0.32 | 38 | 0.13
FDA | RhAg | 0.32 | 39 | 0.131
FDA | RhAg | 0.32 | 40 | 0.132
FDA | RhAg | 0.32 | 41 | 0.133
FDA | RhAg | 0.32 | 42 | 0.134
FDA | RhAg | 0.32 | 43 | 0.135
FDA | RhAg | 0.34 | 20 | 0.096
FDA | RhAg | 0.34 | 21 | 0.101
FDA | RhAg | 0.34 | 22 | 0.104
FDA | RhAg | 0.34 | 23 | 0.108
FDA | RhAg | 0.34 | 24 | 0.112
FDA | RhAg | 0.34 | 25 | 0.116
FDA | RhAg | 0.34 | 26 | 0.119
FDA | RhAg | 0.34 | 27 | 0.122
FDA | RhAg | 0.34 | 28 | 0.125
FDA | RhAg | 0.34 | 29 | 0.126
FDA | RhAg | 0.34 | 30 | 0.128
FDA | RhAg | 0.34 | 31 | 0.129
FDA | RhAg | 0.34 | 32 | 0.131
FDA | RhAg | 0.34 | 33 | 0.132
FDA | RhAg | 0.34 | 34 | 0.133
FDA | RhAg | 0.34 | 35 | 0.134
FDA | RhAg | 0.34 | 36 | 0.135
FDA | RhAg | 0.34 | 37 | 0.136
FDA | RhAg | 0.34 | 38 | 0.137
FDA | RhAg | 0.34 | 39 | 0.138
FDA | RhAg | 0.34 | 40 | 0.139
FDA | RhAg | 0.34 | 41 | 0.14
FDA | RhAg | 0.34 | 42 | 0.141
FDA | RhAg | 0.34 | 43 | 0.142
FDA | RhAg | 0.34 | 44 | 0.142
FDA | RhAg | 0.34 | 45 | 0.143
FDA | RhAg | 0.34 | 46 | 0.144
FDA | RhAg | 0.36 | 20 | 0.104
FDA | RhAg | 0.36 | 21 | 0.108
FDA | RhAg | 0.36 | 22 | 0.111
FDA | RhAg | 0.36 | 23 | 0.115
FDA | RhAg | 0.36 | 24 | 0.12
FDA | RhAg | 0.36 | 25 | 0.123
FDA | RhAg | 0.36 | 26 | 0.127
FDA | RhAg | 0.36 | 27 | 0.13
FDA | RhAg | 0.36 | 28 | 0.132
FDA | RhAg | 0.36 | 29 | 0.134
FDA | RhAg | 0.36 | 30 | 0.135
FDA | RhAg | 0.36 | 31 | 0.137
FDA | RhAg | 0.36 | 32 | 0.138
FDA | RhAg | 0.36 | 33 | 0.139
FDA | RhAg | 0.36 | 34 | 0.14
FDA | RhAg | 0.36 | 35 | 0.141
FDA | RhAg | 0.36 | 36 | 0.142
FDA | RhAg | 0.36 | 37 | 0.143
FDA | RhAg | 0.36 | 38 | 0.144
FDA | RhAg | 0.36 | 39 | 0.145
FDA | RhAg | 0.36 | 40 | 0.146
FDA | RhAg | 0.36 | 41 | 0.147
FDA | RhAg | 0.36 | 42 | 0.148
FDA | RhAg | 0.36 | 43 | 0.148
FDA | RhAg | 0.36 | 44 | 0.149
FDA | RhAg | 0.36 | 45 | 0.15
FDA | RhAg | 0.36 | 46 | 0.15
FDA | RhAg | 0.36 | 47 | 0.151
FDA | RhAg | 0.36 | 48 | 0.151
FDA | RhAg | 0.36 | 49 | 0.152
FDA | RhAg | 0.38 | 20 | 0.111
FDA | RhAg | 0.38 | 21 | 0.116
FDA | RhAg | 0.38 | 22 | 0.119
FDA | RhAg | 0.38 | 23 | 0.122
FDA | RhAg | 0.38 | 24 | 0.127
FDA | RhAg | 0.38 | 25 | 0.131
FDA | RhAg | 0.38 | 26 | 0.134
FDA | RhAg | 0.38 | 27 | 0.137
FDA | RhAg | 0.38 | 28 | 0.139
FDA | RhAg | 0.38 | 29 | 0.141
FDA | RhAg | 0.38 | 30 | 0.142
FDA | RhAg | 0.38 | 31 | 0.144
FDA | RhAg | 0.38 | 32 | 0.145
FDA | RhAg | 0.38 | 33 | 0.146
FDA | RhAg | 0.38 | 34 | 0.147
FDA | RhAg | 0.38 | 35 | 0.148
FDA | RhAg | 0.38 | 36 | 0.149
FDA | RhAg | 0.38 | 37 | 0.15
FDA | RhAg | 0.38 | 38 | 0.151
FDA | RhAg | 0.38 | 39 | 0.152
FDA | RhAg | 0.38 | 40 | 0.152
FDA | RhAg | 0.38 | 41 | 0.153
FDA | RhAg | 0.38 | 42 | 0.154
FDA | RhAg | 0.38 | 43 | 0.155
FDA | RhAg | 0.38 | 44 | 0.155
FDA | RhAg | 0.38 | 45 | 0.156
FDA | RhAg | 0.38 | 46 | 0.157
FDA | RhAg | 0.38 | 47 | 0.157
FDA | RhAg | 0.38 | 48 | 0.158
FDA | RhAg | 0.38 | 49 | 0.158
FDA | RhAg | 0.4 | 20 | 0.119
FDA | RhAg | 0.4 | 21 | 0.123
FDA | RhAg | 0.4 | 22 | 0.126
FDA | RhAg | 0.4 | 23 | 0.13
FDA | RhAg | 0.4 | 24 | 0.134
FDA | RhAg | 0.4 | 25 | 0.138
FDA | RhAg | 0.4 | 26 | 0.141
FDA | RhAg | 0.4 | 27 | 0.144
FDA | RhAg | 0.4 | 28 | 0.146
FDA | RhAg | 0.4 | 29 | 0.148
FDA | RhAg | 0.4 | 30 | 0.149
FDA | RhAg | 0.4 | 31 | 0.15
FDA | RhAg | 0.4 | 32 | 0.151
FDA | RhAg | 0.4 | 33 | 0.152
FDA | RhAg | 0.4 | 34 | 0.153
FDA | RhAg | 0.4 | 35 | 0.154
FDA | RhAg | 0.4 | 36 | 0.155
FDA | RhAg | 0.4 | 37 | 0.156
FDA | RhAg | 0.4 | 38 | 0.157
FDA | RhAg | 0.4 | 39 | 0.158
FDA | RhAg | 0.4 | 40 | 0.159
FDA | RhAg | 0.4 | 41 | 0.159
FDA | RhAg | 0.4 | 42 | 0.16
FDA | RhAg | 0.4 | 43 | 0.161
FDA | RhAg | 0.4 | 44 | 0.161
FDA | RhAg | 0.4 | 45 | 0.162
FDA | RhAg | 0.4 | 46 | 0.163
FDA | RhAg | 0.4 | 47 | 0.163
FDA | RhAg | 0.4 | 48 | 0.164
FDA | RhAg | 0.4 | 49 | 0.164
FDA | RhAg | 0.42 | 21 | 0.13
FDA | RhAg | 0.42 | 22 | 0.134
FDA | RhAg | 0.42 | 23 | 0.137
FDA | RhAg | 0.42 | 24 | 0.142
FDA | RhAg | 0.42 | 25 | 0.145
FDA | RhAg | 0.42 | 26 | 0.148
FDA | RhAg | 0.42 | 27 | 0.151
FDA | RhAg | 0.42 | 28 | 0.153
FDA | RhAg | 0.42 | 29 | 0.155
FDA | RhAg | 0.42 | 30 | 0.156
FDA | RhAg | 0.42 | 31 | 0.157
FDA | RhAg | 0.42 | 32 | 0.158
FDA | RhAg | 0.42 | 33 | 0.159
FDA | RhAg | 0.42 | 34 | 0.16
FDA | RhAg | 0.42 | 35 | 0.161
FDA | RhAg | 0.42 | 36 | 0.161
FDA | RhAg | 0.42 | 37 | 0.162
FDA | RhAg | 0.42 | 38 | 0.163
FDA | RhAg | 0.42 | 39 | 0.164
FDA | RhAg | 0.42 | 40 | 0.165
FDA | RhAg | 0.42 | 41 | 0.165
FDA | RhAg | 0.42 | 42 | 0.166
FDA | RhAg | 0.42 | 43 | 0.167
FDA | RhAg | 0.42 | 44 | 0.167
FDA | RhAg | 0.42 | 45 | 0.168
FDA | RhAg | 0.42 | 46 | 0.169
FDA | RhAg | 0.42 | 47 | 0.169
FDA | RhAg | 0.42 | 48 | 0.17
FDA | RhAg | 0.42 | 49 | 0.17
FDA | RhAg | 0.44 | 21 | 0.138
FDA | RhAg | 0.44 | 22 | 0.141
FDA | RhAg | 0.44 | 23 | 0.145
FDA | RhAg | 0.44 | 24 | 0.149
FDA | RhAg | 0.44 | 25 | 0.152
FDA | RhAg | 0.44 | 26 | 0.156
FDA | RhAg | 0.44 | 27 | 0.158
FDA | RhAg | 0.44 | 28 | 0.16
FDA | RhAg | 0.44 | 29 | 0.161
FDA | RhAg | 0.44 | 30 | 0.162
FDA | RhAg | 0.44 | 31 | 0.163
FDA | RhAg | 0.44 | 32 | 0.164
FDA | RhAg | 0.44 | 33 | 0.165
FDA | RhAg | 0.44 | 34 | 0.166
FDA | RhAg | 0.44 | 35 | 0.167
FDA | RhAg | 0.44 | 36 | 0.168
FDA | RhAg | 0.44 | 37 | 0.168
FDA | RhAg | 0.44 | 38 | 0.169
FDA | RhAg | 0.44 | 39 | 0.17
FDA | RhAg | 0.44 | 40 | 0.171
FDA | RhAg | 0.44 | 41 | 0.171
FDA | RhAg | 0.44 | 42 | 0.172
FDA | RhAg | 0.44 | 43 | 0.173
FDA | RhAg | 0.44 | 44 | 0.173
FDA | RhAg | 0.44 | 45 | 0.174
FDA | RhAg | 0.44 | 46 | 0.175
FDA | RhAg | 0.44 | 47 | 0.175
FDA | RhAg | 0.44 | 48 | 0.176
FDA | RhAg | 0.44 | 49 | 0.176
FDA | RhAg | 0.46 | 22 | 0.148
FDA | RhAg | 0.46 | 23 | 0.152
FDA | RhAg | 0.46 | 24 | 0.156
FDA | RhAg | 0.46 | 25 | 0.159
FDA | RhAg | 0.46 | 26 | 0.162
FDA | RhAg | 0.46 | 27 | 0.165
FDA | RhAg | 0.46 | 28 | 0.167
FDA | RhAg | 0.46 | 29 | 0.168
FDA | RhAg | 0.46 | 30 | 0.169
FDA | RhAg | 0.46 | 31 | 0.17
FDA | RhAg | 0.46 | 32 | 0.171
FDA | RhAg | 0.46 | 33 | 0.171
FDA | RhAg | 0.46 | 34 | 0.172
FDA | RhAg | 0.46 | 35 | 0.173
FDA | RhAg | 0.46 | 36 | 0.174
FDA | RhAg | 0.46 | 37 | 0.174
FDA | RhAg | 0.46 | 38 | 0.175
FDA | RhAg | 0.46 | 39 | 0.176
FDA | RhAg | 0.46 | 40 | 0.177
FDA | RhAg | 0.46 | 41 | 0.177
FDA | RhAg | 0.46 | 42 | 0.178
FDA | RhAg | 0.46 | 43 | 0.179
FDA | RhAg | 0.46 | 44 | 0.179
FDA | RhAg | 0.46 | 45 | 0.18
FDA | RhAg | 0.46 | 46 | 0.18
FDA | RhAg | 0.46 | 47 | 0.181
FDA | RhAg | 0.46 | 48 | 0.181
FDA | RhAg | 0.46 | 49 | 0.182
FDA | RhAg | 0.48 | 22 | 0.156
FDA | RhAg | 0.48 | 23 | 0.159
FDA | RhAg | 0.48 | 24 | 0.163
FDA | RhAg | 0.48 | 25 | 0.166
FDA | RhAg | 0.48 | 26 | 0.169
FDA | RhAg | 0.48 | 27 | 0.172
FDA | RhAg | 0.48 | 28 | 0.173
FDA | RhAg | 0.48 | 29 | 0.174
FDA | RhAg | 0.48 | 30 | 0.175
FDA | RhAg | 0.48 | 31 | 0.176
FDA | RhAg | 0.48 | 32 | 0.177
FDA | RhAg | 0.48 | 33 | 0.178
FDA | RhAg | 0.48 | 34 | 0.178
FDA | RhAg | 0.48 | 35 | 0.179
FDA | RhAg | 0.48 | 36 | 0.18
FDA | RhAg | 0.48 | 37 | 0.18
FDA | RhAg | 0.48 | 38 | 0.181
FDA | RhAg | 0.48 | 39 | 0.182
FDA | RhAg | 0.48 | 40 | 0.182
FDA | RhAg | 0.48 | 41 | 0.183
FDA | RhAg | 0.48 | 42 | 0.184
FDA | RhAg | 0.48 | 43 | 0.184
FDA | RhAg | 0.48 | 44 | 0.185
FDA | RhAg | 0.48 | 45 | 0.185
FDA | RhAg | 0.48 | 46 | 0.186
FDA | RhAg | 0.48 | 47 | 0.187
FDA | RhAg | 0.48 | 48 | 0.187
FDA | RhAg | 0.48 | 49 | 0.188
FDA | RhAg | 0.5 | 23 | 0.167
FDA | RhAg | 0.5 | 24 | 0.17
FDA | RhAg | 0.5 | 25 | 0.173
FDA | RhAg | 0.5 | 26 | 0.176
FDA | RhAg | 0.5 | 27 | 0.178
FDA | RhAg | 0.5 | 28 | 0.18
FDA | RhAg | 0.5 | 29 | 0.181
FDA | RhAg | 0.5 | 30 | 0.182
FDA | RhAg | 0.5 | 31 | 0.183
FDA | RhAg | 0.5 | 32 | 0.183
FDA | RhAg | 0.5 | 33 | 0.184
FDA | RhAg | 0.5 | 34 | 0.184
FDA | RhAg | 0.5 | 35 | 0.185
FDA | RhAg | 0.5 | 36 | 0.186
FDA | RhAg | 0.5 | 37 | 0.186
FDA | RhAg | 0.5 | 38 | 0.187
FDA | RhAg | 0.5 | 39 | 0.188
FDA | RhAg | 0.5 | 40 | 0.188
FDA | RhAg | 0.5 | 41 | 0.189
FDA | RhAg | 0.5 | 42 | 0.189
FDA | RhAg | 0.5 | 43 | 0.19
FDA | RhAg | 0.5 | 44 | 0.191
FDA | RhAg | 0.5 | 45 | 0.191
FDA | RhAg | 0.5 | 46 | 0.192
FDA | RhAg | 0.5 | 47 | 0.192
FDA | RhAg | 0.5 | 48 | 0.193
FDA | RhAg | 0.5 | 49 | 0.193
FDA | RhAg | 0.52 | 23 | 0.174
FDA | RhAg | 0.52 | 24 | 0.177
FDA | RhAg | 0.52 | 25 | 0.18
FDA | RhAg | 0.52 | 26 | 0.183
FDA | RhAg | 0.52 | 27 | 0.185
FDA | RhAg | 0.52 | 28 | 0.187
FDA | RhAg | 0.52 | 29 | 0.187
FDA | RhAg | 0.52 | 30 | 0.188
FDA | RhAg | 0.52 | 31 | 0.189
FDA | RhAg | 0.52 | 32 | 0.189
FDA | RhAg | 0.52 | 33 | 0.19
FDA | RhAg | 0.52 | 34 | 0.19
FDA | RhAg | 0.52 | 35 | 0.191
FDA | RhAg | 0.52 | 36 | 0.192
FDA | RhAg | 0.52 | 37 | 0.192
FDA | RhAg | 0.52 | 38 | 0.193
FDA | RhAg | 0.52 | 39 | 0.193
FDA | RhAg | 0.52 | 40 | 0.194
FDA | RhAg | 0.52 | 41 | 0.194
FDA | RhAg | 0.52 | 42 | 0.195
FDA | RhAg | 0.52 | 43 | 0.196
FDA | RhAg | 0.52 | 44 | 0.196
FDA | RhAg | 0.52 | 45 | 0.197
FDA | RhAg | 0.52 | 46 | 0.197
FDA | RhAg | 0.52 | 47 | 0.198
FDA | RhAg | 0.52 | 48 | 0.198
FDA | RhAg | 0.52 | 49 | 0.199
FDA | RhAg | 0.54 | 24 | 0.184
FDA | RhAg | 0.54 | 25 | 0.187
FDA | RhAg | 0.54 | 26 | 0.19
FDA | RhAg | 0.54 | 27 | 0.192
FDA | RhAg | 0.54 | 28 | 0.193
FDA | RhAg | 0.54 | 29 | 0.194
FDA | RhAg | 0.54 | 30 | 0.194
FDA | RhAg | 0.54 | 31 | 0.195
FDA | RhAg | 0.54 | 32 | 0.196
FDA | RhAg | 0.54 | 33 | 0.196
FDA | RhAg | 0.54 | 34 | 0.196
FDA | RhAg | 0.54 | 35 | 0.197
FDA | RhAg | 0.54 | 36 | 0.197
FDA | RhAg | 0.54 | 37 | 0.198
FDA | RhAg | 0.54 | 38 | 0.198
FDA | RhAg | 0.54 | 39 | 0.199
FDA | RhAg | 0.54 | 40 | 0.199
FDA | RhAg | 0.54 | 41 | 0.2
FDA | RhAg | 0.54 | 42 | 0.201
FDA | RhAg | 0.54 | 43 | 0.201
FDA | RhAg | 0.54 | 44 | 0.202
FDA | RhAg | 0.54 | 45 | 0.202
FDA | RhAg | 0.54 | 46 | 0.203
FDA | RhAg | 0.54 | 47 | 0.203
FDA | RhAg | 0.54 | 48 | 0.204
FDA | RhAg | 0.54 | 49 | 0.204
FDA | RhAg | 0.56 | 26 | 0.196
FDA | RhAg | 0.56 | 27 | 0.198
FDA | RhAg | 0.56 | 28 | 0.199
FDA | RhAg | 0.56 | 29 | 0.2
FDA | RhAg | 0.56 | 30 | 0.201
FDA | RhAg | 0.56 | 31 | 0.201
FDA | RhAg | 0.56 | 32 | 0.202
FDA | RhAg | 0.56 | 33 | 0.202
FDA | RhAg | 0.56 | 34 | 0.202
FDA | RhAg | 0.56 | 35 | 0.203
FDA | RhAg | 0.56 | 36 | 0.203
FDA | RhAg | 0.56 | 37 | 0.204
FDA | RhAg | 0.56 | 38 | 0.204
FDA | RhAg | 0.56 | 39 | 0.205
FDA | RhAg | 0.56 | 40 | 0.205
FDA | RhAg | 0.56 | 41 | 0.206
FDA | RhAg | 0.56 | 42 | 0.206
FDA | RhAg | 0.56 | 43 | 0.207
FDA | RhAg | 0.56 | 44 | 0.207
FDA | RhAg | 0.56 | 45 | 0.208
FDA | RhAg | 0.56 | 46 | 0.208
FDA | RhAg | 0.56 | 47 | 0.208
FDA | RhAg | 0.56 | 48 | 0.209
FDA | RhAg | 0.56 | 49 | 0.209
FDA | RhAg | 0.58 | 28 | 0.206
FDA | RhAg | 0.58 | 29 | 0.207
FDA | RhAg | 0.58 | 30 | 0.207
FDA | RhAg | 0.58 | 31 | 0.207
FDA | RhAg | 0.58 | 32 | 0.208
FDA | RhAg | 0.58 | 33 | 0.208
FDA | RhAg | 0.58 | 34 | 0.208
FDA | RhAg | 0.58 | 35 | 0.209
FDA | RhAg | 0.58 | 36 | 0.209
FDA | RhAg | 0.58 | 37 | 0.209
FDA | RhAg | 0.58 | 38 | 0.21
FDA | RhAg | 0.58 | 39 | 0.21
FDA | RhAg | 0.58 | 40 | 0.211
FDA | RhAg | 0.58 | 41 | 0.211
FDA | RhAg | 0.58 | 42 | 0.212
FDA | RhAg | 0.58 | 43 | 0.212
FDA | RhAg | 0.58 | 44 | 0.213
FDA | RhAg | 0.58 | 45 | 0.213
FDA | RhAg | 0.58 | 46 | 0.214
FDA | RhAg | 0.58 | 47 | 0.214
FDA | RhAg | 0.58 | 48 | 0.214
FDA | RhAg | 0.58 | 49 | 0.215
FDA | RhAg | 0.6 | 30 | 0.213
FDA | RhAg | 0.6 | 31 | 0.214
FDA | RhAg | 0.6 | 32 | 0.214
FDA | RhAg | 0.6 | 33 | 0.214
FDA | RhAg | 0.6 | 34 | 0.214
FDA | RhAg | 0.6 | 35 | 0.214
FDA | RhAg | 0.6 | 36 | 0.215
FDA | RhAg | 0.6 | 37 | 0.215
FDA | RhAg | 0.6 | 38 | 0.216
FDA | RhAg | 0.6 | 39 | 0.216
FDA | RhAg | 0.6 | 40 | 0.216
FDA | RhAg | 0.6 | 41 | 0.217
FDA | RhAg | 0.6 | 42 | 0.217
FDA | RhAg | 0.6 | 43 | 0.218
FDA | RhAg | 0.6 | 44 | 0.218
FDA | RhAg | 0.6 | 45 | 0.218
FDA | RhAg | 0.6 | 46 | 0.219
FDA | RhAg | 0.6 | 47 | 0.219
FDA | RhAg | 0.6 | 48 | 0.22
FDA | RhAg | 0.6 | 49 | 0.22
FDA | RhAg | 0.62 | 33 | 0.22
FDA | RhAg | 0.62 | 34 | 0.22
FDA | RhAg | 0.62 | 35 | 0.22
FDA | RhAg | 0.62 | 36 | 0.221
FDA | RhAg | 0.62 | 37 | 0.221
FDA | RhAg | 0.62 | 38 | 0.221
FDA | RhAg | 0.62 | 39 | 0.222
FDA | RhAg | 0.62 | 40 | 0.222
FDA | RhAg | 0.62 | 41 | 0.222
FDA | RhAg | 0.62 | 42 | 0.223
FDA | RhAg | 0.62 | 43 | 0.223
FDA | RhAg | 0.62 | 44 | 0.223
FDA | RhAg | 0.62 | 45 | 0.224
FDA | RhAg | 0.62 | 46 | 0.224
FDA | RhAg | 0.62 | 47 | 0.225
FDA | RhAg | 0.62 | 48 | 0.225
FDA | RhAg | 0.62 | 49 | 0.226
FDA | RhAg | 0.64 | 35 | 0.226
FDA | RhAg | 0.64 | 36 | 0.226
FDA | RhAg | 0.64 | 37 | 0.227
FDA | RhAg | 0.64 | 38 | 0.227
FDA | RhAg | 0.64 | 39 | 0.227
FDA | RhAg | 0.64 | 40 | 0.227
FDA | RhAg | 0.64 | 41 | 0.228
FDA | RhAg | 0.64 | 42 | 0.228
FDA | RhAg | 0.64 | 43 | 0.229
FDA | RhAg | 0.64 | 44 | 0.229
FDA | RhAg | 0.64 | 45 | 0.229
FDA | RhAg | 0.64 | 46 | 0.23
FDA | RhAg | 0.64 | 47 | 0.23
FDA | RhAg | 0.64 | 48 | 0.231
FDA | RhAg | 0.64 | 49 | 0.231
FDA | RhAg | 0.66 | 38 | 0.232
FDA | RhAg | 0.66 | 39 | 0.233
FDA | RhAg | 0.66 | 40 | 0.233
FDA | RhAg | 0.66 | 41 | 0.233
FDA | RhAg | 0.66 | 42 | 0.234
FDA | RhAg | 0.66 | 43 | 0.234
FDA | RhAg | 0.66 | 44 | 0.234
FDA | RhAg | 0.66 | 45 | 0.235
FDA | RhAg | 0.66 | 46 | 0.235
FDA | RhAg | 0.66 | 47 | 0.236
FDA | RhAg | 0.66 | 48 | 0.236
FDA | RhAg | 0.66 | 49 | 0.236
FDA | RhAg | 0.68 | 44 | 0.24
FDA | RhAg | 0.68 | 45 | 0.24
FDA | RhAg | 0.68 | 46 | 0.241
FDA | RhAg | 0.68 | 47 | 0.241
FDA | RhAg | 0.68 | 48 | 0.241
FDA | RhAg | 0.68 | 49 | 0.242
FDA | RhAg | 0.7 | 47 | 0.246
FDA | RhAg | 0.7 | 48 | 0.247
FDA | RhAg | 0.7 | 49 | 0.247
FDA | WAg | 0.2 | 20 | 0.048
FDA | WAg | 0.2 | 21 | 0.051
FDA | WAg | 0.2 | 22 | 0.054
FDA | WAg | 0.2 | 23 | 0.058
FDA | WAg | 0.2 | 24 | 0.061
FDA | WAg | 0.2 | 25 | 0.063
FDA | WAg | 0.2 | 26 | 0.066
FDA | WAg | 0.2 | 27 | 0.068
FDA | WAg | 0.225 | 20 | 0.057
FDA | WAg | 0.225 | 21 | 0.061
FDA | WAg | 0.225 | 22 | 0.064
FDA | WAg | 0.225 | 23 | 0.067
FDA | WAg | 0.225 | 24 | 0.071
FDA | WAg | 0.225 | 25 | 0.074
FDA | WAg | 0.225 | 26 | 0.077
FDA | WAg | 0.225 | 27 | 0.079
FDA | WAg | 0.25 | 20 | 0.065
FDA | WAg | 0.25 | 21 | 0.069
FDA | WAg | 0.25 | 22 | 0.073
FDA | WAg | 0.25 | 23 | 0.077
FDA | WAg | 0.25 | 24 | 0.08
FDA | WAg | 0.25 | 25 | 0.084
FDA | WAg | 0.25 | 26 | 0.087
FDA | WAg | 0.25 | 27 | 0.089
FDA | WAg | 0.25 | 28 | 0.092
FDA | WAg | 0.25 | 29 | 0.094
FDA | WAg | 0.25 | 30 | 0.096
FDA | WAg | 0.25 | 31 | 0.098
FDA | WAg | 0.275 | 20 | 0.074
FDA | WAg | 0.275 | 21 | 0.078
FDA | WAg | 0.275 | 22 | 0.082
FDA | WAg | 0.275 | 23 | 0.086
FDA | WAg | 0.275 | 24 | 0.09
FDA | WAg | 0.275 | 25 | 0.093
FDA | WAg | 0.275 | 26 | 0.096
FDA | WAg | 0.275 | 27 | 0.099
FDA | WAg | 0.275 | 28 | 0.102
FDA | WAg | 0.275 | 29 | 0.104
FDA | WAg | 0.275 | 30 | 0.106
FDA | WAg | 0.275 | 31 | 0.108
FDA | WAg | 0.275 | 32 | 0.11
FDA | WAg | 0.275 | 33 | 0.112
FDA | WAg | 0.275 | 34 | 0.114
FDA | WAg | 0.3 | 20 | 0.083
FDA | WAg | 0.3 | 21 | 0.087
FDA | WAg | 0.3 | 22 | 0.091
FDA | WAg | 0.3 | 23 | 0.095
FDA | WAg | 0.3 | 24 | 0.099
FDA | WAg | 0.3 | 25 | 0.102
FDA | WAg | 0.3 | 26 | 0.106
FDA | WAg | 0.3 | 27 | 0.109
FDA | WAg | 0.3 | 28 | 0.111
FDA | WAg | 0.3 | 29 | 0.113
FDA | WAg | 0.3 | 30 | 0.115
FDA | WAg | 0.3 | 31 | 0.117
FDA | WAg | 0.3 | 32 | 0.119
FDA | WAg | 0.3 | 33 | 0.121
FDA | WAg | 0.3 | 34 | 0.123
FDA | WAg | 0.3 | 35 | 0.124
FDA | WAg | 0.3 | 36 | 0.126
FDA | WAg | 0.3 | 37 | 0.128
FDA | WAg | 0.3 | 38 | 0.13
FDA | WAg | 0.3 | 39 | 0.131
FDA | WAg | 0.3 | 40 | 0.133
FDA | WAg | 0.3 | 41 | 0.135
FDA | WAg | 0.3 | 42 | 0.136
FDA | WAg | 0.3 | 43 | 0.138
FDA | WAg | 0.325 | 20 | 0.092
FDA | WAg | 0.325 | 21 | 0.096
FDA | WAg | 0.325 | 22 | 0.1
FDA | WAg | 0.325 | 23 | 0.104
FDA | WAg | 0.325 | 24 | 0.108
FDA | WAg | 0.325 | 25 | 0.111
FDA | WAg | 0.325 | 26 | 0.115
FDA | WAg | 0.325 | 27 | 0.118
FDA | WAg | 0.325 | 28 | 0.12
FDA | WAg | 0.325 | 29 | 0.122
FDA | WAg | 0.325 | 30 | 0.124
FDA | WAg | 0.325 | 31 | 0.126
FDA | WAg | 0.325 | 32 | 0.128
FDA | WAg | 0.325 | 33 | 0.13
FDA | WAg | 0.325 | 34 | 0.132
FDA | WAg | 0.325 | 35 | 0.133
FDA | WAg | 0.325 | 36 | 0.135
FDA | WAg | 0.325 | 37 | 0.137
FDA | WAg | 0.325 | 38 | 0.138
FDA | WAg | 0.325 | 39 | 0.14
FDA | WAg | 0.325 | 40 | 0.142
FDA | WAg | 0.325 | 41 | 0.143
FDA | WAg | 0.325 | 42 | 0.145
FDA | WAg | 0.325 | 43 | 0.146
FDA | WAg | 0.325 | 44 | 0.148
FDA | WAg | 0.325 | 45 | 0.149
FDA | WAg | 0.325 | 46 | 0.151
FDA | WAg | 0.325 | 47 | 0.152
FDA | WAg | 0.325 | 48 | 0.154
FDA | WAg | 0.325 | 49 | 0.155
FDA | WAg | 0.35 | 20 | 0.1
FDA | WAg | 0.35 | 21 | 0.105
FDA | WAg | 0.35 | 22 | 0.109
FDA | WAg | 0.35 | 23 | 0.113
FDA | WAg | 0.35 | 24 | 0.117
FDA | WAg | 0.35 | 25 | 0.121
FDA | WAg | 0.35 | 26 | 0.124
FDA | WAg | 0.35 | 27 | 0.127
FDA | WAg | 0.35 | 28 | 0.129
FDA | WAg | 0.35 | 29 | 0.131
FDA | WAg | 0.35 | 30 | 0.133
FDA | WAg | 0.35 | 31 | 0.135
FDA | WAg | 0.35 | 32 | 0.137
FDA | WAg | 0.35 | 33 | 0.138
FDA | WAg | 0.35 | 34 | 0.14
FDA | WAg | 0.35 | 35 | 0.142
FDA | WAg | 0.35 | 36 | 0.143
FDA | WAg | 0.35 | 37 | 0.145
FDA | WAg | 0.35 | 38 | 0.147
FDA | WAg | 0.35 | 39 | 0.148
FDA | WAg | 0.35 | 40 | 0.15
FDA | WAg | 0.35 | 41 | 0.152
FDA | WAg | 0.35 | 42 | 0.153
FDA | WAg | 0.35 | 43 | 0.155
FDA | WAg | 0.35 | 44 | 0.156
FDA | WAg | 0.35 | 45 | 0.158
FDA | WAg | 0.35 | 46 | 0.159
FDA | WAg | 0.35 | 47 | 0.161
FDA | WAg | 0.35 | 48 | 0.162
FDA | WAg | 0.35 | 49 | 0.164
FDA | WAg | 0.375 | 20 | 0.11
FDA | WAg | 0.375 | 21 | 0.114
FDA | WAg | 0.375 | 22 | 0.118
FDA | WAg | 0.375 | 23 | 0.122
FDA | WAg | 0.375 | 24 | 0.126
FDA | WAg | 0.375 | 25 | 0.129
FDA | WAg | 0.375 | 26 | 0.133
FDA | WAg | 0.375 | 27 | 0.136
FDA | WAg | 0.375 | 28 | 0.138
FDA | WAg | 0.375 | 29 | 0.14
FDA | WAg | 0.375 | 30 | 0.142
FDA | WAg | 0.375 | 31 | 0.144
FDA | WAg | 0.375 | 32 | 0.145
FDA | WAg | 0.375 | 33 | 0.147
FDA | WAg | 0.375 | 34 | 0.148
FDA | WAg | 0.375 | 35 | 0.15
FDA | WAg | 0.375 | 36 | 0.152
FDA | WAg | 0.375 | 37 | 0.153
FDA | WAg | 0.375 | 38 | 0.155
FDA | WAg | 0.375 | 39 | 0.156
FDA | WAg | 0.375 | 40 | 0.158
FDA | WAg | 0.375 | 41 | 0.16
FDA | WAg | 0.375 | 42 | 0.161
FDA | WAg | 0.375 | 43 | 0.163
FDA | WAg | 0.375 | 44 | 0.164
FDA | WAg | 0.375 | 45 | 0.166
FDA | WAg | 0.375 | 46 | 0.167
FDA | WAg | 0.375 | 47 | 0.169
FDA | WAg | 0.375 | 48 | 0.17
FDA | WAg | 0.375 | 49 | 0.172
FDA | WAg | 0.4 | 20 | 0.119
FDA | WAg | 0.4 | 21 | 0.123
FDA | WAg | 0.4 | 22 | 0.127
FDA | WAg | 0.4 | 23 | 0.131
FDA | WAg | 0.4 | 24 | 0.135
FDA | WAg | 0.4 | 25 | 0.138
FDA | WAg | 0.4 | 26 | 0.142
FDA | WAg | 0.4 | 27 | 0.145
FDA | WAg | 0.4 | 28 | 0.147
FDA | WAg | 0.4 | 29 | 0.149
FDA | WAg | 0.4 | 30 | 0.15
FDA | WAg | 0.4 | 31 | 0.152
FDA | WAg | 0.4 | 32 | 0.154
FDA | WAg | 0.4 | 33 | 0.155
FDA | WAg | 0.4 | 34 | 0.157
FDA | WAg | 0.4 | 35 | 0.158
FDA | WAg | 0.4 | 36 | 0.16
FDA | WAg | 0.4 | 37 | 0.161
FDA | WAg | 0.4 | 38 | 0.163
FDA | WAg | 0.4 | 39 | 0.164
FDA | WAg | 0.4 | 40 | 0.166
FDA | WAg | 0.4 | 41 | 0.167
FDA | WAg | 0.4 | 42 | 0.169
FDA | WAg | 0.4 | 43 | 0.17
FDA | WAg | 0.4 | 44 | 0.172
FDA | WAg | 0.4 | 45 | 0.173
FDA | WAg | 0.4 | 46 | 0.175
FDA | WAg | 0.4 | 47 | 0.176
FDA | WAg | 0.4 | 48 | 0.178
FDA | WAg | 0.4 | 49 | 0.179
FDA | WAg | 0.425 | 20 | 0.128
FDA | WAg | 0.425 | 21 | 0.132
FDA | WAg | 0.425 | 22 | 0.136
FDA | WAg | 0.425 | 23 | 0.14
FDA | WAg | 0.425 | 24 | 0.144
FDA | WAg | 0.425 | 25 | 0.147
FDA | WAg | 0.425 | 26 | 0.151
FDA | WAg | 0.425 | 27 | 0.153
FDA | WAg | 0.425 | 28 | 0.156
FDA | WAg | 0.425 | 29 | 0.157
FDA | WAg | 0.425 | 30 | 0.159
FDA | WAg | 0.425 | 31 | 0.16
FDA | WAg | 0.425 | 32 | 0.162
FDA | WAg | 0.425 | 33 | 0.163
FDA | WAg | 0.425 | 34 | 0.165
FDA | WAg | 0.425 | 35 | 0.166
FDA | WAg | 0.425 | 36 | 0.167
FDA | WAg | 0.425 | 37 | 0.169
FDA | WAg | 0.425 | 38 | 0.17
FDA | WAg | 0.425 | 39 | 0.172
FDA | WAg | 0.425 | 40 | 0.173
FDA | WAg | 0.425 | 41 | 0.175
FDA | WAg | 0.425 | 42 | 0.177
FDA | WAg | 0.425 | 43 | 0.178
FDA | WAg | 0.425 | 44 | 0.179
FDA | WAg | 0.425 | 45 | 0.181
FDA | WAg | 0.425 | 46 | 0.182
FDA | WAg | 0.425 | 47 | 0.184
FDA | WAg | 0.425 | 48 | 0.185
FDA | WAg | 0.425 | 49 | 0.187
FDA | WAg | 0.45 | 20 | 0.137
FDA | WAg | 0.45 | 21 | 0.141
FDA | WAg | 0.45 | 22 | 0.145
FDA | WAg | 0.45 | 23 | 0.149
FDA | WAg | 0.45 | 24 | 0.153
FDA | WAg | 0.45 | 25 | 0.156
FDA | WAg | 0.45 | 26 | 0.159
FDA | WAg | 0.45 | 27 | 0.162
FDA | WAg | 0.45 | 28 | 0.164
FDA | WAg | 0.45 | 29 | 0.166
FDA | WAg | 0.45 | 30 | 0.167
FDA | WAg | 0.45 | 31 | 0.169
FDA | WAg | 0.45 | 32 | 0.17
FDA | WAg | 0.45 | 33 | 0.171
FDA | WAg | 0.45 | 34 | 0.173
FDA | WAg | 0.45 | 35 | 0.174
FDA | WAg | 0.45 | 36 | 0.175
FDA | WAg | 0.45 | 37 | 0.177
FDA | WAg | 0.45 | 38 | 0.178
FDA | WAg | 0.45 | 39 | 0.18
FDA | WAg | 0.45 | 40 | 0.181
FDA | WAg | 0.45 | 41 | 0.182
FDA | WAg | 0.45 | 42 | 0.184
FDA | WAg | 0.45 | 43 | 0.185
FDA | WAg | 0.45 | 44 | 0.187
FDA | WAg | 0.45 | 45 | 0.188
FDA | WAg | 0.45 | 46 | 0.19
FDA | WAg | 0.45 | 47 | 0.191
FDA | WAg | 0.45 | 48 | 0.193
FDA | WAg | 0.45 | 49 | 0.194
FDA | WAg | 0.475 | 20 | 0.147
FDA | WAg | 0.475 | 21 | 0.151
FDA | WAg | 0.475 | 22 | 0.154
FDA | WAg | 0.475 | 23 | 0.158
FDA | WAg | 0.475 | 24 | 0.162
FDA | WAg | 0.475 | 25 | 0.165
FDA | WAg | 0.475 | 26 | 0.168
FDA | WAg | 0.475 | 27 | 0.171
FDA | WAg | 0.475 | 28 | 0.173
FDA | WAg | 0.475 | 29 | 0.174
FDA | WAg | 0.475 | 30 | 0.176
FDA | WAg | 0.475 | 31 | 0.177
FDA | WAg | 0.475 | 32 | 0.178
FDA | WAg | 0.475 | 33 | 0.179
FDA | WAg | 0.475 | 34 | 0.181
FDA | WAg | 0.475 | 35 | 0.182
FDA | WAg | 0.475 | 36 | 0.183
FDA | WAg | 0.475 | 37 | 0.184
FDA | WAg | 0.475 | 38 | 0.186
FDA | WAg | 0.475 | 39 | 0.187
FDA | WAg | 0.475 | 40 | 0.188
FDA | WAg | 0.475 | 41 | 0.19
FDA | WAg | 0.475 | 42 | 0.191
FDA | WAg | 0.475 | 43 | 0.193
FDA | WAg | 0.475 | 44 | 0.194
FDA | WAg | 0.475 | 45 | 0.196
FDA | WAg | 0.475 | 46 | 0.197
FDA | WAg | 0.475 | 47 | 0.198
FDA | WAg | 0.475 | 48 | 0.2
FDA | WAg | 0.475 | 49 | 0.201
FDA | WAg | 0.5 | 21 | 0.16
FDA | WAg | 0.5 | 22 | 0.163
FDA | WAg | 0.5 | 23 | 0.167
FDA | WAg | 0.5 | 24 | 0.171
FDA | WAg | 0.5 | 25 | 0.174
FDA | WAg | 0.5 | 26 | 0.177
FDA | WAg | 0.5 | 27 | 0.179
FDA | WAg | 0.5 | 28 | 0.181
FDA | WAg | 0.5 | 29 | 0.183
FDA | WAg | 0.5 | 30 | 0.184
FDA | WAg | 0.5 | 31 | 0.185
FDA | WAg | 0.5 | 32 | 0.186
FDA | WAg | 0.5 | 33 | 0.187
FDA | WAg | 0.5 | 34 | 0.188
FDA | WAg | 0.5 | 35 | 0.189
FDA | WAg | 0.5 | 36 | 0.191
FDA | WAg | 0.5 | 37 | 0.192
FDA | WAg | 0.5 | 38 | 0.193
FDA | WAg | 0.5 | 39 | 0.194
FDA | WAg | 0.5 | 40 | 0.196
FDA | WAg | 0.5 | 41 | 0.197
FDA | WAg | 0.5 | 42 | 0.198
FDA | WAg | 0.5 | 43 | 0.2
FDA | WAg | 0.5 | 44 | 0.201
FDA | WAg | 0.5 | 45 | 0.203
FDA | WAg | 0.5 | 46 | 0.204
FDA | WAg | 0.5 | 47 | 0.205
FDA | WAg | 0.5 | 48 | 0.207
FDA | WAg | 0.5 | 49 | 0.208
FDA | WAg | 0.525 | 21 | 0.169
FDA | WAg | 0.525 | 22 | 0.172
FDA | WAg | 0.525 | 23 | 0.176
FDA | WAg | 0.525 | 24 | 0.179
FDA | WAg | 0.525 | 25 | 0.182
FDA | WAg | 0.525 | 26 | 0.185
FDA | WAg | 0.525 | 27 | 0.188
FDA | WAg | 0.525 | 28 | 0.19
FDA | WAg | 0.525 | 29 | 0.191
FDA | WAg | 0.525 | 30 | 0.192
FDA | WAg | 0.525 | 31 | 0.193
FDA | WAg | 0.525 | 32 | 0.194
FDA | WAg | 0.525 | 33 | 0.195
FDA | WAg | 0.525 | 34 | 0.196
FDA | WAg | 0.525 | 35 | 0.197
FDA | WAg | 0.525 | 36 | 0.198
FDA | WAg | 0.525 | 37 | 0.199
FDA | WAg | 0.525 | 38 | 0.2
FDA | WAg | 0.525 | 39 | 0.202
FDA | WAg | 0.525 | 40 | 0.203
FDA | WAg | 0.525 | 41 | 0.204
FDA | WAg | 0.525 | 42 | 0.206
FDA | WAg | 0.525 | 43 | 0.207
FDA | WAg | 0.525 | 44 | 0.208
FDA | WAg | 0.525 | 45 | 0.21
FDA | WAg | 0.525 | 46 | 0.211
FDA | WAg | 0.525 | 47 | 0.212
FDA | WAg | 0.525 | 48 | 0.214
FDA | WAg | 0.525 | 49 | 0.215
FDA | WAg | 0.55 | 22 | 0.182
FDA | WAg | 0.55 | 23 | 0.185
FDA | WAg | 0.55 | 24 | 0.188
FDA | WAg | 0.55 | 25 | 0.191
FDA | WAg | 0.55 | 26 | 0.194
FDA | WAg | 0.55 | 27 | 0.196
FDA | WAg | 0.55 | 28 | 0.198
FDA | WAg | 0.55 | 29 | 0.199
FDA | WAg | 0.55 | 30 | 0.2
FDA | WAg | 0.55 | 31 | 0.201
FDA | WAg | 0.55 | 32 | 0.202
FDA | WAg | 0.55 | 33 | 0.203
FDA | WAg | 0.55 | 34 | 0.204
FDA | WAg | 0.55 | 35 | 0.205
FDA | WAg | 0.55 | 36 | 0.206
FDA | WAg | 0.55 | 37 | 0.207
FDA | WAg | 0.55 | 38 | 0.208
FDA | WAg | 0.55 | 39 | 0.209
FDA | WAg | 0.55 | 40 | 0.21
FDA | WAg | 0.55 | 41 | 0.211
FDA | WAg | 0.55 | 42 | 0.213
FDA | WAg | 0.55 | 43 | 0.214
FDA | WAg | 0.55 | 44 | 0.215
FDA | WAg | 0.55 | 45 | 0.217
FDA | WAg | 0.55 | 46 | 0.218
FDA | WAg | 0.55 | 47 | 0.219
FDA | WAg | 0.55 | 48 | 0.221
FDA | WAg | 0.55 | 49 | 0.222
FDA | WAg | 0.575 | 22 | 0.19
FDA | WAg | 0.575 | 23 | 0.194
FDA | WAg | 0.575 | 24 | 0.197
FDA | WAg | 0.575 | 25 | 0.199
FDA | WAg | 0.575 | 26 | 0.202
FDA | WAg | 0.575 | 27 | 0.205
FDA | WAg | 0.575 | 28 | 0.206
FDA | WAg | 0.575 | 29 | 0.207
FDA | WAg | 0.575 | 30 | 0.208
FDA | WAg | 0.575 | 31 | 0.209
FDA | WAg | 0.575 | 32 | 0.21
FDA | WAg | 0.575 | 33 | 0.211
FDA | WAg | 0.575 | 34 | 0.212
FDA | WAg | 0.575 | 35 | 0.212
FDA | WAg | 0.575 | 36 | 0.213
FDA | WAg | 0.575 | 37 | 0.214
FDA | WAg | 0.575 | 38 | 0.215
FDA | WAg | 0.575 | 39 | 0.216
FDA | WAg | 0.575 | 40 | 0.217
FDA | WAg | 0.575 | 41 | 0.219
FDA | WAg | 0.575 | 42 | 0.22
FDA | WAg | 0.575 | 43 | 0.221
FDA | WAg | 0.575 | 44 | 0.223
FDA | WAg | 0.575 | 45 | 0.224
FDA | WAg | 0.575 | 46 | 0.225
FDA | WAg | 0.575 | 47 | 0.226
FDA | WAg | 0.575 | 48 | 0.228
FDA | WAg | 0.575 | 49 | 0.229
FDA | WAg | 0.6 | 23 | 0.202
FDA | WAg | 0.6 | 24 | 0.205
FDA | WAg | 0.6 | 25 | 0.208
FDA | WAg | 0.6 | 26 | 0.211
FDA | WAg | 0.6 | 27 | 0.213
FDA | WAg | 0.6 | 28 | 0.214
FDA | WAg | 0.6 | 29 | 0.215
FDA | WAg | 0.6 | 30 | 0.216
FDA | WAg | 0.6 | 31 | 0.217
FDA | WAg | 0.6 | 32 | 0.218
FDA | WAg | 0.6 | 33 | 0.218
FDA | WAg | 0.6 | 34 | 0.219
FDA | WAg | 0.6 | 35 | 0.22
FDA | WAg | 0.6 | 36 | 0.221
FDA | WAg | 0.6 | 37 | 0.222
FDA | WAg | 0.6 | 38 | 0.222
FDA | WAg | 0.6 | 39 | 0.224
FDA | WAg | 0.6 | 40 | 0.225
FDA | WAg | 0.6 | 41 | 0.226
FDA | WAg | 0.6 | 42 | 0.227
FDA | WAg | 0.6 | 43 | 0.228
FDA | WAg | 0.6 | 44 | 0.229
FDA | WAg | 0.6 | 45 | 0.231
FDA | WAg | 0.6 | 46 | 0.232
FDA | WAg | 0.6 | 47 | 0.233
FDA | WAg | 0.6 | 48 | 0.234
FDA | WAg | 0.6 | 49 | 0.236
FDA | WAg | 0.625 | 23 | 0.211
FDA | WAg | 0.625 | 24 | 0.214
FDA | WAg | 0.625 | 25 | 0.216
FDA | WAg | 0.625 | 26 | 0.219
FDA | WAg | 0.625 | 27 | 0.221
FDA | WAg | 0.625 | 28 | 0.222
FDA | WAg | 0.625 | 29 | 0.223
FDA | WAg | 0.625 | 30 | 0.224
FDA | WAg | 0.625 | 31 | 0.225
FDA | WAg | 0.625 | 32 | 0.225
FDA | WAg | 0.625 | 33 | 0.226
FDA | WAg | 0.625 | 34 | 0.227
FDA | WAg | 0.625 | 35 | 0.227
FDA | WAg | 0.625 | 36 | 0.228
FDA | WAg | 0.625 | 37 | 0.229
FDA | WAg | 0.625 | 38 | 0.23
FDA | WAg | 0.625 | 39 | 0.231
FDA | WAg | 0.625 | 40 | 0.232
FDA | WAg | 0.625 | 41 | 0.233
FDA | WAg | 0.625 | 42 | 0.234
FDA | WAg | 0.625 | 43 | 0.235
FDA | WAg | 0.625 | 44 | 0.236
FDA | WAg | 0.625 | 45 | 0.237
FDA | WAg | 0.625 | 46 | 0.239
FDA | WAg | 0.625 | 47 | 0.24
FDA | WAg | 0.625 | 48 | 0.241
FDA | WAg | 0.625 | 49 | 0.242
FDA | WAg | 0.65 | 28 | 0.23
FDA | WAg | 0.65 | 29 | 0.231
FDA | WAg | 0.65 | 30 | 0.232
FDA | WAg | 0.65 | 31 | 0.232
FDA | WAg | 0.65 | 32 | 0.233
FDA | WAg | 0.65 | 33 | 0.234
FDA | WAg | 0.65 | 34 | 0.234
FDA | WAg | 0.65 | 35 | 0.235
FDA | WAg | 0.65 | 36 | 0.235
FDA | WAg | 0.65 | 37 | 0.236
FDA | WAg | 0.65 | 38 | 0.237
FDA | WAg | 0.65 | 39 | 0.238
FDA | WAg | 0.65 | 40 | 0.239
FDA | WAg | 0.65 | 41 | 0.24
FDA | WAg | 0.65 | 42 | 0.241
FDA | WAg | 0.65 | 43 | 0.242
FDA | WAg | 0.65 | 44 | 0.243
FDA | WAg | 0.65 | 45 | 0.244
FDA | WAg | 0.65 | 46 | 0.245
FDA | WAg | 0.65 | 47 | 0.247
FDA | WAg | 0.65 | 48 | 0.248
FDA | WAg | 0.65 | 49 | 0.249
FDA | WAg | 0.675 | 28 | 0.238
FDA | WAg | 0.675 | 29 | 0.239
FDA | WAg | 0.675 | 30 | 0.24
FDA | WAg | 0.675 | 31 | 0.24
FDA | WAg | 0.675 | 32 | 0.241
FDA | WAg | 0.675 | 33 | 0.241
FDA | WAg | 0.675 | 34 | 0.242
FDA | WAg | 0.675 | 35 | 0.242
FDA | WAg | 0.675 | 36 | 0.243
FDA | WAg | 0.675 | 37 | 0.243
FDA | WAg | 0.675 | 38 | 0.244
FDA | WAg | 0.675 | 39 | 0.245
FDA | WAg | 0.675 | 40 | 0.246
FDA | WAg | 0.675 | 41 | 0.247
FDA | WAg | 0.675 | 42 | 0.248
FDA | WAg | 0.675 | 43 | 0.249
FDA | WAg | 0.675 | 44 | 0.25
FDA | WAg | 0.675 | 45 | 0.251
FDA | WAg | 0.675 | 46 | 0.252
FDA | WAg | 0.675 | 47 | 0.253
FDA | WAg | 0.675 | 48 | 0.254
FDA | WAg | 0.675 | 49 | 0.256
FDA | WAg | 0.7 | 32 | 0.248
FDA | WAg | 0.7 | 33 | 0.248
FDA | WAg | 0.7 | 34 | 0.249
FDA | WAg | 0.7 | 35 | 0.249
FDA | WAg | 0.7 | 36 | 0.25
FDA | WAg | 0.7 | 37 | 0.25
FDA | WAg | 0.7 | 38 | 0.251
FDA | WAg | 0.7 | 39 | 0.252
FDA | WAg | 0.7 | 40 | 0.253
FDA | WAg | 0.7 | 41 | 0.254
FDA | WAg | 0.7 | 42 | 0.255
FDA | WAg | 0.7 | 43 | 0.256
FDA | WAg | 0.7 | 44 | 0.257
FDA | WAg | 0.7 | 45 | 0.258
FDA | WAg | 0.7 | 46 | 0.259
FDA | WAg | 0.7 | 47 | 0.26
FDA | WAg | 0.7 | 48 | 0.261
FDA | WAg | 0.7 | 49 | 0.262
FDA | WAg | 0.725 | 35 | 0.257
FDA | WAg | 0.725 | 36 | 0.257
FDA | WAg | 0.725 | 37 | 0.257
FDA | WAg | 0.725 | 38 | 0.258
FDA | WAg | 0.725 | 39 | 0.259
FDA | WAg | 0.725 | 40 | 0.26
FDA | WAg | 0.725 | 41 | 0.26
FDA | WAg | 0.725 | 42 | 0.261
FDA | WAg | 0.725 | 43 | 0.262
FDA | WAg | 0.725 | 44 | 0.263
FDA | WAg | 0.725 | 45 | 0.264
FDA | WAg | 0.725 | 46 | 0.265
FDA | WAg | 0.725 | 47 | 0.266
FDA | WAg | 0.725 | 48 | 0.267
FDA | WAg | 0.725 | 49 | 0.269
FDA | WAg | 0.75 | 44 | 0.27
FDA | WAg | 0.75 | 45 | 0.271
FDA | WAg | 0.75 | 46 | 0.272
FDA | WAg | 0.75 | 47 | 0.273
FDA | WAg | 0.75 | 48 | 0.274
FDA | WAg | 0.75 | 49 | 0.275
FDA | WAl | 0.2 | 20 | 0.048
FDA | WAl | 0.2 | 21 | 0.051
FDA | WAl | 0.25 | 20 | 0.065
FDA | WAl | 0.25 | 21 | 0.069
FDA | WAl | 0.25 | 22 | 0.073
FDA | WAl | 0.25 | 23 | 0.076
FDA | WAl | 0.25 | 24 | 0.08
FDA | WAl | 0.25 | 25 | 0.083
FDA | WAl | 0.3 | 20 | 0.082
FDA | WAl | 0.3 | 21 | 0.086
FDA | WAl | 0.3 | 22 | 0.09
FDA | WAl | 0.3 | 23 | 0.094
FDA | WAl | 0.3 | 24 | 0.098
FDA | WAl | 0.3 | 25 | 0.101
FDA | WAl | 0.3 | 26 | 0.105
FDA | WAl | 0.3 | 27 | 0.108
FDA | WAl | 0.3 | 28 | 0.111
FDA | WAl | 0.3 | 29 | 0.114
FDA | WAl | 0.3 | 30 | 0.117
FDA | WAl | 0.35 | 20 | 0.1
FDA | WAl | 0.35 | 21 | 0.104
FDA | WAl | 0.35 | 22 | 0.108
FDA | WAl | 0.35 | 23 | 0.112
FDA | WAl | 0.35 | 24 | 0.116
FDA | WAl | 0.35 | 25 | 0.12
FDA | WAl | 0.35 | 26 | 0.123
FDA | WAl | 0.35 | 27 | 0.126
FDA | WAl | 0.35 | 28 | 0.129
FDA | WAl | 0.35 | 29 | 0.132
FDA | WAl | 0.35 | 30 | 0.135
FDA | WAl | 0.35 | 31 | 0.138
FDA | WAl | 0.35 | 32 | 0.14
FDA | WAl | 0.35 | 33 | 0.143
FDA | WAl | 0.35 | 34 | 0.145
FDA | WAl | 0.35 | 35 | 0.148
FDA | WAl | 0.4 | 20 | 0.118
FDA | WAl | 0.4 | 21 | 0.123
FDA | WAl | 0.4 | 22 | 0.127
FDA | WAl | 0.4 | 23 | 0.13
FDA | WAl | 0.4 | 24 | 0.134
FDA | WAl | 0.4 | 25 | 0.138
FDA | WAl | 0.4 | 26 | 0.141
FDA | WAl | 0.4 | 27 | 0.144
FDA | WAl | 0.4 | 28 | 0.147
FDA | WAl | 0.4 | 29 | 0.15
FDA | WAl | 0.4 | 30 | 0.153
FDA | WAl | 0.4 | 31 | 0.156
FDA | WAl | 0.4 | 32 | 0.158
FDA | WAl | 0.4 | 33 | 0.161
FDA | WAl | 0.4 | 34 | 0.163
FDA | WAl | 0.4 | 35 | 0.165
FDA | WAl | 0.4 | 36 | 0.167
FDA | WAl | 0.4 | 37 | 0.169
FDA | WAl | 0.4 | 38 | 0.171
FDA | WAl | 0.4 | 39 | 0.173
FDA | WAl | 0.4 | 40 | 0.175
FDA | WAl | 0.45 | 20 | 0.137
FDA | WAl | 0.45 | 21 | 0.141
FDA | WAl | 0.45 | 22 | 0.145
FDA | WAl | 0.45 | 23 | 0.149
FDA | WAl | 0.45 | 24 | 0.152
FDA | WAl | 0.45 | 25 | 0.155
FDA | WAl | 0.45 | 26 | 0.158
FDA | WAl | 0.45 | 27 | 0.162
FDA | WAl | 0.45 | 28 | 0.165
FDA | WAl | 0.45 | 29 | 0.168
FDA | WAl | 0.45 | 30 | 0.17
FDA | WAl | 0.45 | 31 | 0.173
FDA | WAl | 0.45 | 32 | 0.175
FDA | WAl | 0.45 | 33 | 0.178
FDA | WAl | 0.45 | 34 | 0.18
FDA | WAl | 0.45 | 35 | 0.182
FDA | WAl | 0.45 | 36 | 0.184
FDA | WAl | 0.45 | 37 | 0.186
FDA | WAl | 0.45 | 38 | 0.188
FDA | WAl | 0.45 | 39 | 0.19
FDA | WAl | 0.45 | 40 | 0.192
FDA | WAl | 0.45 | 41 | 0.194
FDA | WAl | 0.45 | 42 | 0.195
FDA | WAl | 0.45 | 43 | 0.197
FDA | WAl | 0.45 | 44 | 0.198
FDA | WAl | 0.5 | 20 | 0.156
FDA | WAl | 0.5 | 21 | 0.16
FDA | WAl | 0.5 | 22 | 0.163
FDA | WAl | 0.5 | 23 | 0.167
FDA | WAl | 0.5 | 24 | 0.17
FDA | WAl | 0.5 | 25 | 0.173
FDA | WAl | 0.5 | 26 | 0.176
FDA | WAl | 0.5 | 27 | 0.179
FDA | WAl | 0.5 | 28 | 0.182
FDA | WAl | 0.5 | 29 | 0.185
FDA | WAl | 0.5 | 30 | 0.187
FDA | WAl | 0.5 | 31 | 0.19
FDA | WAl | 0.5 | 32 | 0.192
FDA | WAl | 0.5 | 33 | 0.194
FDA | WAl | 0.5 | 34 | 0.197
FDA | WAl | 0.5 | 35 | 0.198
FDA | WAl | 0.5 | 36 | 0.201
FDA | WAl | 0.5 | 37 | 0.202
FDA | WAl | 0.5 | 38 | 0.204
FDA | WAl | 0.5 | 39 | 0.206
FDA | WAl | 0.5 | 40 | 0.208
FDA | WAl | 0.5 | 41 | 0.209
FDA | WAl | 0.5 | 42 | 0.211
FDA | WAl | 0.5 | 43 | 0.213
FDA | WAl | 0.5 | 44 | 0.214
FDA | WAl | 0.5 | 45 | 0.216
FDA | WAl | 0.5 | 46 | 0.217
FDA | WAl | 0.5 | 47 | 0.218
FDA | WAl | 0.5 | 48 | 0.22
FDA | WAl | 0.5 | 49 | 0.221
FDA | WAl | 0.55 | 21 | 0.178
FDA | WAl | 0.55 | 22 | 0.181
FDA | WAl | 0.55 | 23 | 0.184
FDA | WAl | 0.55 | 24 | 0.188
FDA | WAl | 0.55 | 25 | 0.19
FDA | WAl | 0.55 | 26 | 0.193
FDA | WAl | 0.55 | 27 | 0.196
FDA | WAl | 0.55 | 28 | 0.199
FDA | WAl | 0.55 | 29 | 0.201
FDA | WAl | 0.55 | 30 | 0.204
FDA | WAl | 0.55 | 31 | 0.206
FDA | WAl | 0.55 | 32 | 0.209
FDA | WAl | 0.55 | 33 | 0.211
FDA | WAl | 0.55 | 34 | 0.213
FDA | WAl | 0.55 | 35 | 0.215
FDA | WAl | 0.55 | 36 | 0.217
FDA | WAl | 0.55 | 37 | 0.218
FDA | WAl | 0.55 | 38 | 0.22
FDA | WAl | 0.55 | 39 | 0.222
FDA | WAl | 0.55 | 40 | 0.223
FDA | WAl | 0.55 | 41 | 0.225
FDA | WAl | 0.55 | 42 | 0.226
FDA | WAl | 0.55 | 43 | 0.228
FDA | WAl | 0.55 | 44 | 0.229
FDA | WAl | 0.55 | 45 | 0.231
FDA | WAl | 0.55 | 46 | 0.232
FDA | WAl | 0.55 | 47 | 0.233
FDA | WAl | 0.55 | 48 | 0.235
FDA | WAl | 0.55 | 49 | 0.236
FDA | WAl | 0.6 | 22 | 0.199
FDA | WAl | 0.6 | 23 | 0.202
FDA | WAl | 0.6 | 24 | 0.205
FDA | WAl | 0.6 | 25 | 0.207
FDA | WAl | 0.6 | 26 | 0.21
FDA | WAl | 0.6 | 27 | 0.213
FDA | WAl | 0.6 | 28 | 0.215
FDA | WAl | 0.6 | 29 | 0.218
FDA | WAl | 0.6 | 30 | 0.22
FDA | WAl | 0.6 | 31 | 0.222
FDA | WAl | 0.6 | 32 | 0.225
FDA | WAl | 0.6 | 33 | 0.227
FDA | WAl | 0.6 | 34 | 0.229
FDA | WAl | 0.6 | 35 | 0.23
FDA | WAl | 0.6 | 36 | 0.232
FDA | WAl | 0.6 | 37 | 0.234
FDA | WAl | 0.6 | 38 | 0.235
FDA | WAl | 0.6 | 39 | 0.237
FDA | WAl | 0.6 | 40 | 0.238
FDA | WAl | 0.6 | 41 | 0.24
FDA | WAl | 0.6 | 42 | 0.241
FDA | WAl | 0.6 | 43 | 0.243
FDA | WAl | 0.6 | 44 | 0.244
FDA | WAl | 0.6 | 45 | 0.245
FDA | WAl | 0.6 | 46 | 0.247
FDA | WAl | 0.6 | 47 | 0.248
FDA | WAl | 0.6 | 48 | 0.249
FDA | WAl | 0.6 | 49 | 0.25
FDA | WAl | 0.65 | 23 | 0.219
FDA | WAl | 0.65 | 24 | 0.222
FDA | WAl | 0.65 | 25 | 0.224
FDA | WAl | 0.65 | 26 | 0.226
FDA | WAl | 0.65 | 27 | 0.229
FDA | WAl | 0.65 | 28 | 0.232
FDA | WAl | 0.65 | 29 | 0.234
FDA | WAl | 0.65 | 30 | 0.236
FDA | WAl | 0.65 | 31 | 0.238
FDA | WAl | 0.65 | 32 | 0.24
FDA | WAl | 0.65 | 33 | 0.242
FDA | WAl | 0.65 | 34 | 0.244
FDA | WAl | 0.65 | 35 | 0.246
FDA | WAl | 0.65 | 36 | 0.247
FDA | WAl | 0.65 | 37 | 0.249
FDA | WAl | 0.65 | 38 | 0.25
FDA | WAl | 0.65 | 39 | 0.252
FDA | WAl | 0.65 | 40 | 0.253
FDA | WAl | 0.65 | 41 | 0.255
FDA | WAl | 0.65 | 42 | 0.256
FDA | WAl | 0.65 | 43 | 0.257
FDA | WAl | 0.65 | 44 | 0.258
FDA | WAl | 0.65 | 45 | 0.26
FDA | WAl | 0.65 | 46 | 0.261
FDA | WAl | 0.65 | 47 | 0.262
FDA | WAl | 0.65 | 48 | 0.263
FDA | WAl | 0.65 | 49 | 0.264
FDA | WAl | 0.7 | 23 | 0.236
FDA | WAl | 0.7 | 24 | 0.238
FDA | WAl | 0.7 | 25 | 0.24
FDA | WAl | 0.7 | 26 | 0.242
FDA | WAl | 0.7 | 27 | 0.245
FDA | WAl | 0.7 | 28 | 0.247
FDA | WAl | 0.7 | 29 | 0.249
FDA | WAl | 0.7 | 30 | 0.251
FDA | WAl | 0.7 | 31 | 0.253
FDA | WAl | 0.7 | 32 | 0.255
FDA | WAl | 0.7 | 33 | 0.257
FDA | WAl | 0.7 | 34 | 0.259
FDA | WAl | 0.7 | 35 | 0.26
FDA | WAl | 0.7 | 36 | 0.262
FDA | WAl | 0.7 | 37 | 0.263
FDA | WAl | 0.7 | 38 | 0.265
FDA | WAl | 0.7 | 39 | 0.266
FDA | WAl | 0.7 | 40 | 0.267
FDA | WAl | 0.7 | 41 | 0.269
FDA | WAl | 0.7 | 42 | 0.27
FDA | WAl | 0.7 | 43 | 0.271
FDA | WAl | 0.7 | 44 | 0.272
FDA | WAl | 0.7 | 45 | 0.273
FDA | WAl | 0.7 | 46 | 0.274
FDA | WAl | 0.7 | 47 | 0.276
FDA | WAl | 0.7 | 48 | 0.277
FDA | WAl | 0.7 | 49 | 0.278
FDA | WAl | 0.75 | 24 | 0.254
FDA | WAl | 0.75 | 25 | 0.256
FDA | WAl | 0.75 | 26 | 0.258
FDA | WAl | 0.75 | 27 | 0.26
FDA | WAl | 0.75 | 28 | 0.262
FDA | WAl | 0.75 | 29 | 0.264
FDA | WAl | 0.75 | 30 | 0.266
FDA | WAl | 0.75 | 31 | 0.268
FDA | WAl | 0.75 | 32 | 0.27
FDA | WAl | 0.75 | 33 | 0.272
FDA | WAl | 0.75 | 34 | 0.273
FDA | WAl | 0.75 | 35 | 0.275
FDA | WAl | 0.75 | 36 | 0.276
FDA | WAl | 0.75 | 37 | 0.278
FDA | WAl | 0.75 | 38 | 0.279
FDA | WAl | 0.75 | 39 | 0.28
FDA | WAl | 0.75 | 40 | 0.281
FDA | WAl | 0.75 | 41 | 0.283
FDA | WAl | 0.75 | 42 | 0.284
FDA | WAl | 0.75 | 43 | 0.285
FDA | WAl | 0.75 | 44 | 0.286
FDA | WAl | 0.75 | 45 | 0.287
FDA | WAl | 0.75 | 46 | 0.288
FDA | WAl | 0.75 | 47 | 0.289
FDA | WAl | 0.75 | 48 | 0.29
FDA | WAl | 0.75 | 49 | 0.291
FDA | WAl | 0.8 | 25 | 0.271
FDA | WAl | 0.8 | 26 | 0.273
FDA | WAl | 0.8 | 27 | 0.275
FDA | WAl | 0.8 | 28 | 0.277
FDA | WAl | 0.8 | 29 | 0.279
FDA | WAl | 0.8 | 30 | 0.281
FDA | WAl | 0.8 | 31 | 0.282
FDA | WAl | 0.8 | 32 | 0.284
FDA | WAl | 0.8 | 33 | 0.286
FDA | WAl | 0.8 | 34 | 0.287
FDA | WAl | 0.8 | 35 | 0.289
FDA | WAl | 0.8 | 36 | 0.29
FDA | WAl | 0.8 | 37 | 0.291
FDA | WAl | 0.8 | 38 | 0.293
FDA | WAl | 0.8 | 39 | 0.294
FDA | WAl | 0.8 | 40 | 0.295
FDA | WAl | 0.8 | 41 | 0.296
FDA | WAl | 0.8 | 42 | 0.297
FDA | WAl | 0.8 | 43 | 0.298
FDA | WAl | 0.8 | 44 | 0.299
FDA | WAl | 0.8 | 45 | 0.3
FDA | WAl | 0.8 | 46 | 0.301
FDA | WAl | 0.8 | 47 | 0.302
FDA | WAl | 0.8 | 48 | 0.303
FDA | WAl | 0.8 | 49 | 0.303
FDA | WAl | 0.85 | 26 | 0.288
FDA | WAl | 0.85 | 27 | 0.289
FDA | WAl | 0.85 | 28 | 0.291
FDA | WAl | 0.85 | 29 | 0.293
FDA | WAl | 0.85 | 30 | 0.294
FDA | WAl | 0.85 | 31 | 0.296
FDA | WAl | 0.85 | 32 | 0.298
FDA | WAl | 0.85 | 33 | 0.299
FDA | WAl | 0.85 | 34 | 0.301
FDA | WAl | 0.85 | 35 | 0.302
FDA | WAl | 0.85 | 36 | 0.303
FDA | WAl | 0.85 | 37 | 0.304
FDA | WAl | 0.85 | 38 | 0.306
FDA | WAl | 0.85 | 39 | 0.307
FDA | WAl | 0.85 | 40 | 0.308
FDA | WAl | 0.85 | 41 | 0.309
FDA | WAl | 0.85 | 42 | 0.31
FDA | WAl | 0.85 | 43 | 0.311
FDA | WAl | 0.85 | 44 | 0.312
FDA | WAl | 0.85 | 45 | 0.313
FDA | WAl | 0.85 | 46 | 0.313
FDA | WAl | 0.85 | 47 | 0.314
FDA | WAl | 0.85 | 48 | 0.315
FDA | WAl | 0.85 | 49 | 0.316
FDA | WAl | 0.9 | 27 | 0.303
FDA | WAl | 0.9 | 28 | 0.305
FDA | WAl | 0.9 | 29 | 0.307
FDA | WAl | 0.9 | 30 | 0.308
FDA | WAl | 0.9 | 31 | 0.309
FDA | WAl | 0.9 | 32 | 0.311
FDA | WAl | 0.9 | 33 | 0.312
FDA | WAl | 0.9 | 34 | 0.314
FDA | WAl | 0.9 | 35 | 0.315
FDA | WAl | 0.9 | 36 | 0.316
FDA | WAl | 0.9 | 37 | 0.317
FDA | WAl | 0.9 | 38 | 0.318
FDA | WAl | 0.9 | 39 | 0.319
FDA | WAl | 0.9 | 40 | 0.32
FDA | WAl | 0.9 | 41 | 0.321
FDA | WAl | 0.9 | 42 | 0.322
FDA | WAl | 0.9 | 43 | 0.323
FDA | WAl | 0.9 | 44 | 0.324
FDA | WAl | 0.9 | 45 | 0.325
FDA | WAl | 0.9 | 46 | 0.326
FDA | WAl | 0.9 | 47 | 0.326
FDA | WAl | 0.9 | 48 | 0.327
FDA | WAl | 0.9 | 49 | 0.328
FDA | WAl | 0.95 | 27 | 0.317
FDA | WAl | 0.95 | 28 | 0.318
FDA | WAl | 0.95 | 29 | 0.32
FDA | WAl | 0.95 | 30 | 0.321
FDA | WAl | 0.95 | 31 | 0.322
FDA | WAl | 0.95 | 32 | 0.324
FDA | WAl | 0.95 | 33 | 0.325
FDA | WAl | 0.95 | 34 | 0.326
FDA | WAl | 0.95 | 35 | 0.327
FDA | WAl | 0.95 | 36 | 0.328
FDA | WAl | 0.95 | 37 | 0.33
FDA | WAl | 0.95 | 38 | 0.331
FDA | WAl | 0.95 | 39 | 0.332
FDA | WAl | 0.95 | 40 | 0.333
FDA | WAl | 0.95 | 41 | 0.333
FDA | WAl | 0.95 | 42 | 0.334
FDA | WAl | 0.95 | 43 | 0.335
FDA | WAl | 0.95 | 44 | 0.336
FDA | WAl | 0.95 | 45 | 0.337
FDA | WAl | 0.95 | 46 | 0.337
FDA | WAl | 0.95 | 47 | 0.338
FDA | WAl | 0.95 | 48 | 0.339
FDA | WAl | 0.95 | 49 | 0.339
FDA | WAl | 1 | 28 | 0.331
FDA | WAl | 1 | 29 | 0.332
FDA | WAl | 1 | 30 | 0.333
FDA | WAl | 1 | 31 | 0.335
FDA | WAl | 1 | 32 | 0.336
FDA | WAl | 1 | 33 | 0.337
FDA | WAl | 1 | 34 | 0.338
FDA | WAl | 1 | 35 | 0.339
FDA | WAl | 1 | 36 | 0.34
FDA | WAl | 1 | 37 | 0.341
FDA | WAl | 1 | 38 | 0.342
FDA | WAl | 1 | 39 | 0.343
FDA | WAl | 1 | 40 | 0.344
FDA | WAl | 1 | 41 | 0.345
FDA | WAl | 1 | 42 | 0.346
FDA | WAl | 1 | 43 | 0.347
FDA | WAl | 1 | 44 | 0.347
FDA | WAl | 1 | 45 | 0.348
FDA | WAl | 1 | 46 | 0.349
FDA | WAl | 1 | 47 | 0.349
FDA | WAl | 1 | 48 | 0.35
FDA | WAl | 1 | 49 | 0.351
FDA | WAl | 1.05 | 29 | 0.344
FDA | WAl | 1.05 | 30 | 0.346
FDA | WAl | 1.05 | 31 | 0.347
FDA | WAl | 1.05 | 32 | 0.348
FDA | WAl | 1.05 | 33 | 0.349
FDA | WAl | 1.05 | 34 | 0.35
FDA | WAl | 1.05 | 35 | 0.351
FDA | WAl | 1.05 | 36 | 0.352
FDA | WAl | 1.05 | 37 | 0.353
FDA | WAl | 1.05 | 38 | 0.354
FDA | WAl | 1.05 | 39 | 0.355
FDA | WAl | 1.05 | 40 | 0.355
FDA | WAl | 1.05 | 41 | 0.356
FDA | WAl | 1.05 | 42 | 0.357
FDA | WAl | 1.05 | 43 | 0.358
FDA | WAl | 1.05 | 44 | 0.358
FDA | WAl | 1.05 | 45 | 0.359
FDA | WAl | 1.05 | 46 | 0.36
FDA | WAl | 1.05 | 47 | 0.36
FDA | WAl | 1.05 | 48 | 0.361
FDA | WAl | 1.05 | 49 | 0.362
FDA | WAl | 1.1 | 29 | 0.356
FDA | WAl | 1.1 | 30 | 0.357
FDA | WAl | 1.1 | 31 | 0.358
FDA | WAl | 1.1 | 32 | 0.359
FDA | WAl | 1.1 | 33 | 0.36
FDA | WAl | 1.1 | 34 | 0.361
FDA | WAl | 1.1 | 35 | 0.362
FDA | WAl | 1.1 | 36 | 0.363
FDA | WAl | 1.1 | 37 | 0.364
FDA | WAl | 1.1 | 38 | 0.365
FDA | WAl | 1.1 | 39 | 0.366
FDA | WAl | 1.1 | 40 | 0.366
FDA | WAl | 1.1 | 41 | 0.367
FDA | WAl | 1.1 | 42 | 0.368
FDA | WAl | 1.1 | 43 | 0.368
FDA | WAl | 1.1 | 44 | 0.369
FDA | WAl | 1.1 | 45 | 0.37
FDA | WAl | 1.1 | 46 | 0.37
FDA | WAl | 1.1 | 47 | 0.371
FDA | WAl | 1.1 | 48 | 0.371
FDA | WAl | 1.1 | 49 | 0.372
FDA | WAl | 1.15 | 30 | 0.369
FDA | WAl | 1.15 | 31 | 0.37
FDA | WAl | 1.15 | 32 | 0.37
FDA | WAl | 1.15 | 33 | 0.371
FDA | WAl | 1.15 | 34 | 0.372
FDA | WAl | 1.15 | 35 | 0.373
FDA | WAl | 1.15 | 36 | 0.374
FDA | WAl | 1.15 | 37 | 0.375
FDA | WAl | 1.15 | 38 | 0.375
FDA | WAl | 1.15 | 39 | 0.376
FDA | WAl | 1.15 | 40 | 0.377
FDA | WAl | 1.15 | 41 | 0.378
FDA | WAl | 1.15 | 42 | 0.378
FDA | WAl | 1.15 | 43 | 0.379
FDA | WAl | 1.15 | 44 | 0.379
FDA | WAl | 1.15 | 45 | 0.38
FDA | WAl | 1.15 | 46 | 0.38
FDA | WAl | 1.15 | 47 | 0.381
FDA | WAl | 1.15 | 48 | 0.381
FDA | WAl | 1.15 | 49 | 0.382
FDA | WAl | 1.2 | 31 | 0.38
FDA | WAl | 1.2 | 32 | 0.381
FDA | WAl | 1.2 | 33 | 0.382
FDA | WAl | 1.2 | 34 | 0.383
FDA | WAl | 1.2 | 35 | 0.384
FDA | WAl | 1.2 | 36 | 0.384
FDA | WAl | 1.2 | 37 | 0.385
FDA | WAl | 1.2 | 38 | 0.386
FDA | WAl | 1.2 | 39 | 0.386
FDA | WAl | 1.2 | 40 | 0.387
FDA | WAl | 1.2 | 41 | 0.388
FDA | WAl | 1.2 | 42 | 0.388
FDA | WAl | 1.2 | 43 | 0.389
FDA | WAl | 1.2 | 44 | 0.389
FDA | WAl | 1.2 | 45 | 0.39
FDA | WAl | 1.2 | 46 | 0.39
FDA | WAl | 1.2 | 47 | 0.391
FDA | WAl | 1.2 | 48 | 0.391
FDA | WAl | 1.2 | 49 | 0.392
FDA | WAl | 1.25 | 36 | 0.395
FDA | WAl | 1.25 | 37 | 0.395
FDA | WAl | 1.25 | 38 | 0.396
FDA | WAl | 1.25 | 39 | 0.396
FDA | WAl | 1.25 | 40 | 0.397
FDA | WAl | 1.25 | 41 | 0.397
FDA | WAl | 1.25 | 42 | 0.398
FDA | WAl | 1.25 | 43 | 0.399
FDA | WAl | 1.25 | 44 | 0.399
FDA | WAl | 1.25 | 45 | 0.399
FDA | WAl | 1.25 | 46 | 0.4
FDA | WAl | 1.25 | 47 | 0.4
FDA | WAl | 1.25 | 48 | 0.401
FDA | WAl | 1.25 | 49 | 0.401
FDA | WAl | 1.3 | 41 | 0.407
FDA | WAl | 1.3 | 42 | 0.407
FDA | WAl | 1.3 | 43 | 0.408
FDA | WAl | 1.3 | 44 | 0.408
FDA | WAl | 1.3 | 45 | 0.409
FDA | WAl | 1.3 | 46 | 0.409
FDA | WAl | 1.3 | 47 | 0.41
FDA | WAl | 1.3 | 48 | 0.41
FDA | WAl | 1.3 | 49 | 0.41
FDA | WAl | 1.35 | 45 | 0.418
FDA | WAl | 1.35 | 46 | 0.418
FDA | WAl | 1.35 | 47 | 0.418
FDA | WAl | 1.35 | 48 | 0.419
FDA | WAl | 1.35 | 49 | 0.419
FDA | WRh | 0.2 | 20 | 0.048
FDA | WRh | 0.2 | 21 | 0.051
FDA | WRh | 0.225 | 20 | 0.057
FDA | WRh | 0.225 | 21 | 0.06
FDA | WRh | 0.225 | 22 | 0.064
FDA | WRh | 0.225 | 23 | 0.067
FDA | WRh | 0.225 | 24 | 0.071
FDA | WRh | 0.25 | 20 | 0.065
FDA | WRh | 0.25 | 21 | 0.069
FDA | WRh | 0.25 | 22 | 0.073
FDA | WRh | 0.25 | 23 | 0.077
FDA | WRh | 0.25 | 24 | 0.08
FDA | WRh | 0.25 | 25 | 0.083
FDA | WRh | 0.25 | 26 | 0.086
FDA | WRh | 0.275 | 20 | 0.074
FDA | WRh | 0.275 | 21 | 0.078
FDA | WRh | 0.275 | 22 | 0.082
FDA | WRh | 0.275 | 23 | 0.086
FDA | WRh | 0.275 | 24 | 0.09
FDA | WRh | 0.275 | 25 | 0.093
FDA | WRh | 0.275 | 26 | 0.095
FDA | WRh | 0.275 | 27 | 0.097
FDA | WRh | 0.275 | 28 | 0.099
FDA | WRh | 0.275 | 29 | 0.101
FDA | WRh | 0.275 | 30 | 0.103
FDA | WRh | 0.275 | 31 | 0.104
FDA | WRh | 0.275 | 32 | 0.106
FDA | WRh | 0.3 | 20 | 0.083
FDA | WRh | 0.3 | 21 | 0.087
FDA | WRh | 0.3 | 22 | 0.091
FDA | WRh | 0.3 | 23 | 0.095
FDA | WRh | 0.3 | 24 | 0.099
FDA | WRh | 0.3 | 25 | 0.102
FDA | WRh | 0.3 | 26 | 0.104
FDA | WRh | 0.3 | 27 | 0.106
FDA | WRh | 0.3 | 28 | 0.108
FDA | WRh | 0.3 | 29 | 0.11
FDA | WRh | 0.3 | 30 | 0.112
FDA | WRh | 0.3 | 31 | 0.113
FDA | WRh | 0.3 | 32 | 0.115
FDA | WRh | 0.3 | 33 | 0.117
FDA | WRh | 0.3 | 34 | 0.119
FDA | WRh | 0.3 | 35 | 0.12
FDA | WRh | 0.3 | 36 | 0.122
FDA | WRh | 0.325 | 20 | 0.092
FDA | WRh | 0.325 | 21 | 0.096
FDA | WRh | 0.325 | 22 | 0.1
FDA | WRh | 0.325 | 23 | 0.104
FDA | WRh | 0.325 | 24 | 0.108
FDA | WRh | 0.325 | 25 | 0.111
FDA | WRh | 0.325 | 26 | 0.114
FDA | WRh | 0.325 | 27 | 0.115
FDA | WRh | 0.325 | 28 | 0.117
FDA | WRh | 0.325 | 29 | 0.119
FDA | WRh | 0.325 | 30 | 0.12
FDA | WRh | 0.325 | 31 | 0.122
FDA | WRh | 0.325 | 32 | 0.124
FDA | WRh | 0.325 | 33 | 0.125
FDA | WRh | 0.325 | 34 | 0.127
FDA | WRh | 0.325 | 35 | 0.129
FDA | WRh | 0.325 | 36 | 0.13
FDA | WRh | 0.325 | 37 | 0.132
FDA | WRh | 0.325 | 38 | 0.134
FDA | WRh | 0.325 | 39 | 0.136
FDA | WRh | 0.325 | 40 | 0.137
FDA | WRh | 0.325 | 41 | 0.139
FDA | WRh | 0.325 | 42 | 0.141
FDA | WRh | 0.35 | 20 | 0.1
FDA | WRh | 0.35 | 21 | 0.105
FDA | WRh | 0.35 | 22 | 0.109
FDA | WRh | 0.35 | 23 | 0.113
FDA | WRh | 0.35 | 24 | 0.117
FDA | WRh | 0.35 | 25 | 0.12
FDA | WRh | 0.35 | 26 | 0.122
FDA | WRh | 0.35 | 27 | 0.124
FDA | WRh | 0.35 | 28 | 0.126
FDA | WRh | 0.35 | 29 | 0.127
FDA | WRh | 0.35 | 30 | 0.129
FDA | WRh | 0.35 | 31 | 0.13
FDA | WRh | 0.35 | 32 | 0.132
FDA | WRh | 0.35 | 33 | 0.133
FDA | WRh | 0.35 | 34 | 0.135
FDA | WRh | 0.35 | 35 | 0.137
FDA | WRh | 0.35 | 36 | 0.138
FDA | WRh | 0.35 | 37 | 0.14
FDA | WRh | 0.35 | 38 | 0.142
FDA | WRh | 0.35 | 39 | 0.143
FDA | WRh | 0.35 | 40 | 0.145
FDA | WRh | 0.35 | 41 | 0.147
FDA | WRh | 0.35 | 42 | 0.149
FDA | WRh | 0.35 | 43 | 0.15
FDA | WRh | 0.35 | 44 | 0.152
FDA | WRh | 0.35 | 45 | 0.153
FDA | WRh | 0.35 | 46 | 0.155
FDA | WRh | 0.35 | 47 | 0.157
FDA | WRh | 0.375 | 20 | 0.11
FDA | WRh | 0.375 | 21 | 0.114
FDA | WRh | 0.375 | 22 | 0.118
FDA | WRh | 0.375 | 23 | 0.122
FDA | WRh | 0.375 | 24 | 0.126
FDA | WRh | 0.375 | 25 | 0.129
FDA | WRh | 0.375 | 26 | 0.131
FDA | WRh | 0.375 | 27 | 0.133
FDA | WRh | 0.375 | 28 | 0.134
FDA | WRh | 0.375 | 29 | 0.136
FDA | WRh | 0.375 | 30 | 0.137
FDA | WRh | 0.375 | 31 | 0.138
FDA | WRh | 0.375 | 32 | 0.14
FDA | WRh | 0.375 | 33 | 0.141
FDA | WRh | 0.375 | 34 | 0.143
FDA | WRh | 0.375 | 35 | 0.144
FDA | WRh | 0.375 | 36 | 0.146
FDA | WRh | 0.375 | 37 | 0.148
FDA | WRh | 0.375 | 38 | 0.149
FDA | WRh | 0.375 | 39 | 0.151
FDA | WRh | 0.375 | 40 | 0.153
FDA | WRh | 0.375 | 41 | 0.154
FDA | WRh | 0.375 | 42 | 0.156
FDA | WRh | 0.375 | 43 | 0.158
FDA | WRh | 0.375 | 44 | 0.159
FDA | WRh | 0.375 | 45 | 0.161
FDA | WRh | 0.375 | 46 | 0.163
FDA | WRh | 0.375 | 47 | 0.164
FDA | WRh | 0.375 | 48 | 0.166
FDA | WRh | 0.375 | 49 | 0.167
FDA | WRh | 0.4 | 20 | 0.119
FDA | WRh | 0.4 | 21 | 0.123
FDA | WRh | 0.4 | 22 | 0.127
FDA | WRh | 0.4 | 23 | 0.131
FDA | WRh | 0.4 | 24 | 0.135
FDA | WRh | 0.4 | 25 | 0.138
FDA | WRh | 0.4 | 26 | 0.14
FDA | WRh | 0.4 | 27 | 0.142
FDA | WRh | 0.4 | 28 | 0.143
FDA | WRh | 0.4 | 29 | 0.144
FDA | WRh | 0.4 | 30 | 0.145
FDA | WRh | 0.4 | 31 | 0.147
FDA | WRh | 0.4 | 32 | 0.148
FDA | WRh | 0.4 | 33 | 0.149
FDA | WRh | 0.4 | 34 | 0.151
FDA | WRh | 0.4 | 35 | 0.152
FDA | WRh | 0.4 | 36 | 0.154
FDA | WRh | 0.4 | 37 | 0.155
FDA | WRh | 0.4 | 38 | 0.157
FDA | WRh | 0.4 | 39 | 0.159
FDA | WRh | 0.4 | 40 | 0.16
FDA | WRh | 0.4 | 41 | 0.162
FDA | WRh | 0.4 | 42 | 0.164
FDA | WRh | 0.4 | 43 | 0.165
FDA | WRh | 0.4 | 44 | 0.167
FDA | WRh | 0.4 | 45 | 0.169
FDA | WRh | 0.4 | 46 | 0.17
FDA | WRh | 0.4 | 47 | 0.172
FDA | WRh | 0.4 | 48 | 0.173
FDA | WRh | 0.4 | 49 | 0.175
FDA | WRh | 0.425 | 20 | 0.128
FDA | WRh | 0.425 | 21 | 0.132
FDA | WRh | 0.425 | 22 | 0.136
FDA | WRh | 0.425 | 23 | 0.14
FDA | WRh | 0.425 | 24 | 0.144
FDA | WRh | 0.425 | 25 | 0.147
FDA | WRh | 0.425 | 26 | 0.149
FDA | WRh | 0.425 | 27 | 0.15
FDA | WRh | 0.425 | 28 | 0.151
FDA | WRh | 0.425 | 29 | 0.153
FDA | WRh | 0.425 | 30 | 0.154
FDA | WRh | 0.425 | 31 | 0.155
FDA | WRh | 0.425 | 32 | 0.156
FDA | WRh | 0.425 | 33 | 0.157
FDA | WRh | 0.425 | 34 | 0.159
FDA | WRh | 0.425 | 35 | 0.16
FDA | WRh | 0.425 | 36 | 0.161
FDA | WRh | 0.425 | 37 | 0.163
FDA | WRh | 0.425 | 38 | 0.165
FDA | WRh | 0.425 | 39 | 0.166
FDA | WRh | 0.425 | 40 | 0.168
FDA | WRh | 0.425 | 41 | 0.169
FDA | WRh | 0.425 | 42 | 0.171
FDA | WRh | 0.425 | 43 | 0.173
FDA | WRh | 0.425 | 44 | 0.174
FDA | WRh | 0.425 | 45 | 0.176
FDA | WRh | 0.425 | 46 | 0.177
FDA | WRh | 0.425 | 47 | 0.179
FDA | WRh | 0.425 | 48 | 0.181
FDA | WRh | 0.425 | 49 | 0.182
FDA | WRh | 0.45 | 20 | 0.137
FDA | WRh | 0.45 | 21 | 0.141
FDA | WRh | 0.45 | 22 | 0.145
FDA | WRh | 0.45 | 23 | 0.149
FDA | WRh | 0.45 | 24 | 0.153
FDA | WRh | 0.45 | 25 | 0.156
FDA | WRh | 0.45 | 26 | 0.158
FDA | WRh | 0.45 | 27 | 0.159
FDA | WRh | 0.45 | 28 | 0.16
FDA | WRh | 0.45 | 29 | 0.161
FDA | WRh | 0.45 | 30 | 0.162
FDA | WRh | 0.45 | 31 | 0.163
FDA | WRh | 0.45 | 32 | 0.164
FDA | WRh | 0.45 | 33 | 0.165
FDA | WRh | 0.45 | 34 | 0.166
FDA | WRh | 0.45 | 35 | 0.168
FDA | WRh | 0.45 | 36 | 0.169
FDA | WRh | 0.45 | 37 | 0.17
FDA | WRh | 0.45 | 38 | 0.172
FDA | WRh | 0.45 | 39 | 0.173
FDA | WRh | 0.45 | 40 | 0.175
FDA | WRh | 0.45 | 41 | 0.177
FDA | WRh | 0.45 | 42 | 0.178
FDA | WRh | 0.45 | 43 | 0.18
FDA | WRh | 0.45 | 44 | 0.181
FDA | WRh | 0.45 | 45 | 0.183
FDA | WRh | 0.45 | 46 | 0.185
FDA | WRh | 0.45 | 47 | 0.186
FDA | WRh | 0.45 | 48 | 0.188
FDA | WRh | 0.45 | 49 | 0.189
FDA | WRh | 0.475 | 20 | 0.147
FDA | WRh | 0.475 | 21 | 0.151
FDA | WRh | 0.475 | 22 | 0.154
FDA | WRh | 0.475 | 23 | 0.158
FDA | WRh | 0.475 | 24 | 0.162
FDA | WRh | 0.475 | 25 | 0.164
FDA | WRh | 0.475 | 26 | 0.166
FDA | WRh | 0.475 | 27 | 0.167
FDA | WRh | 0.475 | 28 | 0.168
FDA | WRh | 0.475 | 29 | 0.169
FDA | WRh | 0.475 | 30 | 0.17
FDA | WRh | 0.475 | 31 | 0.171
FDA | WRh | 0.475 | 32 | 0.172
FDA | WRh | 0.475 | 33 | 0.173
FDA | WRh | 0.475 | 34 | 0.174
FDA | WRh | 0.475 | 35 | 0.175
FDA | WRh | 0.475 | 36 | 0.176
FDA | WRh | 0.475 | 37 | 0.178
FDA | WRh | 0.475 | 38 | 0.179
FDA | WRh | 0.475 | 39 | 0.181
FDA | WRh | 0.475 | 40 | 0.182
FDA | WRh | 0.475 | 41 | 0.184
FDA | WRh | 0.475 | 42 | 0.185
FDA | WRh | 0.475 | 43 | 0.187
FDA | WRh | 0.475 | 44 | 0.189
FDA | WRh | 0.475 | 45 | 0.19
FDA | WRh | 0.475 | 46 | 0.192
FDA | WRh | 0.475 | 47 | 0.193
FDA | WRh | 0.475 | 48 | 0.195
FDA | WRh | 0.475 | 49 | 0.197
FDA | WRh | 0.5 | 20 | 0.156
FDA | WRh | 0.5 | 21 | 0.16
FDA | WRh | 0.5 | 22 | 0.163
FDA | WRh | 0.5 | 23 | 0.167
FDA | WRh | 0.5 | 24 | 0.171
FDA | WRh | 0.5 | 25 | 0.173
FDA | WRh | 0.5 | 26 | 0.175
FDA | WRh | 0.5 | 27 | 0.176
FDA | WRh | 0.5 | 28 | 0.177
FDA | WRh | 0.5 | 29 | 0.177
FDA | WRh | 0.5 | 30 | 0.178
FDA | WRh | 0.5 | 31 | 0.179
FDA | WRh | 0.5 | 32 | 0.18
FDA | WRh | 0.5 | 33 | 0.181
FDA | WRh | 0.5 | 34 | 0.182
FDA | WRh | 0.5 | 35 | 0.183
FDA | WRh | 0.5 | 36 | 0.184
FDA | WRh | 0.5 | 37 | 0.185
FDA | WRh | 0.5 | 38 | 0.187
FDA | WRh | 0.5 | 39 | 0.188
FDA | WRh | 0.5 | 40 | 0.189
FDA | WRh | 0.5 | 41 | 0.191
FDA | WRh | 0.5 | 42 | 0.192
FDA | WRh | 0.5 | 43 | 0.194
FDA | WRh | 0.5 | 44 | 0.196
FDA | WRh | 0.5 | 45 | 0.197
FDA | WRh | 0.5 | 46 | 0.199
FDA | WRh | 0.5 | 47 | 0.2
FDA | WRh | 0.5 | 48 | 0.202
FDA | WRh | 0.5 | 49 | 0.203
FDA | WRh | 0.525 | 21 | 0.169
FDA | WRh | 0.525 | 22 | 0.173
FDA | WRh | 0.525 | 23 | 0.176
FDA | WRh | 0.525 | 24 | 0.179
FDA | WRh | 0.525 | 25 | 0.182
FDA | WRh | 0.525 | 26 | 0.183
FDA | WRh | 0.525 | 27 | 0.184
FDA | WRh | 0.525 | 28 | 0.185
FDA | WRh | 0.525 | 29 | 0.186
FDA | WRh | 0.525 | 30 | 0.186
FDA | WRh | 0.525 | 31 | 0.187
FDA | WRh | 0.525 | 32 | 0.188
FDA | WRh | 0.525 | 33 | 0.189
FDA | WRh | 0.525 | 34 | 0.189
FDA | WRh | 0.525 | 35 | 0.19
FDA | WRh | 0.525 | 36 | 0.191
FDA | WRh | 0.525 | 37 | 0.193
FDA | WRh | 0.525 | 38 | 0.194
FDA | WRh | 0.525 | 39 | 0.195
FDA | WRh | 0.525 | 40 | 0.196
FDA | WRh | 0.525 | 41 | 0.198
FDA | WRh | 0.525 | 42 | 0.2
FDA | WRh | 0.525 | 43 | 0.201
FDA | WRh | 0.525 | 44 | 0.203
FDA | WRh | 0.525 | 45 | 0.204
FDA | WRh | 0.525 | 46 | 0.206
FDA | WRh | 0.525 | 47 | 0.207
FDA | WRh | 0.525 | 48 | 0.209
FDA | WRh | 0.525 | 49 | 0.21
FDA | WRh | 0.55 | 21 | 0.178
FDA | WRh | 0.55 | 22 | 0.181
FDA | WRh | 0.55 | 23 | 0.185
FDA | WRh | 0.55 | 24 | 0.188
FDA | WRh | 0.55 | 25 | 0.19
FDA | WRh | 0.55 | 26 | 0.192
FDA | WRh | 0.55 | 27 | 0.193
FDA | WRh | 0.55 | 28 | 0.193
FDA | WRh | 0.55 | 29 | 0.194
FDA | WRh | 0.55 | 30 | 0.194
FDA | WRh | 0.55 | 31 | 0.195
FDA | WRh | 0.55 | 32 | 0.196
FDA | WRh | 0.55 | 33 | 0.196
FDA | WRh | 0.55 | 34 | 0.197
FDA | WRh | 0.55 | 35 | 0.198
FDA | WRh | 0.55 | 36 | 0.199
FDA | WRh | 0.55 | 37 | 0.2
FDA | WRh | 0.55 | 38 | 0.201
FDA | WRh | 0.55 | 39 | 0.202
FDA | WRh | 0.55 | 40 | 0.204
FDA | WRh | 0.55 | 41 | 0.205
FDA | WRh | 0.55 | 42 | 0.207
FDA | WRh | 0.55 | 43 | 0.208
FDA | WRh | 0.55 | 44 | 0.21
FDA | WRh | 0.55 | 45 | 0.211
FDA | WRh | 0.55 | 46 | 0.213
FDA | WRh | 0.55 | 47 | 0.214
FDA | WRh | 0.55 | 48 | 0.216
FDA | WRh | 0.55 | 49 | 0.217
FDA | WRh | 0.575 | 22 | 0.19
FDA | WRh | 0.575 | 23 | 0.194
FDA | WRh | 0.575 | 24 | 0.197
FDA | WRh | 0.575 | 25 | 0.199
FDA | WRh | 0.575 | 26 | 0.2
FDA | WRh | 0.575 | 27 | 0.201
FDA | WRh | 0.575 | 28 | 0.202
FDA | WRh | 0.575 | 29 | 0.202
FDA | WRh | 0.575 | 30 | 0.203
FDA | WRh | 0.575 | 31 | 0.203
FDA | WRh | 0.575 | 32 | 0.204
FDA | WRh | 0.575 | 33 | 0.204
FDA | WRh | 0.575 | 34 | 0.205
FDA | WRh | 0.575 | 35 | 0.206
FDA | WRh | 0.575 | 36 | 0.206
FDA | WRh | 0.575 | 37 | 0.207
FDA | WRh | 0.575 | 38 | 0.208
FDA | WRh | 0.575 | 39 | 0.21
FDA | WRh | 0.575 | 40 | 0.211
FDA | WRh | 0.575 | 41 | 0.212
FDA | WRh | 0.575 | 42 | 0.214
FDA | WRh | 0.575 | 43 | 0.215
FDA | WRh | 0.575 | 44 | 0.216
FDA | WRh | 0.575 | 45 | 0.218
FDA | WRh | 0.575 | 46 | 0.219
FDA | WRh | 0.575 | 47 | 0.221
FDA | WRh | 0.575 | 48 | 0.223
FDA | WRh | 0.575 | 49 | 0.224
FDA | WRh | 0.6 | 22 | 0.199
FDA | WRh | 0.6 | 23 | 0.202
FDA | WRh | 0.6 | 24 | 0.205
FDA | WRh | 0.6 | 25 | 0.207
FDA | WRh | 0.6 | 26 | 0.209
FDA | WRh | 0.6 | 27 | 0.209
FDA | WRh | 0.6 | 28 | 0.21
FDA | WRh | 0.6 | 29 | 0.21
FDA | WRh | 0.6 | 30 | 0.211
FDA | WRh | 0.6 | 31 | 0.211
FDA | WRh | 0.6 | 32 | 0.211
FDA | WRh | 0.6 | 33 | 0.212
FDA | WRh | 0.6 | 34 | 0.213
FDA | WRh | 0.6 | 35 | 0.213
FDA | WRh | 0.6 | 36 | 0.214
FDA | WRh | 0.6 | 37 | 0.215
FDA | WRh | 0.6 | 38 | 0.216
FDA | WRh | 0.6 | 39 | 0.217
FDA | WRh | 0.6 | 40 | 0.218
FDA | WRh | 0.6 | 41 | 0.219
FDA | WRh | 0.6 | 42 | 0.221
FDA | WRh | 0.6 | 43 | 0.222
FDA | WRh | 0.6 | 44 | 0.223
FDA | WRh | 0.6 | 45 | 0.225
FDA | WRh | 0.6 | 46 | 0.226
FDA | WRh | 0.6 | 47 | 0.228
FDA | WRh | 0.6 | 48 | 0.229
FDA | WRh | 0.6 | 49 | 0.231
FDA | WRh | 0.625 | 22 | 0.208
FDA | WRh | 0.625 | 23 | 0.211
FDA | WRh | 0.625 | 24 | 0.214
FDA | WRh | 0.625 | 25 | 0.216
FDA | WRh | 0.625 | 26 | 0.217
FDA | WRh | 0.625 | 27 | 0.217
FDA | WRh | 0.625 | 28 | 0.218
FDA | WRh | 0.625 | 29 | 0.218
FDA | WRh | 0.625 | 30 | 0.218
FDA | WRh | 0.625 | 31 | 0.219
FDA | WRh | 0.625 | 32 | 0.219
FDA | WRh | 0.625 | 33 | 0.22
FDA | WRh | 0.625 | 34 | 0.22
FDA | WRh | 0.625 | 35 | 0.221
FDA | WRh | 0.625 | 36 | 0.221
FDA | WRh | 0.625 | 37 | 0.222
FDA | WRh | 0.625 | 38 | 0.223
FDA | WRh | 0.625 | 39 | 0.224
FDA | WRh | 0.625 | 40 | 0.225
FDA | WRh | 0.625 | 41 | 0.226
FDA | WRh | 0.625 | 42 | 0.228
FDA | WRh | 0.625 | 43 | 0.229
FDA | WRh | 0.625 | 44 | 0.23
FDA | WRh | 0.625 | 45 | 0.232
FDA | WRh | 0.625 | 46 | 0.233
FDA | WRh | 0.625 | 47 | 0.235
FDA | WRh | 0.625 | 48 | 0.236
FDA | WRh | 0.625 | 49 | 0.238
FDA | WRh | 0.65 | 23 | 0.219
FDA | WRh | 0.65 | 24 | 0.222
FDA | WRh | 0.65 | 25 | 0.224
FDA | WRh | 0.65 | 26 | 0.225
FDA | WRh | 0.65 | 27 | 0.226
FDA | WRh | 0.65 | 28 | 0.226
FDA | WRh | 0.65 | 29 | 0.226
FDA | WRh | 0.65 | 30 | 0.227
FDA | WRh | 0.65 | 31 | 0.227
FDA | WRh | 0.65 | 32 | 0.227
FDA | WRh | 0.65 | 33 | 0.227
FDA | WRh | 0.65 | 34 | 0.228
FDA | WRh | 0.65 | 35 | 0.228
FDA | WRh | 0.65 | 36 | 0.229
FDA | WRh | 0.65 | 37 | 0.229
FDA | WRh | 0.65 | 38 | 0.23
FDA | WRh | 0.65 | 39 | 0.231
FDA | WRh | 0.65 | 40 | 0.232
FDA | WRh | 0.65 | 41 | 0.233
FDA | WRh | 0.65 | 42 | 0.234
FDA | WRh | 0.65 | 43 | 0.236
FDA | WRh | 0.65 | 44 | 0.237
FDA | WRh | 0.65 | 45 | 0.238
FDA | WRh | 0.65 | 46 | 0.24
FDA | WRh | 0.65 | 47 | 0.241
FDA | WRh | 0.65 | 48 | 0.243
FDA | WRh | 0.65 | 49 | 0.244
FDA | WRh | 0.675 | 25 | 0.232
FDA | WRh | 0.675 | 26 | 0.233
FDA | WRh | 0.675 | 27 | 0.233
FDA | WRh | 0.675 | 28 | 0.234
FDA | WRh | 0.675 | 29 | 0.234
FDA | WRh | 0.675 | 30 | 0.234
FDA | WRh | 0.675 | 31 | 0.234
FDA | WRh | 0.675 | 32 | 0.235
FDA | WRh | 0.675 | 33 | 0.235
FDA | WRh | 0.675 | 34 | 0.235
FDA | WRh | 0.675 | 35 | 0.236
FDA | WRh | 0.675 | 36 | 0.236
FDA | WRh | 0.675 | 37 | 0.237
FDA | WRh | 0.675 | 38 | 0.237
FDA | WRh | 0.675 | 39 | 0.238
FDA | WRh | 0.675 | 40 | 0.239
FDA | WRh | 0.675 | 41 | 0.24
FDA | WRh | 0.675 | 42 | 0.241
FDA | WRh | 0.675 | 43 | 0.243
FDA | WRh | 0.675 | 44 | 0.244
FDA | WRh | 0.675 | 45 | 0.245
FDA | WRh | 0.675 | 46 | 0.247
FDA | WRh | 0.675 | 47 | 0.248
FDA | WRh | 0.675 | 48 | 0.25
FDA | WRh | 0.675 | 49 | 0.251
FDA | WRh | 0.7 | 27 | 0.241
FDA | WRh | 0.7 | 28 | 0.242
FDA | WRh | 0.7 | 29 | 0.242
FDA | WRh | 0.7 | 30 | 0.242
FDA | WRh | 0.7 | 31 | 0.242
FDA | WRh | 0.7 | 32 | 0.242
FDA | WRh | 0.7 | 33 | 0.243
FDA | WRh | 0.7 | 34 | 0.243
FDA | WRh | 0.7 | 35 | 0.243
FDA | WRh | 0.7 | 36 | 0.244
FDA | WRh | 0.7 | 37 | 0.244
FDA | WRh | 0.7 | 38 | 0.245
FDA | WRh | 0.7 | 39 | 0.245
FDA | WRh | 0.7 | 40 | 0.246
FDA | WRh | 0.7 | 41 | 0.247
FDA | WRh | 0.7 | 42 | 0.248
FDA | WRh | 0.7 | 43 | 0.249
FDA | WRh | 0.7 | 44 | 0.251
FDA | WRh | 0.7 | 45 | 0.252
FDA | WRh | 0.7 | 46 | 0.253
FDA | WRh | 0.7 | 47 | 0.255
FDA | WRh | 0.7 | 48 | 0.256
FDA | WRh | 0.7 | 49 | 0.258
FDA | WRh | 0.725 | 33 | 0.25
FDA | WRh | 0.725 | 34 | 0.25
FDA | WRh | 0.725 | 35 | 0.251
FDA | WRh | 0.725 | 36 | 0.251
FDA | WRh | 0.725 | 37 | 0.251
FDA | WRh | 0.725 | 38 | 0.252
FDA | WRh | 0.725 | 39 | 0.253
FDA | WRh | 0.725 | 40 | 0.253
FDA | WRh | 0.725 | 41 | 0.254
FDA | WRh | 0.725 | 42 | 0.255
FDA | WRh | 0.725 | 43 | 0.256
FDA | WRh | 0.725 | 44 | 0.257
FDA | WRh | 0.725 | 45 | 0.259
FDA | WRh | 0.725 | 46 | 0.26
FDA | WRh | 0.725 | 47 | 0.261
FDA | WRh | 0.725 | 48 | 0.263
FDA | WRh | 0.725 | 49 | 0.265
FDA | WRh | 0.75 | 37 | 0.259
FDA | WRh | 0.75 | 38 | 0.259
FDA | WRh | 0.75 | 39 | 0.26
FDA | WRh | 0.75 | 40 | 0.26
FDA | WRh | 0.75 | 41 | 0.261
FDA | WRh | 0.75 | 42 | 0.262
FDA | WRh | 0.75 | 43 | 0.263
FDA | WRh | 0.75 | 44 | 0.264
FDA | WRh | 0.75 | 45 | 0.265
FDA | WRh | 0.75 | 46 | 0.267
FDA | WRh | 0.75 | 47 | 0.268
FDA | WRh | 0.75 | 48 | 0.27
FDA | WRh | 0.75 | 49 | 0.271
FDA | WRh | 0.775 | 43 | 0.27
FDA | WRh | 0.775 | 44 | 0.271
FDA | WRh | 0.775 | 45 | 0.272
FDA | WRh | 0.775 | 46 | 0.273
FDA | WRh | 0.775 | 47 | 0.275
FDA | WRh | 0.775 | 48 | 0.276
FDA | WRh | 0.775 | 49 | 0.278
FDA | WRh | 0.8 | 48 | 0.283
FDA | WRh | 0.8 | 49 | 0.284
US50 | MoMo | 0.15 | 20 | 0.032
US50 | MoMo | 0.15 | 21 | 0.034
US50 | MoMo | 0.16 | 20 | 0.037
US50 | MoMo | 0.16 | 21 | 0.039
US50 | MoMo | 0.17 | 20 | 0.042
US50 | MoMo | 0.17 | 21 | 0.044
US50 | MoMo | 0.18 | 20 | 0.046
US50 | MoMo | 0.18 | 21 | 0.049
US50 | MoMo | 0.18 | 22 | 0.052
US50 | MoMo | 0.19 | 20 | 0.051
US50 | MoMo | 0.19 | 21 | 0.054
US50 | MoMo | 0.19 | 22 | 0.057
US50 | MoMo | 0.19 | 23 | 0.059
US50 | MoMo | 0.2 | 20 | 0.055
US50 | MoMo | 0.2 | 21 | 0.059
US50 | MoMo | 0.2 | 22 | 0.061
US50 | MoMo | 0.2 | 23 | 0.064
US50 | MoMo | 0.2 | 24 | 0.066
US50 | MoMo | 0.2 | 25 | 0.067
US50 | MoMo | 0.21 | 20 | 0.06
US50 | MoMo | 0.21 | 21 | 0.063
US50 | MoMo | 0.21 | 22 | 0.066
US50 | MoMo | 0.21 | 23 | 0.068
US50 | MoMo | 0.21 | 24 | 0.07
US50 | MoMo | 0.21 | 25 | 0.072
US50 | MoMo | 0.22 | 20 | 0.064
US50 | MoMo | 0.22 | 21 | 0.068
US50 | MoMo | 0.22 | 22 | 0.071
US50 | MoMo | 0.22 | 23 | 0.073
US50 | MoMo | 0.22 | 24 | 0.075
US50 | MoMo | 0.22 | 25 | 0.076
US50 | MoMo | 0.22 | 26 | 0.078
US50 | MoMo | 0.23 | 20 | 0.068
US50 | MoMo | 0.23 | 21 | 0.072
US50 | MoMo | 0.23 | 22 | 0.075
US50 | MoMo | 0.23 | 23 | 0.077
US50 | MoMo | 0.23 | 24 | 0.079
US50 | MoMo | 0.23 | 25 | 0.081
US50 | MoMo | 0.23 | 26 | 0.082
US50 | MoMo | 0.23 | 27 | 0.083
US50 | MoMo | 0.23 | 28 | 0.085
US50 | MoMo | 0.24 | 20 | 0.073
US50 | MoMo | 0.24 | 21 | 0.077
US50 | MoMo | 0.24 | 22 | 0.079
US50 | MoMo | 0.24 | 23 | 0.082
US50 | MoMo | 0.24 | 24 | 0.083
US50 | MoMo | 0.24 | 25 | 0.085
US50 | MoMo | 0.24 | 26 | 0.086
US50 | MoMo | 0.24 | 27 | 0.088
US50 | MoMo | 0.24 | 28 | 0.089
US50 | MoMo | 0.24 | 29 | 0.09
US50 | MoMo | 0.25 | 20 | 0.077
US50 | MoMo | 0.25 | 21 | 0.081
US50 | MoMo | 0.25 | 22 | 0.084
US50 | MoMo | 0.25 | 23 | 0.086
US50 | MoMo | 0.25 | 24 | 0.087
US50 | MoMo | 0.25 | 25 | 0.089
US50 | MoMo | 0.25 | 26 | 0.09
US50 | MoMo | 0.25 | 27 | 0.092
US50 | MoMo | 0.25 | 28 | 0.093
US50 | MoMo | 0.25 | 29 | 0.094
US50 | MoMo | 0.25 | 30 | 0.095
US50 | MoMo | 0.25 | 31 | 0.096
US50 | MoMo | 0.26 | 20 | 0.081
US50 | MoMo | 0.26 | 21 | 0.086
US50 | MoMo | 0.26 | 22 | 0.088
US50 | MoMo | 0.26 | 23 | 0.09
US50 | MoMo | 0.26 | 24 | 0.092
US50 | MoMo | 0.26 | 25 | 0.093
US50 | MoMo | 0.26 | 26 | 0.094
US50 | MoMo | 0.26 | 27 | 0.095
US50 | MoMo | 0.26 | 28 | 0.097
US50 | MoMo | 0.26 | 29 | 0.098
US50 | MoMo | 0.26 | 30 | 0.099
US50 | MoMo | 0.26 | 31 | 0.1
US50 | MoMo | 0.26 | 32 | 0.101
US50 | MoMo | 0.27 | 20 | 0.085
US50 | MoMo | 0.27 | 21 | 0.09
US50 | MoMo | 0.27 | 22 | 0.092
US50 | MoMo | 0.27 | 23 | 0.094
US50 | MoMo | 0.27 | 24 | 0.096
US50 | MoMo | 0.27 | 25 | 0.097
US50 | MoMo | 0.27 | 26 | 0.098
US50 | MoMo | 0.27 | 27 | 0.099
US50 | MoMo | 0.27 | 28 | 0.1
US50 | MoMo | 0.27 | 29 | 0.102
US50 | MoMo | 0.27 | 30 | 0.103
US50 | MoMo | 0.27 | 31 | 0.104
US50 | MoMo | 0.27 | 32 | 0.105
US50 | MoMo | 0.27 | 33 | 0.106
US50 | MoMo | 0.27 | 34 | 0.107
US50 | MoMo | 0.28 | 20 | 0.09
US50 | MoMo | 0.28 | 21 | 0.094
US50 | MoMo | 0.28 | 22 | 0.097
US50 | MoMo | 0.28 | 23 | 0.098
US50 | MoMo | 0.28 | 24 | 0.1
US50 | MoMo | 0.28 | 25 | 0.101
US50 | MoMo | 0.28 | 26 | 0.102
US50 | MoMo | 0.28 | 27 | 0.103
US50 | MoMo | 0.28 | 28 | 0.104
US50 | MoMo | 0.28 | 29 | 0.105
US50 | MoMo | 0.28 | 30 | 0.106
US50 | MoMo | 0.28 | 31 | 0.107
US50 | MoMo | 0.28 | 32 | 0.108
US50 | MoMo | 0.28 | 33 | 0.109
US50 | MoMo | 0.28 | 34 | 0.11
US50 | MoMo | 0.28 | 35 | 0.111
US50 | MoMo | 0.28 | 36 | 0.112
US50 | MoMo | 0.29 | 20 | 0.094
US50 | MoMo | 0.29 | 21 | 0.098
US50 | MoMo | 0.29 | 22 | 0.101
US50 | MoMo | 0.29 | 23 | 0.102
US50 | MoMo | 0.29 | 24 | 0.104
US50 | MoMo | 0.29 | 25 | 0.105
US50 | MoMo | 0.29 | 26 | 0.106
US50 | MoMo | 0.29 | 27 | 0.107
US50 | MoMo | 0.29 | 28 | 0.108
US50 | MoMo | 0.29 | 29 | 0.109
US50 | MoMo | 0.29 | 30 | 0.11
US50 | MoMo | 0.29 | 31 | 0.111
US50 | MoMo | 0.29 | 32 | 0.112
US50 | MoMo | 0.29 | 33 | 0.113
US50 | MoMo | 0.29 | 34 | 0.114
US50 | MoMo | 0.29 | 35 | 0.115
US50 | MoMo | 0.29 | 36 | 0.116
US50 | MoMo | 0.29 | 37 | 0.116
US50 | MoMo | 0.29 | 38 | 0.117
US50 | MoMo | 0.3 | 20 | 0.098
US50 | MoMo | 0.3 | 21 | 0.103
US50 | MoMo | 0.3 | 22 | 0.105
US50 | MoMo | 0.3 | 23 | 0.107
US50 | MoMo | 0.3 | 24 | 0.108
US50 | MoMo | 0.3 | 25 | 0.109
US50 | MoMo | 0.3 | 26 | 0.11
US50 | MoMo | 0.3 | 27 | 0.111
US50 | MoMo | 0.3 | 28 | 0.112
US50 | MoMo | 0.3 | 29 | 0.113
US50 | MoMo | 0.3 | 30 | 0.113
US50 | MoMo | 0.3 | 31 | 0.114
US50 | MoMo | 0.3 | 32 | 0.115
US50 | MoMo | 0.3 | 33 | 0.116
US50 | MoMo | 0.3 | 34 | 0.117
US50 | MoMo | 0.3 | 35 | 0.118
US50 | MoMo | 0.3 | 36 | 0.119
US50 | MoMo | 0.3 | 37 | 0.12
US50 | MoMo | 0.3 | 38 | 0.12
US50 | MoMo | 0.3 | 39 | 0.121
US50 | MoMo | 0.3 | 40 | 0.122
US50 | MoMo | 0.3 | 41 | 0.123
US50 | MoMo | 0.3 | 42 | 0.123
US50 | MoMo | 0.31 | 20 | 0.102
US50 | MoMo | 0.31 | 21 | 0.107
US50 | MoMo | 0.31 | 22 | 0.109
US50 | MoMo | 0.31 | 23 | 0.111
US50 | MoMo | 0.31 | 24 | 0.112
US50 | MoMo | 0.31 | 25 | 0.113
US50 | MoMo | 0.31 | 26 | 0.114
US50 | MoMo | 0.31 | 27 | 0.114
US50 | MoMo | 0.31 | 28 | 0.115
US50 | MoMo | 0.31 | 29 | 0.116
US50 | MoMo | 0.31 | 30 | 0.117
US50 | MoMo | 0.31 | 31 | 0.118
US50 | MoMo | 0.31 | 32 | 0.119
US50 | MoMo | 0.31 | 33 | 0.12
US50 | MoMo | 0.31 | 34 | 0.12
US50 | MoMo | 0.31 | 35 | 0.121
US50 | MoMo | 0.31 | 36 | 0.122
US50 | MoMo | 0.31 | 37 | 0.123
US50 | MoMo | 0.31 | 38 | 0.124
US50 | MoMo | 0.31 | 39 | 0.124
US50 | MoMo | 0.31 | 40 | 0.125
US50 | MoMo | 0.31 | 41 | 0.126
US50 | MoMo | 0.31 | 42 | 0.126
US50 | MoMo | 0.31 | 43 | 0.127
US50 | MoMo | 0.32 | 20 | 0.107
US50 | MoMo | 0.32 | 21 | 0.111
US50 | MoMo | 0.32 | 22 | 0.113
US50 | MoMo | 0.32 | 23 | 0.115
US50 | MoMo | 0.32 | 24 | 0.116
US50 | MoMo | 0.32 | 25 | 0.117
US50 | MoMo | 0.32 | 26 | 0.117
US50 | MoMo | 0.32 | 27 | 0.118
US50 | MoMo | 0.32 | 28 | 0.119
US50 | MoMo | 0.32 | 29 | 0.12
US50 | MoMo | 0.32 | 30 | 0.12
US50 | MoMo | 0.32 | 31 | 0.121
US50 | MoMo | 0.32 | 32 | 0.122
US50 | MoMo | 0.32 | 33 | 0.123
US50 | MoMo | 0.32 | 34 | 0.124
US50 | MoMo | 0.32 | 35 | 0.125
US50 | MoMo | 0.32 | 36 | 0.125
US50 | MoMo | 0.32 | 37 | 0.126
US50 | MoMo | 0.32 | 38 | 0.127
US50 | MoMo | 0.32 | 39 | 0.128
US50 | MoMo | 0.32 | 40 | 0.128
US50 | MoMo | 0.32 | 41 | 0.129
US50 | MoMo | 0.32 | 42 | 0.13
US50 | MoMo | 0.32 | 43 | 0.13
US50 | MoMo | 0.32 | 44 | 0.131
US50 | MoMo | 0.32 | 45 | 0.131
US50 | MoMo | 0.32 | 46 | 0.132
US50 | MoMo | 0.33 | 22 | 0.117
US50 | MoMo | 0.33 | 23 | 0.119
US50 | MoMo | 0.33 | 24 | 0.12
US50 | MoMo | 0.33 | 25 | 0.12
US50 | MoMo | 0.33 | 26 | 0.121
US50 | MoMo | 0.33 | 27 | 0.122
US50 | MoMo | 0.33 | 28 | 0.122
US50 | MoMo | 0.33 | 29 | 0.123
US50 | MoMo | 0.33 | 30 | 0.124
US50 | MoMo | 0.33 | 31 | 0.125
US50 | MoMo | 0.33 | 32 | 0.125
US50 | MoMo | 0.33 | 33 | 0.126
US50 | MoMo | 0.33 | 34 | 0.127
US50 | MoMo | 0.33 | 35 | 0.128
US50 | MoMo | 0.33 | 36 | 0.129
US50 | MoMo | 0.33 | 37 | 0.129
US50 | MoMo | 0.33 | 38 | 0.13
US50 | MoMo | 0.33 | 39 | 0.131
US50 | MoMo | 0.33 | 40 | 0.132
US50 | MoMo | 0.33 | 41 | 0.132
US50 | MoMo | 0.33 | 42 | 0.133
US50 | MoMo | 0.33 | 43 | 0.133
US50 | MoMo | 0.33 | 44 | 0.134
US50 | MoMo | 0.33 | 45 | 0.135
US50 | MoMo | 0.33 | 46 | 0.135
US50 | MoMo | 0.33 | 47 | 0.135
US50 | MoMo | 0.33 | 48 | 0.136
US50 | MoMo | 0.33 | 49 | 0.136
US50 | MoMo | 0.34 | 22 | 0.122
US50 | MoMo | 0.34 | 23 | 0.123
US50 | MoMo | 0.34 | 24 | 0.124
US50 | MoMo | 0.34 | 25 | 0.124
US50 | MoMo | 0.34 | 26 | 0.125
US50 | MoMo | 0.34 | 27 | 0.125
US50 | MoMo | 0.34 | 28 | 0.126
US50 | MoMo | 0.34 | 29 | 0.127
US50 | MoMo | 0.34 | 30 | 0.127
US50 | MoMo | 0.34 | 31 | 0.128
US50 | MoMo | 0.34 | 32 | 0.129
US50 | MoMo | 0.34 | 33 | 0.13
US50 | MoMo | 0.34 | 34 | 0.13
US50 | MoMo | 0.34 | 35 | 0.131
US50 | MoMo | 0.34 | 36 | 0.132
US50 | MoMo | 0.34 | 37 | 0.133
US50 | MoMo | 0.34 | 38 | 0.133
US50 | MoMo | 0.34 | 39 | 0.134
US50 | MoMo | 0.34 | 40 | 0.135
US50 | MoMo | 0.34 | 41 | 0.135
US50 | MoMo | 0.34 | 42 | 0.136
US50 | MoMo | 0.34 | 43 | 0.137
US50 | MoMo | 0.34 | 44 | 0.137
US50 | MoMo | 0.34 | 45 | 0.138
US50 | MoMo | 0.34 | 46 | 0.138
US50 | MoMo | 0.34 | 47 | 0.139
US50 | MoMo | 0.34 | 48 | 0.139
US50 | MoMo | 0.34 | 49 | 0.14
US50 | MoMo | 0.35 | 22 | 0.126
US50 | MoMo | 0.35 | 23 | 0.127
US50 | MoMo | 0.35 | 24 | 0.128
US50 | MoMo | 0.35 | 25 | 0.128
US50 | MoMo | 0.35 | 26 | 0.129
US50 | MoMo | 0.35 | 27 | 0.129
US50 | MoMo | 0.35 | 28 | 0.13
US50 | MoMo | 0.35 | 29 | 0.13
US50 | MoMo | 0.35 | 30 | 0.131
US50 | MoMo | 0.35 | 31 | 0.131
US50 | MoMo | 0.35 | 32 | 0.132
US50 | MoMo | 0.35 | 33 | 0.133
US50 | MoMo | 0.35 | 34 | 0.133
US50 | MoMo | 0.35 | 35 | 0.134
US50 | MoMo | 0.35 | 36 | 0.135
US50 | MoMo | 0.35 | 37 | 0.136
US50 | MoMo | 0.35 | 38 | 0.136
US50 | MoMo | 0.35 | 39 | 0.137
US50 | MoMo | 0.35 | 40 | 0.138
US50 | MoMo | 0.35 | 41 | 0.138
US50 | MoMo | 0.35 | 42 | 0.139
US50 | MoMo | 0.35 | 43 | 0.14
US50 | MoMo | 0.35 | 44 | 0.14
US50 | MoMo | 0.35 | 45 | 0.141
US50 | MoMo | 0.35 | 46 | 0.141
US50 | MoMo | 0.35 | 47 | 0.142
US50 | MoMo | 0.35 | 48 | 0.142
US50 | MoMo | 0.35 | 49 | 0.143
US50 | MoMo | 0.36 | 23 | 0.131
US50 | MoMo | 0.36 | 24 | 0.132
US50 | MoMo | 0.36 | 25 | 0.132
US50 | MoMo | 0.36 | 26 | 0.132
US50 | MoMo | 0.36 | 27 | 0.133
US50 | MoMo | 0.36 | 28 | 0.133
US50 | MoMo | 0.36 | 29 | 0.134
US50 | MoMo | 0.36 | 30 | 0.134
US50 | MoMo | 0.36 | 31 | 0.135
US50 | MoMo | 0.36 | 32 | 0.135
US50 | MoMo | 0.36 | 33 | 0.136
US50 | MoMo | 0.36 | 34 | 0.137
US50 | MoMo | 0.36 | 35 | 0.137
US50 | MoMo | 0.36 | 36 | 0.138
US50 | MoMo | 0.36 | 37 | 0.139
US50 | MoMo | 0.36 | 38 | 0.14
US50 | MoMo | 0.36 | 39 | 0.14
US50 | MoMo | 0.36 | 40 | 0.141
US50 | MoMo | 0.36 | 41 | 0.142
US50 | MoMo | 0.36 | 42 | 0.142
US50 | MoMo | 0.36 | 43 | 0.143
US50 | MoMo | 0.36 | 44 | 0.143
US50 | MoMo | 0.36 | 45 | 0.144
US50 | MoMo | 0.36 | 46 | 0.144
US50 | MoMo | 0.36 | 47 | 0.145
US50 | MoMo | 0.36 | 48 | 0.145
US50 | MoMo | 0.36 | 49 | 0.146
US50 | MoMo | 0.37 | 24 | 0.135
US50 | MoMo | 0.37 | 25 | 0.136
US50 | MoMo | 0.37 | 26 | 0.136
US50 | MoMo | 0.37 | 27 | 0.136
US50 | MoMo | 0.37 | 28 | 0.137
US50 | MoMo | 0.37 | 29 | 0.137
US50 | MoMo | 0.37 | 30 | 0.138
US50 | MoMo | 0.37 | 31 | 0.138
US50 | MoMo | 0.37 | 32 | 0.139
US50 | MoMo | 0.37 | 33 | 0.139
US50 | MoMo | 0.37 | 34 | 0.14
US50 | MoMo | 0.37 | 35 | 0.141
US50 | MoMo | 0.37 | 36 | 0.141
US50 | MoMo | 0.37 | 37 | 0.142
US50 | MoMo | 0.37 | 38 | 0.143
US50 | MoMo | 0.37 | 39 | 0.143
US50 | MoMo | 0.37 | 40 | 0.144
US50 | MoMo | 0.37 | 41 | 0.145
US50 | MoMo | 0.37 | 42 | 0.145
US50 | MoMo | 0.37 | 43 | 0.146
US50 | MoMo | 0.37 | 44 | 0.146
US50 | MoMo | 0.37 | 45 | 0.147
US50 | MoMo | 0.37 | 46 | 0.147
US50 | MoMo | 0.37 | 47 | 0.148
US50 | MoMo | 0.37 | 48 | 0.148
US50 | MoMo | 0.37 | 49 | 0.149
US50 | MoMo | 0.38 | 26 | 0.14
US50 | MoMo | 0.38 | 27 | 0.14
US50 | MoMo | 0.38 | 28 | 0.14
US50 | MoMo | 0.38 | 29 | 0.141
US50 | MoMo | 0.38 | 30 | 0.141
US50 | MoMo | 0.38 | 31 | 0.142
US50 | MoMo | 0.38 | 32 | 0.142
US50 | MoMo | 0.38 | 33 | 0.143
US50 | MoMo | 0.38 | 34 | 0.143
US50 | MoMo | 0.38 | 35 | 0.144
US50 | MoMo | 0.38 | 36 | 0.144
US50 | MoMo | 0.38 | 37 | 0.145
US50 | MoMo | 0.38 | 38 | 0.146
US50 | MoMo | 0.38 | 39 | 0.147
US50 | MoMo | 0.38 | 40 | 0.147
US50 | MoMo | 0.38 | 41 | 0.148
US50 | MoMo | 0.38 | 42 | 0.148
US50 | MoMo | 0.38 | 43 | 0.149
US50 | MoMo | 0.38 | 44 | 0.15
US50 | MoMo | 0.38 | 45 | 0.15
US50 | MoMo | 0.38 | 46 | 0.151
US50 | MoMo | 0.38 | 47 | 0.151
US50 | MoMo | 0.38 | 48 | 0.151
US50 | MoMo | 0.38 | 49 | 0.152
US50 | MoMo | 0.39 | 26 | 0.144
US50 | MoMo | 0.39 | 27 | 0.144
US50 | MoMo | 0.39 | 28 | 0.144
US50 | MoMo | 0.39 | 29 | 0.144
US50 | MoMo | 0.39 | 30 | 0.145
US50 | MoMo | 0.39 | 31 | 0.145
US50 | MoMo | 0.39 | 32 | 0.145
US50 | MoMo | 0.39 | 33 | 0.146
US50 | MoMo | 0.39 | 34 | 0.147
US50 | MoMo | 0.39 | 35 | 0.147
US50 | MoMo | 0.39 | 36 | 0.148
US50 | MoMo | 0.39 | 37 | 0.148
US50 | MoMo | 0.39 | 38 | 0.149
US50 | MoMo | 0.39 | 39 | 0.15
US50 | MoMo | 0.39 | 40 | 0.15
US50 | MoMo | 0.39 | 41 | 0.151
US50 | MoMo | 0.39 | 42 | 0.151
US50 | MoMo | 0.39 | 43 | 0.152
US50 | MoMo | 0.39 | 44 | 0.153
US50 | MoMo | 0.39 | 45 | 0.153
US50 | MoMo | 0.39 | 46 | 0.154
US50 | MoMo | 0.39 | 47 | 0.154
US50 | MoMo | 0.39 | 48 | 0.155
US50 | MoMo | 0.39 | 49 | 0.155
US50 | MoMo | 0.4 | 27 | 0.148
US50 | MoMo | 0.4 | 28 | 0.148
US50 | MoMo | 0.4 | 29 | 0.148
US50 | MoMo | 0.4 | 30 | 0.148
US50 | MoMo | 0.4 | 31 | 0.148
US50 | MoMo | 0.4 | 32 | 0.149
US50 | MoMo | 0.4 | 33 | 0.149
US50 | MoMo | 0.4 | 34 | 0.15
US50 | MoMo | 0.4 | 35 | 0.15
US50 | MoMo | 0.4 | 36 | 0.151
US50 | MoMo | 0.4 | 37 | 0.152
US50 | MoMo | 0.4 | 38 | 0.152
US50 | MoMo | 0.4 | 39 | 0.153
US50 | MoMo | 0.4 | 40 | 0.153
US50 | MoMo | 0.4 | 41 | 0.154
US50 | MoMo | 0.4 | 42 | 0.155
US50 | MoMo | 0.4 | 43 | 0.155
US50 | MoMo | 0.4 | 44 | 0.156
US50 | MoMo | 0.4 | 45 | 0.156
US50 | MoMo | 0.4 | 46 | 0.157
US50 | MoMo | 0.4 | 47 | 0.157
US50 | MoMo | 0.4 | 48 | 0.158
US50 | MoMo | 0.4 | 49 | 0.158
US50 | MoMo | 0.41 | 29 | 0.152
US50 | MoMo | 0.41 | 30 | 0.152
US50 | MoMo | 0.41 | 31 | 0.152
US50 | MoMo | 0.41 | 32 | 0.152
US50 | MoMo | 0.41 | 33 | 0.153
US50 | MoMo | 0.41 | 34 | 0.153
US50 | MoMo | 0.41 | 35 | 0.154
US50 | MoMo | 0.41 | 36 | 0.154
US50 | MoMo | 0.41 | 37 | 0.155
US50 | MoMo | 0.41 | 38 | 0.155
US50 | MoMo | 0.41 | 39 | 0.156
US50 | MoMo | 0.41 | 40 | 0.157
US50 | MoMo | 0.41 | 41 | 0.157
US50 | MoMo | 0.41 | 42 | 0.158
US50 | MoMo | 0.41 | 43 | 0.158
US50 | MoMo | 0.41 | 44 | 0.159
US50 | MoMo | 0.41 | 45 | 0.159
US50 | MoMo | 0.41 | 46 | 0.16
US50 | MoMo | 0.41 | 47 | 0.16
US50 | MoMo | 0.41 | 48 | 0.161
US50 | MoMo | 0.41 | 49 | 0.161
US50 | MoMo | 0.42 | 30 | 0.155
US50 | MoMo | 0.42 | 31 | 0.156
US50 | MoMo | 0.42 | 32 | 0.156
US50 | MoMo | 0.42 | 33 | 0.156
US50 | MoMo | 0.42 | 34 | 0.157
US50 | MoMo | 0.42 | 35 | 0.157
US50 | MoMo | 0.42 | 36 | 0.158
US50 | MoMo | 0.42 | 37 | 0.158
US50 | MoMo | 0.42 | 38 | 0.159
US50 | MoMo | 0.42 | 39 | 0.159
US50 | MoMo | 0.42 | 40 | 0.16
US50 | MoMo | 0.42 | 41 | 0.16
US50 | MoMo | 0.42 | 42 | 0.161
US50 | MoMo | 0.42 | 43 | 0.161
US50 | MoMo | 0.42 | 44 | 0.162
US50 | MoMo | 0.42 | 45 | 0.163
US50 | MoMo | 0.42 | 46 | 0.163
US50 | MoMo | 0.42 | 47 | 0.164
US50 | MoMo | 0.42 | 48 | 0.164
US50 | MoMo | 0.42 | 49 | 0.165
US50 | MoMo | 0.43 | 32 | 0.159
US50 | MoMo | 0.43 | 33 | 0.16
US50 | MoMo | 0.43 | 34 | 0.16
US50 | MoMo | 0.43 | 35 | 0.161
US50 | MoMo | 0.43 | 36 | 0.161
US50 | MoMo | 0.43 | 37 | 0.161
US50 | MoMo | 0.43 | 38 | 0.162
US50 | MoMo | 0.43 | 39 | 0.163
US50 | MoMo | 0.43 | 40 | 0.163
US50 | MoMo | 0.43 | 41 | 0.164
US50 | MoMo | 0.43 | 42 | 0.164
US50 | MoMo | 0.43 | 43 | 0.165
US50 | MoMo | 0.43 | 44 | 0.165
US50 | MoMo | 0.43 | 45 | 0.166
US50 | MoMo | 0.43 | 46 | 0.166
US50 | MoMo | 0.43 | 47 | 0.167
US50 | MoMo | 0.43 | 48 | 0.167
US50 | MoMo | 0.43 | 49 | 0.168
US50 | MoMo | 0.44 | 33 | 0.163
US50 | MoMo | 0.44 | 34 | 0.164
US50 | MoMo | 0.44 | 35 | 0.164
US50 | MoMo | 0.44 | 36 | 0.164
US50 | MoMo | 0.44 | 37 | 0.165
US50 | MoMo | 0.44 | 38 | 0.165
US50 | MoMo | 0.44 | 39 | 0.166
US50 | MoMo | 0.44 | 40 | 0.166
US50 | MoMo | 0.44 | 41 | 0.167
US50 | MoMo | 0.44 | 42 | 0.167
US50 | MoMo | 0.44 | 43 | 0.168
US50 | MoMo | 0.44 | 44 | 0.169
US50 | MoMo | 0.44 | 45 | 0.169
US50 | MoMo | 0.44 | 46 | 0.17
US50 | MoMo | 0.44 | 47 | 0.17
US50 | MoMo | 0.44 | 48 | 0.171
US50 | MoMo | 0.44 | 49 | 0.171
US50 | MoMo | 0.45 | 35 | 0.168
US50 | MoMo | 0.45 | 36 | 0.168
US50 | MoMo | 0.45 | 37 | 0.168
US50 | MoMo | 0.45 | 38 | 0.169
US50 | MoMo | 0.45 | 39 | 0.169
US50 | MoMo | 0.45 | 40 | 0.17
US50 | MoMo | 0.45 | 41 | 0.17
US50 | MoMo | 0.45 | 42 | 0.171
US50 | MoMo | 0.45 | 43 | 0.172
US50 | MoMo | 0.45 | 44 | 0.172
US50 | MoMo | 0.45 | 45 | 0.173
US50 | MoMo | 0.45 | 46 | 0.173
US50 | MoMo | 0.45 | 47 | 0.174
US50 | MoMo | 0.45 | 48 | 0.174
US50 | MoMo | 0.45 | 49 | 0.175
US50 | MoMo | 0.46 | 37 | 0.172
US50 | MoMo | 0.46 | 38 | 0.173
US50 | MoMo | 0.46 | 39 | 0.173
US50 | MoMo | 0.46 | 40 | 0.173
US50 | MoMo | 0.46 | 41 | 0.174
US50 | MoMo | 0.46 | 42 | 0.174
US50 | MoMo | 0.46 | 43 | 0.175
US50 | MoMo | 0.46 | 44 | 0.176
US50 | MoMo | 0.46 | 45 | 0.176
US50 | MoMo | 0.46 | 46 | 0.177
US50 | MoMo | 0.46 | 47 | 0.177
US50 | MoMo | 0.46 | 48 | 0.178
US50 | MoMo | 0.46 | 49 | 0.178
US50 | MoMo | 0.47 | 39 | 0.177
US50 | MoMo | 0.47 | 40 | 0.177
US50 | MoMo | 0.47 | 41 | 0.178
US50 | MoMo | 0.47 | 42 | 0.178
US50 | MoMo | 0.47 | 43 | 0.179
US50 | MoMo | 0.47 | 44 | 0.179
US50 | MoMo | 0.47 | 45 | 0.18
US50 | MoMo | 0.47 | 46 | 0.18
US50 | MoMo | 0.47 | 47 | 0.181
US50 | MoMo | 0.47 | 48 | 0.181
US50 | MoMo | 0.47 | 49 | 0.182
US50 | MoMo | 0.48 | 43 | 0.183
US50 | MoMo | 0.48 | 44 | 0.183
US50 | MoMo | 0.48 | 45 | 0.184
US50 | MoMo | 0.48 | 46 | 0.184
US50 | MoMo | 0.48 | 47 | 0.185
US50 | MoMo | 0.48 | 48 | 0.185
US50 | MoMo | 0.48 | 49 | 0.186
US50 | MoMo | 0.49 | 44 | 0.187
US50 | MoMo | 0.49 | 45 | 0.188
US50 | MoMo | 0.49 | 46 | 0.188
US50 | MoMo | 0.49 | 47 | 0.189
US50 | MoMo | 0.49 | 48 | 0.19
US50 | MoMo | 0.49 | 49 | 0.19
US50 | MoMo | 0.5 | 47 | 0.193
US50 | MoMo | 0.5 | 48 | 0.194
US50 | MoMo | 0.5 | 49 | 0.195
US50 | MoRh | 0.15 | 20 | 0.032
US50 | MoRh | 0.15 | 21 | 0.034
US50 | MoRh | 0.15 | 22 | 0.036
US50 | MoRh | 0.165 | 20 | 0.039
US50 | MoRh | 0.165 | 21 | 0.042
US50 | MoRh | 0.165 | 22 | 0.044
US50 | MoRh | 0.18 | 20 | 0.046
US50 | MoRh | 0.18 | 21 | 0.049
US50 | MoRh | 0.18 | 22 | 0.052
US50 | MoRh | 0.195 | 20 | 0.053
US50 | MoRh | 0.195 | 21 | 0.056
US50 | MoRh | 0.195 | 22 | 0.06
US50 | MoRh | 0.21 | 20 | 0.06
US50 | MoRh | 0.21 | 21 | 0.063
US50 | MoRh | 0.21 | 22 | 0.067
US50 | MoRh | 0.21 | 23 | 0.07
US50 | MoRh | 0.21 | 24 | 0.072
US50 | MoRh | 0.225 | 20 | 0.066
US50 | MoRh | 0.225 | 21 | 0.07
US50 | MoRh | 0.225 | 22 | 0.074
US50 | MoRh | 0.225 | 23 | 0.077
US50 | MoRh | 0.225 | 24 | 0.079
US50 | MoRh | 0.225 | 25 | 0.081
US50 | MoRh | 0.225 | 26 | 0.083
US50 | MoRh | 0.225 | 27 | 0.084
US50 | MoRh | 0.24 | 20 | 0.073
US50 | MoRh | 0.24 | 21 | 0.077
US50 | MoRh | 0.24 | 22 | 0.08
US50 | MoRh | 0.24 | 23 | 0.084
US50 | MoRh | 0.24 | 24 | 0.086
US50 | MoRh | 0.24 | 25 | 0.088
US50 | MoRh | 0.24 | 26 | 0.09
US50 | MoRh | 0.24 | 27 | 0.091
US50 | MoRh | 0.255 | 20 | 0.079
US50 | MoRh | 0.255 | 21 | 0.083
US50 | MoRh | 0.255 | 22 | 0.087
US50 | MoRh | 0.255 | 23 | 0.09
US50 | MoRh | 0.255 | 24 | 0.093
US50 | MoRh | 0.255 | 25 | 0.095
US50 | MoRh | 0.255 | 26 | 0.097
US50 | MoRh | 0.255 | 27 | 0.098
US50 | MoRh | 0.255 | 28 | 0.099
US50 | MoRh | 0.27 | 20 | 0.085
US50 | MoRh | 0.27 | 21 | 0.09
US50 | MoRh | 0.27 | 22 | 0.094
US50 | MoRh | 0.27 | 23 | 0.097
US50 | MoRh | 0.27 | 24 | 0.1
US50 | MoRh | 0.27 | 25 | 0.102
US50 | MoRh | 0.27 | 26 | 0.103
US50 | MoRh | 0.27 | 27 | 0.104
US50 | MoRh | 0.27 | 28 | 0.105
US50 | MoRh | 0.27 | 29 | 0.106
US50 | MoRh | 0.27 | 30 | 0.107
US50 | MoRh | 0.27 | 31 | 0.108
US50 | MoRh | 0.285 | 20 | 0.092
US50 | MoRh | 0.285 | 21 | 0.096
US50 | MoRh | 0.285 | 22 | 0.1
US50 | MoRh | 0.285 | 23 | 0.103
US50 | MoRh | 0.285 | 24 | 0.106
US50 | MoRh | 0.285 | 25 | 0.108
US50 | MoRh | 0.285 | 26 | 0.109
US50 | MoRh | 0.285 | 27 | 0.111
US50 | MoRh | 0.285 | 28 | 0.111
US50 | MoRh | 0.285 | 29 | 0.112
US50 | MoRh | 0.285 | 30 | 0.113
US50 | MoRh | 0.285 | 31 | 0.114
US50 | MoRh | 0.285 | 32 | 0.115
US50 | MoRh | 0.285 | 33 | 0.115
US50 | MoRh | 0.3 | 20 | 0.098
US50 | MoRh | 0.3 | 21 | 0.103
US50 | MoRh | 0.3 | 22 | 0.106
US50 | MoRh | 0.3 | 23 | 0.11
US50 | MoRh | 0.3 | 24 | 0.112
US50 | MoRh | 0.3 | 25 | 0.114
US50 | MoRh | 0.3 | 26 | 0.116
US50 | MoRh | 0.3 | 27 | 0.117
US50 | MoRh | 0.3 | 28 | 0.118
US50 | MoRh | 0.3 | 29 | 0.118
US50 | MoRh | 0.3 | 30 | 0.119
US50 | MoRh | 0.3 | 31 | 0.12
US50 | MoRh | 0.3 | 32 | 0.12
US50 | MoRh | 0.3 | 33 | 0.121
US50 | MoRh | 0.3 | 34 | 0.122
US50 | MoRh | 0.3 | 35 | 0.122
US50 | MoRh | 0.3 | 36 | 0.123
US50 | MoRh | 0.3 | 37 | 0.123
US50 | MoRh | 0.3 | 38 | 0.124
US50 | MoRh | 0.3 | 39 | 0.124
US50 | MoRh | 0.3 | 40 | 0.125
US50 | MoRh | 0.3 | 41 | 0.125
US50 | MoRh | 0.3 | 42 | 0.126
US50 | MoRh | 0.315 | 20 | 0.105
US50 | MoRh | 0.315 | 21 | 0.109
US50 | MoRh | 0.315 | 22 | 0.113
US50 | MoRh | 0.315 | 23 | 0.116
US50 | MoRh | 0.315 | 24 | 0.119
US50 | MoRh | 0.315 | 25 | 0.121
US50 | MoRh | 0.315 | 26 | 0.122
US50 | MoRh | 0.315 | 27 | 0.123
US50 | MoRh | 0.315 | 28 | 0.124
US50 | MoRh | 0.315 | 29 | 0.124
US50 | MoRh | 0.315 | 30 | 0.125
US50 | MoRh | 0.315 | 31 | 0.125
US50 | MoRh | 0.315 | 32 | 0.126
US50 | MoRh | 0.315 | 33 | 0.127
US50 | MoRh | 0.315 | 34 | 0.127
US50 | MoRh | 0.315 | 35 | 0.128
US50 | MoRh | 0.315 | 36 | 0.128
US50 | MoRh | 0.315 | 37 | 0.129
US50 | MoRh | 0.315 | 38 | 0.129
US50 | MoRh | 0.315 | 39 | 0.13
US50 | MoRh | 0.315 | 40 | 0.13
US50 | MoRh | 0.315 | 41 | 0.131
US50 | MoRh | 0.315 | 42 | 0.131
US50 | MoRh | 0.33 | 20 | 0.111
US50 | MoRh | 0.33 | 21 | 0.115
US50 | MoRh | 0.33 | 22 | 0.119
US50 | MoRh | 0.33 | 23 | 0.122
US50 | MoRh | 0.33 | 24 | 0.125
US50 | MoRh | 0.33 | 25 | 0.127
US50 | MoRh | 0.33 | 26 | 0.128
US50 | MoRh | 0.33 | 27 | 0.129
US50 | MoRh | 0.33 | 28 | 0.129
US50 | MoRh | 0.33 | 29 | 0.13
US50 | MoRh | 0.33 | 30 | 0.131
US50 | MoRh | 0.33 | 31 | 0.131
US50 | MoRh | 0.33 | 32 | 0.132
US50 | MoRh | 0.33 | 33 | 0.132
US50 | MoRh | 0.33 | 34 | 0.133
US50 | MoRh | 0.33 | 35 | 0.133
US50 | MoRh | 0.33 | 36 | 0.134
US50 | MoRh | 0.33 | 37 | 0.134
US50 | MoRh | 0.33 | 38 | 0.135
US50 | MoRh | 0.33 | 39 | 0.135
US50 | MoRh | 0.33 | 40 | 0.136
US50 | MoRh | 0.33 | 41 | 0.136
US50 | MoRh | 0.33 | 42 | 0.137
US50 | MoRh | 0.345 | 20 | 0.117
US50 | MoRh | 0.345 | 21 | 0.122
US50 | MoRh | 0.345 | 22 | 0.125
US50 | MoRh | 0.345 | 23 | 0.128
US50 | MoRh | 0.345 | 24 | 0.131
US50 | MoRh | 0.345 | 25 | 0.133
US50 | MoRh | 0.345 | 26 | 0.134
US50 | MoRh | 0.345 | 27 | 0.135
US50 | MoRh | 0.345 | 28 | 0.135
US50 | MoRh | 0.345 | 29 | 0.136
US50 | MoRh | 0.345 | 30 | 0.136
US50 | MoRh | 0.345 | 31 | 0.137
US50 | MoRh | 0.345 | 32 | 0.137
US50 | MoRh | 0.345 | 33 | 0.138
US50 | MoRh | 0.345 | 34 | 0.138
US50 | MoRh | 0.345 | 35 | 0.139
US50 | MoRh | 0.345 | 36 | 0.139
US50 | MoRh | 0.345 | 37 | 0.14
US50 | MoRh | 0.345 | 38 | 0.14
US50 | MoRh | 0.345 | 39 | 0.141
US50 | MoRh | 0.345 | 40 | 0.141
US50 | MoRh | 0.345 | 41 | 0.141
US50 | MoRh | 0.345 | 42 | 0.142
US50 | MoRh | 0.345 | 43 | 0.142
US50 | MoRh | 0.345 | 44 | 0.143
US50 | MoRh | 0.345 | 45 | 0.143
US50 | MoRh | 0.345 | 46 | 0.143
US50 | MoRh | 0.36 | 20 | 0.124
US50 | MoRh | 0.36 | 21 | 0.128
US50 | MoRh | 0.36 | 22 | 0.132
US50 | MoRh | 0.36 | 23 | 0.135
US50 | MoRh | 0.36 | 24 | 0.137
US50 | MoRh | 0.36 | 25 | 0.139
US50 | MoRh | 0.36 | 26 | 0.14
US50 | MoRh | 0.36 | 27 | 0.141
US50 | MoRh | 0.36 | 28 | 0.141
US50 | MoRh | 0.36 | 29 | 0.142
US50 | MoRh | 0.36 | 30 | 0.142
US50 | MoRh | 0.36 | 31 | 0.142
US50 | MoRh | 0.36 | 32 | 0.143
US50 | MoRh | 0.36 | 33 | 0.143
US50 | MoRh | 0.36 | 34 | 0.144
US50 | MoRh | 0.36 | 35 | 0.144
US50 | MoRh | 0.36 | 36 | 0.144
US50 | MoRh | 0.36 | 37 | 0.145
US50 | MoRh | 0.36 | 38 | 0.145
US50 | MoRh | 0.36 | 39 | 0.146
US50 | MoRh | 0.36 | 40 | 0.146
US50 | MoRh | 0.36 | 41 | 0.147
US50 | MoRh | 0.36 | 42 | 0.147
US50 | MoRh | 0.36 | 43 | 0.147
US50 | MoRh | 0.36 | 44 | 0.148
US50 | MoRh | 0.36 | 45 | 0.148
US50 | MoRh | 0.36 | 46 | 0.148
US50 | MoRh | 0.36 | 47 | 0.149
US50 | MoRh | 0.36 | 48 | 0.149
US50 | MoRh | 0.36 | 49 | 0.149
US50 | MoRh | 0.375 | 20 | 0.13
US50 | MoRh | 0.375 | 21 | 0.134
US50 | MoRh | 0.375 | 22 | 0.138
US50 | MoRh | 0.375 | 23 | 0.141
US50 | MoRh | 0.375 | 24 | 0.143
US50 | MoRh | 0.375 | 25 | 0.145
US50 | MoRh | 0.375 | 26 | 0.146
US50 | MoRh | 0.375 | 27 | 0.147
US50 | MoRh | 0.375 | 28 | 0.147
US50 | MoRh | 0.375 | 29 | 0.147
US50 | MoRh | 0.375 | 30 | 0.148
US50 | MoRh | 0.375 | 31 | 0.148
US50 | MoRh | 0.375 | 32 | 0.148
US50 | MoRh | 0.375 | 33 | 0.149
US50 | MoRh | 0.375 | 34 | 0.149
US50 | MoRh | 0.375 | 35 | 0.149
US50 | MoRh | 0.375 | 36 | 0.15
US50 | MoRh | 0.375 | 37 | 0.15
US50 | MoRh | 0.375 | 38 | 0.151
US50 | MoRh | 0.375 | 39 | 0.151
US50 | MoRh | 0.375 | 40 | 0.151
US50 | MoRh | 0.375 | 41 | 0.152
US50 | MoRh | 0.375 | 42 | 0.152
US50 | MoRh | 0.375 | 43 | 0.153
US50 | MoRh | 0.375 | 44 | 0.153
US50 | MoRh | 0.375 | 45 | 0.153
US50 | MoRh | 0.375 | 46 | 0.154
US50 | MoRh | 0.375 | 47 | 0.154
US50 | MoRh | 0.375 | 48 | 0.154
US50 | MoRh | 0.375 | 49 | 0.155
US50 | MoRh | 0.39 | 20 | 0.137
US50 | MoRh | 0.39 | 21 | 0.141
US50 | MoRh | 0.39 | 22 | 0.144
US50 | MoRh | 0.39 | 23 | 0.147
US50 | MoRh | 0.39 | 24 | 0.15
US50 | MoRh | 0.39 | 25 | 0.151
US50 | MoRh | 0.39 | 26 | 0.152
US50 | MoRh | 0.39 | 27 | 0.153
US50 | MoRh | 0.39 | 28 | 0.153
US50 | MoRh | 0.39 | 29 | 0.153
US50 | MoRh | 0.39 | 30 | 0.153
US50 | MoRh | 0.39 | 31 | 0.153
US50 | MoRh | 0.39 | 32 | 0.154
US50 | MoRh | 0.39 | 33 | 0.154
US50 | MoRh | 0.39 | 34 | 0.154
US50 | MoRh | 0.39 | 35 | 0.155
US50 | MoRh | 0.39 | 36 | 0.155
US50 | MoRh | 0.39 | 37 | 0.155
US50 | MoRh | 0.39 | 38 | 0.156
US50 | MoRh | 0.39 | 39 | 0.156
US50 | MoRh | 0.39 | 40 | 0.157
US50 | MoRh | 0.39 | 41 | 0.157
US50 | MoRh | 0.39 | 42 | 0.157
US50 | MoRh | 0.39 | 43 | 0.158
US50 | MoRh | 0.39 | 44 | 0.158
US50 | MoRh | 0.39 | 45 | 0.158
US50 | MoRh | 0.39 | 46 | 0.159
US50 | MoRh | 0.39 | 47 | 0.159
US50 | MoRh | 0.39 | 48 | 0.159
US50 | MoRh | 0.39 | 49 | 0.16
US50 | MoRh | 0.405 | 20 | 0.143
US50 | MoRh | 0.405 | 21 | 0.147
US50 | MoRh | 0.405 | 22 | 0.15
US50 | MoRh | 0.405 | 23 | 0.153
US50 | MoRh | 0.405 | 24 | 0.156
US50 | MoRh | 0.405 | 25 | 0.157
US50 | MoRh | 0.405 | 26 | 0.158
US50 | MoRh | 0.405 | 27 | 0.158
US50 | MoRh | 0.405 | 28 | 0.159
US50 | MoRh | 0.405 | 29 | 0.159
US50 | MoRh | 0.405 | 30 | 0.159
US50 | MoRh | 0.405 | 31 | 0.159
US50 | MoRh | 0.405 | 32 | 0.159
US50 | MoRh | 0.405 | 33 | 0.16
US50 | MoRh | 0.405 | 34 | 0.16
US50 | MoRh | 0.405 | 35 | 0.16
US50 | MoRh | 0.405 | 36 | 0.16
US50 | MoRh | 0.405 | 37 | 0.161
US50 | MoRh | 0.405 | 38 | 0.161
US50 | MoRh | 0.405 | 39 | 0.161
US50 | MoRh | 0.405 | 40 | 0.162
US50 | MoRh | 0.405 | 41 | 0.162
US50 | MoRh | 0.405 | 42 | 0.163
US50 | MoRh | 0.405 | 43 | 0.163
US50 | MoRh | 0.405 | 44 | 0.163
US50 | MoRh | 0.405 | 45 | 0.164
US50 | MoRh | 0.405 | 46 | 0.164
US50 | MoRh | 0.405 | 47 | 0.164
US50 | MoRh | 0.405 | 48 | 0.165
US50 | MoRh | 0.405 | 49 | 0.165
US50 | MoRh | 0.42 | 23 | 0.159
US50 | MoRh | 0.42 | 24 | 0.162
US50 | MoRh | 0.42 | 25 | 0.163
US50 | MoRh | 0.42 | 26 | 0.164
US50 | MoRh | 0.42 | 27 | 0.164
US50 | MoRh | 0.42 | 28 | 0.165
US50 | MoRh | 0.42 | 29 | 0.165
US50 | MoRh | 0.42 | 30 | 0.165
US50 | MoRh | 0.42 | 31 | 0.165
US50 | MoRh | 0.42 | 32 | 0.165
US50 | MoRh | 0.42 | 33 | 0.165
US50 | MoRh | 0.42 | 34 | 0.165
US50 | MoRh | 0.42 | 35 | 0.166
US50 | MoRh | 0.42 | 36 | 0.166
US50 | MoRh | 0.42 | 37 | 0.166
US50 | MoRh | 0.42 | 38 | 0.166
US50 | MoRh | 0.42 | 39 | 0.167
US50 | MoRh | 0.42 | 40 | 0.167
US50 | MoRh | 0.42 | 41 | 0.167
US50 | MoRh | 0.42 | 42 | 0.168
US50 | MoRh | 0.42 | 43 | 0.168
US50 | MoRh | 0.42 | 44 | 0.168
US50 | MoRh | 0.42 | 45 | 0.169
US50 | MoRh | 0.42 | 46 | 0.169
US50 | MoRh | 0.42 | 47 | 0.169
US50 | MoRh | 0.42 | 48 | 0.17
US50 | MoRh | 0.42 | 49 | 0.17
US50 | MoRh | 0.435 | 23 | 0.166
US50 | MoRh | 0.435 | 24 | 0.168
US50 | MoRh | 0.435 | 25 | 0.17
US50 | MoRh | 0.435 | 26 | 0.17
US50 | MoRh | 0.435 | 27 | 0.17
US50 | MoRh | 0.435 | 28 | 0.17
US50 | MoRh | 0.435 | 29 | 0.171
US50 | MoRh | 0.435 | 30 | 0.171
US50 | MoRh | 0.435 | 31 | 0.171
US50 | MoRh | 0.435 | 32 | 0.171
US50 | MoRh | 0.435 | 33 | 0.171
US50 | MoRh | 0.435 | 34 | 0.171
US50 | MoRh | 0.435 | 35 | 0.171
US50 | MoRh | 0.435 | 36 | 0.171
US50 | MoRh | 0.435 | 37 | 0.172
US50 | MoRh | 0.435 | 38 | 0.172
US50 | MoRh | 0.435 | 39 | 0.172
US50 | MoRh | 0.435 | 40 | 0.172
US50 | MoRh | 0.435 | 41 | 0.173
US50 | MoRh | 0.435 | 42 | 0.173
US50 | MoRh | 0.435 | 43 | 0.173
US50 | MoRh | 0.435 | 44 | 0.174
US50 | MoRh | 0.435 | 45 | 0.174
US50 | MoRh | 0.435 | 46 | 0.174
US50 | MoRh | 0.435 | 47 | 0.175
US50 | MoRh | 0.435 | 48 | 0.175
US50 | MoRh | 0.435 | 49 | 0.175
US50 | MoRh | 0.45 | 23 | 0.172
US50 | MoRh | 0.45 | 24 | 0.174
US50 | MoRh | 0.45 | 25 | 0.176
US50 | MoRh | 0.45 | 26 | 0.176
US50 | MoRh | 0.45 | 27 | 0.176
US50 | MoRh | 0.45 | 28 | 0.177
US50 | MoRh | 0.45 | 29 | 0.176
US50 | MoRh | 0.45 | 30 | 0.176
US50 | MoRh | 0.45 | 31 | 0.176
US50 | MoRh | 0.45 | 32 | 0.176
US50 | MoRh | 0.45 | 33 | 0.176
US50 | MoRh | 0.45 | 34 | 0.177
US50 | MoRh | 0.45 | 35 | 0.177
US50 | MoRh | 0.45 | 36 | 0.177
US50 | MoRh | 0.45 | 37 | 0.177
US50 | MoRh | 0.45 | 38 | 0.177
US50 | MoRh | 0.45 | 39 | 0.178
US50 | MoRh | 0.45 | 40 | 0.178
US50 | MoRh | 0.45 | 41 | 0.178
US50 | MoRh | 0.45 | 42 | 0.178
US50 | MoRh | 0.45 | 43 | 0.179
US50 | MoRh | 0.45 | 44 | 0.179
US50 | MoRh | 0.45 | 45 | 0.179
US50 | MoRh | 0.45 | 46 | 0.18
US50 | MoRh | 0.45 | 47 | 0.18
US50 | MoRh | 0.45 | 48 | 0.18
US50 | MoRh | 0.45 | 49 | 0.181
US50 | MoRh | 0.465 | 23 | 0.178
US50 | MoRh | 0.465 | 24 | 0.18
US50 | MoRh | 0.465 | 25 | 0.182
US50 | MoRh | 0.465 | 26 | 0.182
US50 | MoRh | 0.465 | 27 | 0.183
US50 | MoRh | 0.465 | 28 | 0.182
US50 | MoRh | 0.465 | 29 | 0.182
US50 | MoRh | 0.465 | 30 | 0.182
US50 | MoRh | 0.465 | 31 | 0.182
US50 | MoRh | 0.465 | 32 | 0.182
US50 | MoRh | 0.465 | 33 | 0.182
US50 | MoRh | 0.465 | 34 | 0.182
US50 | MoRh | 0.465 | 35 | 0.182
US50 | MoRh | 0.465 | 36 | 0.183
US50 | MoRh | 0.465 | 37 | 0.183
US50 | MoRh | 0.465 | 38 | 0.183
US50 | MoRh | 0.465 | 39 | 0.183
US50 | MoRh | 0.465 | 40 | 0.183
US50 | MoRh | 0.465 | 41 | 0.184
US50 | MoRh | 0.465 | 42 | 0.184
US50 | MoRh | 0.465 | 43 | 0.184
US50 | MoRh | 0.465 | 44 | 0.185
US50 | MoRh | 0.465 | 45 | 0.185
US50 | MoRh | 0.465 | 46 | 0.185
US50 | MoRh | 0.465 | 47 | 0.185
US50 | MoRh | 0.465 | 48 | 0.186
US50 | MoRh | 0.465 | 49 | 0.186
US50 | MoRh | 0.48 | 25 | 0.188
US50 | MoRh | 0.48 | 26 | 0.188
US50 | MoRh | 0.48 | 27 | 0.189
US50 | MoRh | 0.48 | 28 | 0.189
US50 | MoRh | 0.48 | 29 | 0.188
US50 | MoRh | 0.48 | 30 | 0.188
US50 | MoRh | 0.48 | 31 | 0.188
US50 | MoRh | 0.48 | 32 | 0.188
US50 | MoRh | 0.48 | 33 | 0.188
US50 | MoRh | 0.48 | 34 | 0.188
US50 | MoRh | 0.48 | 35 | 0.188
US50 | MoRh | 0.48 | 36 | 0.188
US50 | MoRh | 0.48 | 37 | 0.188
US50 | MoRh | 0.48 | 38 | 0.189
US50 | MoRh | 0.48 | 39 | 0.189
US50 | MoRh | 0.48 | 40 | 0.189
US50 | MoRh | 0.48 | 41 | 0.189
US50 | MoRh | 0.48 | 42 | 0.19
US50 | MoRh | 0.48 | 43 | 0.19
US50 | MoRh | 0.48 | 44 | 0.19
US50 | MoRh | 0.48 | 45 | 0.19
US50 | MoRh | 0.48 | 46 | 0.191
US50 | MoRh | 0.48 | 47 | 0.191
US50 | MoRh | 0.48 | 48 | 0.191
US50 | MoRh | 0.48 | 49 | 0.192
US50 | MoRh | 0.495 | 28 | 0.195
US50 | MoRh | 0.495 | 29 | 0.195
US50 | MoRh | 0.495 | 30 | 0.195
US50 | MoRh | 0.495 | 31 | 0.194
US50 | MoRh | 0.495 | 32 | 0.194
US50 | MoRh | 0.495 | 33 | 0.194
US50 | MoRh | 0.495 | 34 | 0.194
US50 | MoRh | 0.495 | 35 | 0.194
US50 | MoRh | 0.495 | 36 | 0.194
US50 | MoRh | 0.495 | 37 | 0.195
US50 | MoRh | 0.495 | 38 | 0.195
US50 | MoRh | 0.495 | 39 | 0.195
US50 | MoRh | 0.495 | 40 | 0.195
US50 | MoRh | 0.495 | 41 | 0.195
US50 | MoRh | 0.495 | 42 | 0.195
US50 | MoRh | 0.495 | 43 | 0.196
US50 | MoRh | 0.495 | 44 | 0.196
US50 | MoRh | 0.495 | 45 | 0.196
US50 | MoRh | 0.495 | 46 | 0.197
US50 | MoRh | 0.495 | 47 | 0.197
US50 | MoRh | 0.495 | 48 | 0.197
US50 | MoRh | 0.495 | 49 | 0.198
US50 | MoRh | 0.51 | 28 | 0.201
US50 | MoRh | 0.51 | 29 | 0.201
US50 | MoRh | 0.51 | 30 | 0.201
US50 | MoRh | 0.51 | 31 | 0.2
US50 | MoRh | 0.51 | 32 | 0.2
US50 | MoRh | 0.51 | 33 | 0.2
US50 | MoRh | 0.51 | 34 | 0.2
US50 | MoRh | 0.51 | 35 | 0.2
US50 | MoRh | 0.51 | 36 | 0.2
US50 | MoRh | 0.51 | 37 | 0.201
US50 | MoRh | 0.51 | 38 | 0.201
US50 | MoRh | 0.51 | 39 | 0.201
US50 | MoRh | 0.51 | 40 | 0.201
US50 | MoRh | 0.51 | 41 | 0.201
US50 | MoRh | 0.51 | 42 | 0.201
US50 | MoRh | 0.51 | 43 | 0.202
US50 | MoRh | 0.51 | 44 | 0.202
US50 | MoRh | 0.51 | 45 | 0.202
US50 | MoRh | 0.51 | 46 | 0.203
US50 | MoRh | 0.51 | 47 | 0.203
US50 | MoRh | 0.51 | 48 | 0.203
US50 | MoRh | 0.51 | 49 | 0.204
US50 | MoRh | 0.525 | 29 | 0.207
US50 | MoRh | 0.525 | 30 | 0.207
US50 | MoRh | 0.525 | 31 | 0.207
US50 | MoRh | 0.525 | 32 | 0.207
US50 | MoRh | 0.525 | 33 | 0.207
US50 | MoRh | 0.525 | 34 | 0.207
US50 | MoRh | 0.525 | 35 | 0.207
US50 | MoRh | 0.525 | 36 | 0.207
US50 | MoRh | 0.525 | 37 | 0.207
US50 | MoRh | 0.525 | 38 | 0.207
US50 | MoRh | 0.525 | 39 | 0.207
US50 | MoRh | 0.525 | 40 | 0.207
US50 | MoRh | 0.525 | 41 | 0.207
US50 | MoRh | 0.525 | 42 | 0.208
US50 | MoRh | 0.525 | 43 | 0.208
US50 | MoRh | 0.525 | 44 | 0.208
US50 | MoRh | 0.525 | 45 | 0.209
US50 | MoRh | 0.525 | 46 | 0.209
US50 | MoRh | 0.525 | 47 | 0.209
US50 | MoRh | 0.525 | 48 | 0.21
US50 | MoRh | 0.525 | 49 | 0.21
US50 | MoRh | 0.54 | 32 | 0.213
US50 | MoRh | 0.54 | 33 | 0.213
US50 | MoRh | 0.54 | 34 | 0.213
US50 | MoRh | 0.54 | 35 | 0.213
US50 | MoRh | 0.54 | 36 | 0.213
US50 | MoRh | 0.54 | 37 | 0.213
US50 | MoRh | 0.54 | 38 | 0.213
US50 | MoRh | 0.54 | 39 | 0.213
US50 | MoRh | 0.54 | 40 | 0.213
US50 | MoRh | 0.54 | 41 | 0.214
US50 | MoRh | 0.54 | 42 | 0.214
US50 | MoRh | 0.54 | 43 | 0.214
US50 | MoRh | 0.54 | 44 | 0.215
US50 | MoRh | 0.54 | 45 | 0.215
US50 | MoRh | 0.54 | 46 | 0.215
US50 | MoRh | 0.54 | 47 | 0.216
US50 | MoRh | 0.54 | 48 | 0.216
US50 | MoRh | 0.54 | 49 | 0.216
US50 | MoRh | 0.555 | 34 | 0.219
US50 | MoRh | 0.555 | 35 | 0.219
US50 | MoRh | 0.555 | 36 | 0.219
US50 | MoRh | 0.555 | 37 | 0.219
US50 | MoRh | 0.555 | 38 | 0.219
US50 | MoRh | 0.555 | 39 | 0.219
US50 | MoRh | 0.555 | 40 | 0.22
US50 | MoRh | 0.555 | 41 | 0.22
US50 | MoRh | 0.555 | 42 | 0.22
US50 | MoRh | 0.555 | 43 | 0.221
US50 | MoRh | 0.555 | 44 | 0.221
US50 | MoRh | 0.555 | 45 | 0.221
US50 | MoRh | 0.555 | 46 | 0.222
US50 | MoRh | 0.555 | 47 | 0.222
US50 | MoRh | 0.555 | 48 | 0.222
US50 | MoRh | 0.555 | 49 | 0.223
US50 | MoRh | 0.57 | 43 | 0.227
US50 | MoRh | 0.57 | 44 | 0.227
US50 | MoRh | 0.57 | 45 | 0.227
US50 | MoRh | 0.57 | 46 | 0.228
US50 | MoRh | 0.57 | 47 | 0.228
US50 | MoRh | 0.57 | 48 | 0.229
US50 | MoRh | 0.57 | 49 | 0.229
US50 | MoRh | 0.585 | 43 | 0.233
US50 | MoRh | 0.585 | 44 | 0.233
US50 | MoRh | 0.585 | 45 | 0.234
US50 | MoRh | 0.585 | 46 | 0.234
US50 | MoRh | 0.585 | 47 | 0.235
US50 | MoRh | 0.585 | 48 | 0.235
US50 | MoRh | 0.585 | 49 | 0.236
US50 | MoRh | 0.6 | 44 | 0.239
US50 | MoRh | 0.6 | 45 | 0.24
US50 | MoRh | 0.6 | 46 | 0.241
US50 | MoRh | 0.6 | 47 | 0.241
US50 | MoRh | 0.6 | 48 | 0.241
US50 | MoRh | 0.6 | 49 | 0.242
US50 | RhAg | 0.16 | 20 | 0.037
US50 | RhAg | 0.18 | 20 | 0.046
US50 | RhAg | 0.18 | 21 | 0.049
US50 | RhAg | 0.18 | 22 | 0.051
US50 | RhAg | 0.18 | 23 | 0.053
US50 | RhAg | 0.2 | 20 | 0.055
US50 | RhAg | 0.2 | 21 | 0.058
US50 | RhAg | 0.2 | 22 | 0.06
US50 | RhAg | 0.2 | 23 | 0.063
US50 | RhAg | 0.2 | 24 | 0.066
US50 | RhAg | 0.2 | 25 | 0.069
US50 | RhAg | 0.22 | 20 | 0.064
US50 | RhAg | 0.22 | 21 | 0.067
US50 | RhAg | 0.22 | 22 | 0.07
US50 | RhAg | 0.22 | 23 | 0.073
US50 | RhAg | 0.22 | 24 | 0.076
US50 | RhAg | 0.22 | 25 | 0.08
US50 | RhAg | 0.22 | 26 | 0.083
US50 | RhAg | 0.22 | 27 | 0.086
US50 | RhAg | 0.24 | 20 | 0.072
US50 | RhAg | 0.24 | 21 | 0.076
US50 | RhAg | 0.24 | 22 | 0.079
US50 | RhAg | 0.24 | 23 | 0.082
US50 | RhAg | 0.24 | 24 | 0.086
US50 | RhAg | 0.24 | 25 | 0.09
US50 | RhAg | 0.24 | 26 | 0.093
US50 | RhAg | 0.24 | 27 | 0.096
US50 | RhAg | 0.24 | 28 | 0.098
US50 | RhAg | 0.24 | 29 | 0.101
US50 | RhAg | 0.26 | 20 | 0.081
US50 | RhAg | 0.26 | 21 | 0.085
US50 | RhAg | 0.26 | 22 | 0.088
US50 | RhAg | 0.26 | 23 | 0.091
US50 | RhAg | 0.26 | 24 | 0.096
US50 | RhAg | 0.26 | 25 | 0.099
US50 | RhAg | 0.26 | 26 | 0.103
US50 | RhAg | 0.26 | 27 | 0.106
US50 | RhAg | 0.26 | 28 | 0.108
US50 | RhAg | 0.26 | 29 | 0.111
US50 | RhAg | 0.26 | 30 | 0.113
US50 | RhAg | 0.26 | 31 | 0.114
US50 | RhAg | 0.26 | 32 | 0.116
US50 | RhAg | 0.28 | 20 | 0.089
US50 | RhAg | 0.28 | 21 | 0.094
US50 | RhAg | 0.28 | 22 | 0.097
US50 | RhAg | 0.28 | 23 | 0.1
US50 | RhAg | 0.28 | 24 | 0.105
US50 | RhAg | 0.28 | 25 | 0.109
US50 | RhAg | 0.28 | 26 | 0.112
US50 | RhAg | 0.28 | 27 | 0.116
US50 | RhAg | 0.28 | 28 | 0.118
US50 | RhAg | 0.28 | 29 | 0.12
US50 | RhAg | 0.28 | 30 | 0.122
US50 | RhAg | 0.28 | 31 | 0.124
US50 | RhAg | 0.28 | 32 | 0.125
US50 | RhAg | 0.28 | 33 | 0.127
US50 | RhAg | 0.28 | 34 | 0.128
US50 | RhAg | 0.3 | 20 | 0.098
US50 | RhAg | 0.3 | 21 | 0.103
US50 | RhAg | 0.3 | 22 | 0.106
US50 | RhAg | 0.3 | 23 | 0.109
US50 | RhAg | 0.3 | 24 | 0.114
US50 | RhAg | 0.3 | 25 | 0.118
US50 | RhAg | 0.3 | 26 | 0.122
US50 | RhAg | 0.3 | 27 | 0.125
US50 | RhAg | 0.3 | 28 | 0.127
US50 | RhAg | 0.3 | 29 | 0.129
US50 | RhAg | 0.3 | 30 | 0.131
US50 | RhAg | 0.3 | 31 | 0.133
US50 | RhAg | 0.3 | 32 | 0.134
US50 | RhAg | 0.3 | 33 | 0.136
US50 | RhAg | 0.3 | 34 | 0.137
US50 | RhAg | 0.3 | 35 | 0.138
US50 | RhAg | 0.3 | 36 | 0.14
US50 | RhAg | 0.3 | 37 | 0.141
US50 | RhAg | 0.32 | 20 | 0.107
US50 | RhAg | 0.32 | 21 | 0.111
US50 | RhAg | 0.32 | 22 | 0.114
US50 | RhAg | 0.32 | 23 | 0.118
US50 | RhAg | 0.32 | 24 | 0.123
US50 | RhAg | 0.32 | 25 | 0.127
US50 | RhAg | 0.32 | 26 | 0.131
US50 | RhAg | 0.32 | 27 | 0.134
US50 | RhAg | 0.32 | 28 | 0.136
US50 | RhAg | 0.32 | 29 | 0.138
US50 | RhAg | 0.32 | 30 | 0.14
US50 | RhAg | 0.32 | 31 | 0.141
US50 | RhAg | 0.32 | 32 | 0.143
US50 | RhAg | 0.32 | 33 | 0.144
US50 | RhAg | 0.32 | 34 | 0.146
US50 | RhAg | 0.32 | 35 | 0.147
US50 | RhAg | 0.32 | 36 | 0.148
US50 | RhAg | 0.32 | 37 | 0.149
US50 | RhAg | 0.32 | 38 | 0.15
US50 | RhAg | 0.32 | 39 | 0.151
US50 | RhAg | 0.32 | 40 | 0.152
US50 | RhAg | 0.32 | 41 | 0.153
US50 | RhAg | 0.32 | 42 | 0.154
US50 | RhAg | 0.32 | 43 | 0.154
US50 | RhAg | 0.34 | 20 | 0.115
US50 | RhAg | 0.34 | 21 | 0.12
US50 | RhAg | 0.34 | 22 | 0.123
US50 | RhAg | 0.34 | 23 | 0.127
US50 | RhAg | 0.34 | 24 | 0.132
US50 | RhAg | 0.34 | 25 | 0.136
US50 | RhAg | 0.34 | 26 | 0.139
US50 | RhAg | 0.34 | 27 | 0.142
US50 | RhAg | 0.34 | 28 | 0.145
US50 | RhAg | 0.34 | 29 | 0.147
US50 | RhAg | 0.34 | 30 | 0.148
US50 | RhAg | 0.34 | 31 | 0.15
US50 | RhAg | 0.34 | 32 | 0.151
US50 | RhAg | 0.34 | 33 | 0.153
US50 | RhAg | 0.34 | 34 | 0.154
US50 | RhAg | 0.34 | 35 | 0.155
US50 | RhAg | 0.34 | 36 | 0.156
US50 | RhAg | 0.34 | 37 | 0.157
US50 | RhAg | 0.34 | 38 | 0.158
US50 | RhAg | 0.34 | 39 | 0.159
US50 | RhAg | 0.34 | 40 | 0.16
US50 | RhAg | 0.34 | 41 | 0.161
US50 | RhAg | 0.34 | 42 | 0.162
US50 | RhAg | 0.34 | 43 | 0.162
US50 | RhAg | 0.34 | 44 | 0.163
US50 | RhAg | 0.34 | 45 | 0.164
US50 | RhAg | 0.34 | 46 | 0.164
US50 | RhAg | 0.36 | 20 | 0.124
US50 | RhAg | 0.36 | 21 | 0.128
US50 | RhAg | 0.36 | 22 | 0.132
US50 | RhAg | 0.36 | 23 | 0.136
US50 | RhAg | 0.36 | 24 | 0.14
US50 | RhAg | 0.36 | 25 | 0.144
US50 | RhAg | 0.36 | 26 | 0.148
US50 | RhAg | 0.36 | 27 | 0.151
US50 | RhAg | 0.36 | 28 | 0.153
US50 | RhAg | 0.36 | 29 | 0.155
US50 | RhAg | 0.36 | 30 | 0.157
US50 | RhAg | 0.36 | 31 | 0.158
US50 | RhAg | 0.36 | 32 | 0.159
US50 | RhAg | 0.36 | 33 | 0.16
US50 | RhAg | 0.36 | 34 | 0.162
US50 | RhAg | 0.36 | 35 | 0.163
US50 | RhAg | 0.36 | 36 | 0.164
US50 | RhAg | 0.36 | 37 | 0.165
US50 | RhAg | 0.36 | 38 | 0.166
US50 | RhAg | 0.36 | 39 | 0.167
US50 | RhAg | 0.36 | 40 | 0.168
US50 | RhAg | 0.36 | 41 | 0.168
US50 | RhAg | 0.36 | 42 | 0.169
US50 | RhAg | 0.36 | 43 | 0.17
US50 | RhAg | 0.36 | 44 | 0.171
US50 | RhAg | 0.36 | 45 | 0.171
US50 | RhAg | 0.36 | 46 | 0.172
US50 | RhAg | 0.36 | 47 | 0.172
US50 | RhAg | 0.36 | 48 | 0.173
US50 | RhAg | 0.36 | 49 | 0.174
US50 | RhAg | 0.38 | 20 | 0.132
US50 | RhAg | 0.38 | 21 | 0.137
US50 | RhAg | 0.38 | 22 | 0.141
US50 | RhAg | 0.38 | 23 | 0.144
US50 | RhAg | 0.38 | 24 | 0.149
US50 | RhAg | 0.38 | 25 | 0.153
US50 | RhAg | 0.38 | 26 | 0.156
US50 | RhAg | 0.38 | 27 | 0.159
US50 | RhAg | 0.38 | 28 | 0.161
US50 | RhAg | 0.38 | 29 | 0.163
US50 | RhAg | 0.38 | 30 | 0.164
US50 | RhAg | 0.38 | 31 | 0.166
US50 | RhAg | 0.38 | 32 | 0.167
US50 | RhAg | 0.38 | 33 | 0.168
US50 | RhAg | 0.38 | 34 | 0.169
US50 | RhAg | 0.38 | 35 | 0.17
US50 | RhAg | 0.38 | 36 | 0.171
US50 | RhAg | 0.38 | 37 | 0.172
US50 | RhAg | 0.38 | 38 | 0.173
US50 | RhAg | 0.38 | 39 | 0.174
US50 | RhAg | 0.38 | 40 | 0.175
US50 | RhAg | 0.38 | 41 | 0.176
US50 | RhAg | 0.38 | 42 | 0.176
US50 | RhAg | 0.38 | 43 | 0.177
US50 | RhAg | 0.38 | 44 | 0.178
US50 | RhAg | 0.38 | 45 | 0.178
US50 | RhAg | 0.38 | 46 | 0.179
US50 | RhAg | 0.38 | 47 | 0.18
US50 | RhAg | 0.38 | 48 | 0.18
US50 | RhAg | 0.38 | 49 | 0.181
US50 | RhAg | 0.4 | 20 | 0.141
US50 | RhAg | 0.4 | 21 | 0.146
US50 | RhAg | 0.4 | 22 | 0.149
US50 | RhAg | 0.4 | 23 | 0.153
US50 | RhAg | 0.4 | 24 | 0.157
US50 | RhAg | 0.4 | 25 | 0.161
US50 | RhAg | 0.4 | 26 | 0.165
US50 | RhAg | 0.4 | 27 | 0.167
US50 | RhAg | 0.4 | 28 | 0.169
US50 | RhAg | 0.4 | 29 | 0.171
US50 | RhAg | 0.4 | 30 | 0.172
US50 | RhAg | 0.4 | 31 | 0.174
US50 | RhAg | 0.4 | 32 | 0.175
US50 | RhAg | 0.4 | 33 | 0.176
US50 | RhAg | 0.4 | 34 | 0.177
US50 | RhAg | 0.4 | 35 | 0.178
US50 | RhAg | 0.4 | 36 | 0.179
US50 | RhAg | 0.4 | 37 | 0.18
US50 | RhAg | 0.4 | 38 | 0.18
US50 | RhAg | 0.4 | 39 | 0.181
US50 | RhAg | 0.4 | 40 | 0.182
US50 | RhAg | 0.4 | 41 | 0.183
US50 | RhAg | 0.4 | 42 | 0.184
US50 | RhAg | 0.4 | 43 | 0.184
US50 | RhAg | 0.4 | 44 | 0.185
US50 | RhAg | 0.4 | 45 | 0.185
US50 | RhAg | 0.4 | 46 | 0.186
US50 | RhAg | 0.4 | 47 | 0.187
US50 | RhAg | 0.4 | 48 | 0.187
US50 | RhAg | 0.4 | 49 | 0.188
US50 | RhAg | 0.42 | 21 | 0.154
US50 | RhAg | 0.42 | 22 | 0.158
US50 | RhAg | 0.42 | 23 | 0.161
US50 | RhAg | 0.42 | 24 | 0.166
US50 | RhAg | 0.42 | 25 | 0.169
US50 | RhAg | 0.42 | 26 | 0.173
US50 | RhAg | 0.42 | 27 | 0.175
US50 | RhAg | 0.42 | 28 | 0.177
US50 | RhAg | 0.42 | 29 | 0.179
US50 | RhAg | 0.42 | 30 | 0.18
US50 | RhAg | 0.42 | 31 | 0.181
US50 | RhAg | 0.42 | 32 | 0.182
US50 | RhAg | 0.42 | 33 | 0.183
US50 | RhAg | 0.42 | 34 | 0.184
US50 | RhAg | 0.42 | 35 | 0.185
US50 | RhAg | 0.42 | 36 | 0.186
US50 | RhAg | 0.42 | 37 | 0.187
US50 | RhAg | 0.42 | 38 | 0.187
US50 | RhAg | 0.42 | 39 | 0.188
US50 | RhAg | 0.42 | 40 | 0.189
US50 | RhAg | 0.42 | 41 | 0.19
US50 | RhAg | 0.42 | 42 | 0.19
US50 | RhAg | 0.42 | 43 | 0.191
US50 | RhAg | 0.42 | 44 | 0.192
US50 | RhAg | 0.42 | 45 | 0.192
US50 | RhAg | 0.42 | 46 | 0.193
US50 | RhAg | 0.42 | 47 | 0.193
US50 | RhAg | 0.42 | 48 | 0.194
US50 | RhAg | 0.42 | 49 | 0.194
US50 | RhAg | 0.44 | 21 | 0.163
US50 | RhAg | 0.44 | 22 | 0.166
US50 | RhAg | 0.44 | 23 | 0.17
US50 | RhAg | 0.44 | 24 | 0.174
US50 | RhAg | 0.44 | 25 | 0.177
US50 | RhAg | 0.44 | 26 | 0.181
US50 | RhAg | 0.44 | 27 | 0.183
US50 | RhAg | 0.44 | 28 | 0.185
US50 | RhAg | 0.44 | 29 | 0.186
US50 | RhAg | 0.44 | 30 | 0.188
US50 | RhAg | 0.44 | 31 | 0.189
US50 | RhAg | 0.44 | 32 | 0.19
US50 | RhAg | 0.44 | 33 | 0.19
US50 | RhAg | 0.44 | 34 | 0.191
US50 | RhAg | 0.44 | 35 | 0.192
US50 | RhAg | 0.44 | 36 | 0.193
US50 | RhAg | 0.44 | 37 | 0.194
US50 | RhAg | 0.44 | 38 | 0.194
US50 | RhAg | 0.44 | 39 | 0.195
US50 | RhAg | 0.44 | 40 | 0.196
US50 | RhAg | 0.44 | 41 | 0.196
US50 | RhAg | 0.44 | 42 | 0.197
US50 | RhAg | 0.44 | 43 | 0.198
US50 | RhAg | 0.44 | 44 | 0.198
US50 | RhAg | 0.44 | 45 | 0.199
US50 | RhAg | 0.44 | 46 | 0.2
US50 | RhAg | 0.44 | 47 | 0.2
US50 | RhAg | 0.44 | 48 | 0.201
US50 | RhAg | 0.44 | 49 | 0.201
US50 | RhAg | 0.46 | 22 | 0.175
US50 | RhAg | 0.46 | 23 | 0.178
US50 | RhAg | 0.46 | 24 | 0.182
US50 | RhAg | 0.46 | 25 | 0.185
US50 | RhAg | 0.46 | 26 | 0.189
US50 | RhAg | 0.46 | 27 | 0.191
US50 | RhAg | 0.46 | 28 | 0.193
US50 | RhAg | 0.46 | 29 | 0.194
US50 | RhAg | 0.46 | 30 | 0.195
US50 | RhAg | 0.46 | 31 | 0.196
US50 | RhAg | 0.46 | 32 | 0.197
US50 | RhAg | 0.46 | 33 | 0.198
US50 | RhAg | 0.46 | 34 | 0.198
US50 | RhAg | 0.46 | 35 | 0.199
US50 | RhAg | 0.46 | 36 | 0.2
US50 | RhAg | 0.46 | 37 | 0.2
US50 | RhAg | 0.46 | 38 | 0.201
US50 | RhAg | 0.46 | 39 | 0.202
US50 | RhAg | 0.46 | 40 | 0.203
US50 | RhAg | 0.46 | 41 | 0.203
US50 | RhAg | 0.46 | 42 | 0.204
US50 | RhAg | 0.46 | 43 | 0.204
US50 | RhAg | 0.46 | 44 | 0.205
US50 | RhAg | 0.46 | 45 | 0.206
US50 | RhAg | 0.46 | 46 | 0.206
US50 | RhAg | 0.46 | 47 | 0.207
US50 | RhAg | 0.46 | 48 | 0.207
US50 | RhAg | 0.46 | 49 | 0.208
US50 | RhAg | 0.48 | 22 | 0.183
US50 | RhAg | 0.48 | 23 | 0.187
US50 | RhAg | 0.48 | 24 | 0.19
US50 | RhAg | 0.48 | 25 | 0.193
US50 | RhAg | 0.48 | 26 | 0.196
US50 | RhAg | 0.48 | 27 | 0.199
US50 | RhAg | 0.48 | 28 | 0.2
US50 | RhAg | 0.48 | 29 | 0.201
US50 | RhAg | 0.48 | 30 | 0.202
US50 | RhAg | 0.48 | 31 | 0.203
US50 | RhAg | 0.48 | 32 | 0.204
US50 | RhAg | 0.48 | 33 | 0.205
US50 | RhAg | 0.48 | 34 | 0.205
US50 | RhAg | 0.48 | 35 | 0.206
US50 | RhAg | 0.48 | 36 | 0.207
US50 | RhAg | 0.48 | 37 | 0.207
US50 | RhAg | 0.48 | 38 | 0.208
US50 | RhAg | 0.48 | 39 | 0.209
US50 | RhAg | 0.48 | 40 | 0.209
US50 | RhAg | 0.48 | 41 | 0.21
US50 | RhAg | 0.48 | 42 | 0.21
US50 | RhAg | 0.48 | 43 | 0.211
US50 | RhAg | 0.48 | 44 | 0.211
US50 | RhAg | 0.48 | 45 | 0.212
US50 | RhAg | 0.48 | 46 | 0.213
US50 | RhAg | 0.48 | 47 | 0.213
US50 | RhAg | 0.48 | 48 | 0.213
US50 | RhAg | 0.48 | 49 | 0.214
US50 | RhAg | 0.5 | 23 | 0.195
US50 | RhAg | 0.5 | 24 | 0.198
US50 | RhAg | 0.5 | 25 | 0.201
US50 | RhAg | 0.5 | 26 | 0.204
US50 | RhAg | 0.5 | 27 | 0.206
US50 | RhAg | 0.5 | 28 | 0.208
US50 | RhAg | 0.5 | 29 | 0.209
US50 | RhAg | 0.5 | 30 | 0.21
US50 | RhAg | 0.5 | 31 | 0.21
US50 | RhAg | 0.5 | 32 | 0.211
US50 | RhAg | 0.5 | 33 | 0.212
US50 | RhAg | 0.5 | 34 | 0.212
US50 | RhAg | 0.5 | 35 | 0.213
US50 | RhAg | 0.5 | 36 | 0.213
US50 | RhAg | 0.5 | 37 | 0.214
US50 | RhAg | 0.5 | 38 | 0.214
US50 | RhAg | 0.5 | 39 | 0.215
US50 | RhAg | 0.5 | 40 | 0.216
US50 | RhAg | 0.5 | 41 | 0.216
US50 | RhAg | 0.5 | 42 | 0.217
US50 | RhAg | 0.5 | 43 | 0.217
US50 | RhAg | 0.5 | 44 | 0.218
US50 | RhAg | 0.5 | 45 | 0.218
US50 | RhAg | 0.5 | 46 | 0.219
US50 | RhAg | 0.5 | 47 | 0.219
US50 | RhAg | 0.5 | 48 | 0.22
US50 | RhAg | 0.5 | 49 | 0.22
US50 | RhAg | 0.52 | 23 | 0.203
US50 | RhAg | 0.52 | 24 | 0.206
US50 | RhAg | 0.52 | 25 | 0.209
US50 | RhAg | 0.52 | 26 | 0.212
US50 | RhAg | 0.52 | 27 | 0.214
US50 | RhAg | 0.52 | 28 | 0.215
US50 | RhAg | 0.52 | 29 | 0.216
US50 | RhAg | 0.52 | 30 | 0.217
US50 | RhAg | 0.52 | 31 | 0.217
US50 | RhAg | 0.52 | 32 | 0.218
US50 | RhAg | 0.52 | 33 | 0.219
US50 | RhAg | 0.52 | 34 | 0.219
US50 | RhAg | 0.52 | 35 | 0.219
US50 | RhAg | 0.52 | 36 | 0.22
US50 | RhAg | 0.52 | 37 | 0.221
US50 | RhAg | 0.52 | 38 | 0.221
US50 | RhAg | 0.52 | 39 | 0.222
US50 | RhAg | 0.52 | 40 | 0.222
US50 | RhAg | 0.52 | 41 | 0.223
US50 | RhAg | 0.52 | 42 | 0.223
US50 | RhAg | 0.52 | 43 | 0.224
US50 | RhAg | 0.52 | 44 | 0.224
US50 | RhAg | 0.52 | 45 | 0.225
US50 | RhAg | 0.52 | 46 | 0.225
US50 | RhAg | 0.52 | 47 | 0.226
US50 | RhAg | 0.52 | 48 | 0.226
US50 | RhAg | 0.52 | 49 | 0.227
US50 | RhAg | 0.54 | 24 | 0.214
US50 | RhAg | 0.54 | 25 | 0.217
US50 | RhAg | 0.54 | 26 | 0.219
US50 | RhAg | 0.54 | 27 | 0.221
US50 | RhAg | 0.54 | 28 | 0.222
US50 | RhAg | 0.54 | 29 | 0.223
US50 | RhAg | 0.54 | 30 | 0.224
US50 | RhAg | 0.54 | 31 | 0.225
US50 | RhAg | 0.54 | 32 | 0.225
US50 | RhAg | 0.54 | 33 | 0.225
US50 | RhAg | 0.54 | 34 | 0.226
US50 | RhAg | 0.54 | 35 | 0.226
US50 | RhAg | 0.54 | 36 | 0.227
US50 | RhAg | 0.54 | 37 | 0.227
US50 | RhAg | 0.54 | 38 | 0.228
US50 | RhAg | 0.54 | 39 | 0.228
US50 | RhAg | 0.54 | 40 | 0.229
US50 | RhAg | 0.54 | 41 | 0.229
US50 | RhAg | 0.54 | 42 | 0.23
US50 | RhAg | 0.54 | 43 | 0.23
US50 | RhAg | 0.54 | 44 | 0.231
US50 | RhAg | 0.54 | 45 | 0.231
US50 | RhAg | 0.54 | 46 | 0.231
US50 | RhAg | 0.54 | 47 | 0.232
US50 | RhAg | 0.54 | 48 | 0.232
US50 | RhAg | 0.54 | 49 | 0.233
US50 | RhAg | 0.56 | 26 | 0.227
US50 | RhAg | 0.56 | 27 | 0.229
US50 | RhAg | 0.56 | 28 | 0.23
US50 | RhAg | 0.56 | 29 | 0.23
US50 | RhAg | 0.56 | 30 | 0.231
US50 | RhAg | 0.56 | 31 | 0.231
US50 | RhAg | 0.56 | 32 | 0.232
US50 | RhAg | 0.56 | 33 | 0.232
US50 | RhAg | 0.56 | 34 | 0.232
US50 | RhAg | 0.56 | 35 | 0.233
US50 | RhAg | 0.56 | 36 | 0.233
US50 | RhAg | 0.56 | 37 | 0.234
US50 | RhAg | 0.56 | 38 | 0.234
US50 | RhAg | 0.56 | 39 | 0.234
US50 | RhAg | 0.56 | 40 | 0.235
US50 | RhAg | 0.56 | 41 | 0.235
US50 | RhAg | 0.56 | 42 | 0.236
US50 | RhAg | 0.56 | 43 | 0.236
US50 | RhAg | 0.56 | 44 | 0.237
US50 | RhAg | 0.56 | 45 | 0.237
US50 | RhAg | 0.56 | 46 | 0.238
US50 | RhAg | 0.56 | 47 | 0.238
US50 | RhAg | 0.56 | 48 | 0.238
US50 | RhAg | 0.56 | 49 | 0.239
US50 | RhAg | 0.58 | 28 | 0.237
US50 | RhAg | 0.58 | 29 | 0.237
US50 | RhAg | 0.58 | 30 | 0.238
US50 | RhAg | 0.58 | 31 | 0.238
US50 | RhAg | 0.58 | 32 | 0.239
US50 | RhAg | 0.58 | 33 | 0.239
US50 | RhAg | 0.58 | 34 | 0.239
US50 | RhAg | 0.58 | 35 | 0.239
US50 | RhAg | 0.58 | 36 | 0.24
US50 | RhAg | 0.58 | 37 | 0.24
US50 | RhAg | 0.58 | 38 | 0.24
US50 | RhAg | 0.58 | 39 | 0.241
US50 | RhAg | 0.58 | 40 | 0.241
US50 | RhAg | 0.58 | 41 | 0.242
US50 | RhAg | 0.58 | 42 | 0.242
US50 | RhAg | 0.58 | 43 | 0.243
US50 | RhAg | 0.58 | 44 | 0.243
US50 | RhAg | 0.58 | 45 | 0.243
US50 | RhAg | 0.58 | 46 | 0.244
US50 | RhAg | 0.58 | 47 | 0.244
US50 | RhAg | 0.58 | 48 | 0.245
US50 | RhAg | 0.58 | 49 | 0.245
US50 | RhAg | 0.6 | 30 | 0.245
US50 | RhAg | 0.6 | 31 | 0.245
US50 | RhAg | 0.6 | 32 | 0.245
US50 | RhAg | 0.6 | 33 | 0.246
US50 | RhAg | 0.6 | 34 | 0.246
US50 | RhAg | 0.6 | 35 | 0.246
US50 | RhAg | 0.6 | 36 | 0.246
US50 | RhAg | 0.6 | 37 | 0.246
US50 | RhAg | 0.6 | 38 | 0.247
US50 | RhAg | 0.6 | 39 | 0.247
US50 | RhAg | 0.6 | 40 | 0.248
US50 | RhAg | 0.6 | 41 | 0.248
US50 | RhAg | 0.6 | 42 | 0.248
US50 | RhAg | 0.6 | 43 | 0.249
US50 | RhAg | 0.6 | 44 | 0.249
US50 | RhAg | 0.6 | 45 | 0.249
US50 | RhAg | 0.6 | 46 | 0.25
US50 | RhAg | 0.6 | 47 | 0.25
US50 | RhAg | 0.6 | 48 | 0.251
US50 | RhAg | 0.6 | 49 | 0.251
US50 | RhAg | 0.62 | 33 | 0.252
US50 | RhAg | 0.62 | 34 | 0.252
US50 | RhAg | 0.62 | 35 | 0.252
US50 | RhAg | 0.62 | 36 | 0.253
US50 | RhAg | 0.62 | 37 | 0.253
US50 | RhAg | 0.62 | 38 | 0.253
US50 | RhAg | 0.62 | 39 | 0.254
US50 | RhAg | 0.62 | 40 | 0.254
US50 | RhAg | 0.62 | 41 | 0.254
US50 | RhAg | 0.62 | 42 | 0.255
US50 | RhAg | 0.62 | 43 | 0.255
US50 | RhAg | 0.62 | 44 | 0.255
US50 | RhAg | 0.62 | 45 | 0.256
US50 | RhAg | 0.62 | 46 | 0.256
US50 | RhAg | 0.62 | 47 | 0.256
US50 | RhAg | 0.62 | 48 | 0.257
US50 | RhAg | 0.62 | 49 | 0.257
US50 | RhAg | 0.64 | 35 | 0.259
US50 | RhAg | 0.64 | 36 | 0.259
US50 | RhAg | 0.64 | 37 | 0.259
US50 | RhAg | 0.64 | 38 | 0.26
US50 | RhAg | 0.64 | 39 | 0.26
US50 | RhAg | 0.64 | 40 | 0.26
US50 | RhAg | 0.64 | 41 | 0.26
US50 | RhAg | 0.64 | 42 | 0.261
US50 | RhAg | 0.64 | 43 | 0.261
US50 | RhAg | 0.64 | 44 | 0.261
US50 | RhAg | 0.64 | 45 | 0.262
US50 | RhAg | 0.64 | 46 | 0.262
US50 | RhAg | 0.64 | 47 | 0.262
US50 | RhAg | 0.64 | 48 | 0.263
US50 | RhAg | 0.64 | 49 | 0.263
US50 | RhAg | 0.66 | 38 | 0.266
US50 | RhAg | 0.66 | 39 | 0.266
US50 | RhAg | 0.66 | 40 | 0.266
US50 | RhAg | 0.66 | 41 | 0.267
US50 | RhAg | 0.66 | 42 | 0.267
US50 | RhAg | 0.66 | 43 | 0.267
US50 | RhAg | 0.66 | 44 | 0.267
US50 | RhAg | 0.66 | 45 | 0.268
US50 | RhAg | 0.66 | 46 | 0.268
US50 | RhAg | 0.66 | 47 | 0.268
US50 | RhAg | 0.66 | 48 | 0.269
US50 | RhAg | 0.66 | 49 | 0.269
US50 | RhAg | 0.68 | 44 | 0.274
US50 | RhAg | 0.68 | 45 | 0.274
US50 | RhAg | 0.68 | 46 | 0.274
US50 | RhAg | 0.68 | 47 | 0.275
US50 | RhAg | 0.68 | 48 | 0.275
US50 | RhAg | 0.68 | 49 | 0.275
US50 | RhAg | 0.7 | 47 | 0.281
US50 | RhAg | 0.7 | 48 | 0.281
US50 | RhAg | 0.7 | 49 | 0.281
US50 | RhRh | 0.16 | 20 | 0.037
US50 | RhRh | 0.18 | 20 | 0.046
US50 | RhRh | 0.18 | 21 | 0.049
US50 | RhRh | 0.18 | 22 | 0.051
US50 | RhRh | 0.18 | 23 | 0.053
US50 | RhRh | 0.2 | 20 | 0.055
US50 | RhRh | 0.2 | 21 | 0.058
US50 | RhRh | 0.2 | 22 | 0.06
US50 | RhRh | 0.2 | 23 | 0.063
US50 | RhRh | 0.2 | 24 | 0.066
US50 | RhRh | 0.2 | 25 | 0.069
US50 | RhRh | 0.22 | 20 | 0.064
US50 | RhRh | 0.22 | 21 | 0.067
US50 | RhRh | 0.22 | 22 | 0.07
US50 | RhRh | 0.22 | 23 | 0.073
US50 | RhRh | 0.22 | 24 | 0.076
US50 | RhRh | 0.22 | 25 | 0.08
US50 | RhRh | 0.22 | 26 | 0.082
US50 | RhRh | 0.22 | 27 | 0.085
US50 | RhRh | 0.24 | 20 | 0.072
US50 | RhRh | 0.24 | 21 | 0.076
US50 | RhRh | 0.24 | 22 | 0.079
US50 | RhRh | 0.24 | 23 | 0.082
US50 | RhRh | 0.24 | 24 | 0.086
US50 | RhRh | 0.24 | 25 | 0.09
US50 | RhRh | 0.24 | 26 | 0.093
US50 | RhRh | 0.24 | 27 | 0.095
US50 | RhRh | 0.24 | 28 | 0.097
US50 | RhRh | 0.24 | 29 | 0.1
US50 | RhRh | 0.26 | 20 | 0.081
US50 | RhRh | 0.26 | 21 | 0.085
US50 | RhRh | 0.26 | 22 | 0.088
US50 | RhRh | 0.26 | 23 | 0.091
US50 | RhRh | 0.26 | 24 | 0.096
US50 | RhRh | 0.26 | 25 | 0.099
US50 | RhRh | 0.26 | 26 | 0.102
US50 | RhRh | 0.26 | 27 | 0.105
US50 | RhRh | 0.26 | 28 | 0.107
US50 | RhRh | 0.26 | 29 | 0.109
US50 | RhRh | 0.26 | 30 | 0.111
US50 | RhRh | 0.26 | 31 | 0.113
US50 | RhRh | 0.26 | 32 | 0.115
US50 | RhRh | 0.28 | 20 | 0.089
US50 | RhRh | 0.28 | 21 | 0.094
US50 | RhRh | 0.28 | 22 | 0.097
US50 | RhRh | 0.28 | 23 | 0.1
US50 | RhRh | 0.28 | 24 | 0.105
US50 | RhRh | 0.28 | 25 | 0.109
US50 | RhRh | 0.28 | 26 | 0.112
US50 | RhRh | 0.28 | 27 | 0.114
US50 | RhRh | 0.28 | 28 | 0.116
US50 | RhRh | 0.28 | 29 | 0.118
US50 | RhRh | 0.28 | 30 | 0.12
US50 | RhRh | 0.28 | 31 | 0.122
US50 | RhRh | 0.28 | 32 | 0.124
US50 | RhRh | 0.28 | 33 | 0.125
US50 | RhRh | 0.28 | 34 | 0.127
US50 | RhRh | 0.3 | 20 | 0.098
US50 | RhRh | 0.3 | 21 | 0.103
US50 | RhRh | 0.3 | 22 | 0.106
US50 | RhRh | 0.3 | 23 | 0.109
US50 | RhRh | 0.3 | 24 | 0.114
US50 | RhRh | 0.3 | 25 | 0.118
US50 | RhRh | 0.3 | 26 | 0.121
US50 | RhRh | 0.3 | 27 | 0.123
US50 | RhRh | 0.3 | 28 | 0.125
US50 | RhRh | 0.3 | 29 | 0.127
US50 | RhRh | 0.3 | 30 | 0.129
US50 | RhRh | 0.3 | 31 | 0.131
US50 | RhRh | 0.3 | 32 | 0.132
US50 | RhRh | 0.3 | 33 | 0.134
US50 | RhRh | 0.3 | 34 | 0.135
US50 | RhRh | 0.3 | 35 | 0.137
US50 | RhRh | 0.3 | 36 | 0.138
US50 | RhRh | 0.3 | 37 | 0.139
US50 | RhRh | 0.32 | 20 | 0.107
US50 | RhRh | 0.32 | 21 | 0.111
US50 | RhRh | 0.32 | 22 | 0.114
US50 | RhRh | 0.32 | 23 | 0.118
US50 | RhRh | 0.32 | 24 | 0.123
US50 | RhRh | 0.32 | 25 | 0.127
US50 | RhRh | 0.32 | 26 | 0.13
US50 | RhRh | 0.32 | 27 | 0.132
US50 | RhRh | 0.32 | 28 | 0.134
US50 | RhRh | 0.32 | 29 | 0.136
US50 | RhRh | 0.32 | 30 | 0.138
US50 | RhRh | 0.32 | 31 | 0.139
US50 | RhRh | 0.32 | 32 | 0.141
US50 | RhRh | 0.32 | 33 | 0.142
US50 | RhRh | 0.32 | 34 | 0.144
US50 | RhRh | 0.32 | 35 | 0.145
US50 | RhRh | 0.32 | 36 | 0.146
US50 | RhRh | 0.32 | 37 | 0.147
US50 | RhRh | 0.32 | 38 | 0.148
US50 | RhRh | 0.32 | 39 | 0.15
US50 | RhRh | 0.34 | 20 | 0.115
US50 | RhRh | 0.34 | 21 | 0.12
US50 | RhRh | 0.34 | 22 | 0.123
US50 | RhRh | 0.34 | 23 | 0.127
US50 | RhRh | 0.34 | 24 | 0.132
US50 | RhRh | 0.34 | 25 | 0.135
US50 | RhRh | 0.34 | 26 | 0.138
US50 | RhRh | 0.34 | 27 | 0.141
US50 | RhRh | 0.34 | 28 | 0.142
US50 | RhRh | 0.34 | 29 | 0.144
US50 | RhRh | 0.34 | 30 | 0.146
US50 | RhRh | 0.34 | 31 | 0.147
US50 | RhRh | 0.34 | 32 | 0.149
US50 | RhRh | 0.34 | 33 | 0.15
US50 | RhRh | 0.34 | 34 | 0.151
US50 | RhRh | 0.34 | 35 | 0.153
US50 | RhRh | 0.34 | 36 | 0.154
US50 | RhRh | 0.34 | 37 | 0.155
US50 | RhRh | 0.34 | 38 | 0.156
US50 | RhRh | 0.34 | 39 | 0.157
US50 | RhRh | 0.34 | 40 | 0.158
US50 | RhRh | 0.34 | 41 | 0.159
US50 | RhRh | 0.34 | 42 | 0.16
US50 | RhRh | 0.36 | 20 | 0.124
US50 | RhRh | 0.36 | 21 | 0.128
US50 | RhRh | 0.36 | 22 | 0.132
US50 | RhRh | 0.36 | 23 | 0.136
US50 | RhRh | 0.36 | 24 | 0.14
US50 | RhRh | 0.36 | 25 | 0.144
US50 | RhRh | 0.36 | 26 | 0.147
US50 | RhRh | 0.36 | 27 | 0.149
US50 | RhRh | 0.36 | 28 | 0.151
US50 | RhRh | 0.36 | 29 | 0.152
US50 | RhRh | 0.36 | 30 | 0.154
US50 | RhRh | 0.36 | 31 | 0.155
US50 | RhRh | 0.36 | 32 | 0.157
US50 | RhRh | 0.36 | 33 | 0.158
US50 | RhRh | 0.36 | 34 | 0.159
US50 | RhRh | 0.36 | 35 | 0.16
US50 | RhRh | 0.36 | 36 | 0.161
US50 | RhRh | 0.36 | 37 | 0.163
US50 | RhRh | 0.36 | 38 | 0.164
US50 | RhRh | 0.36 | 39 | 0.165
US50 | RhRh | 0.36 | 40 | 0.166
US50 | RhRh | 0.36 | 41 | 0.167
US50 | RhRh | 0.36 | 42 | 0.167
US50 | RhRh | 0.36 | 43 | 0.168
US50 | RhRh | 0.36 | 44 | 0.169
US50 | RhRh | 0.36 | 45 | 0.17
US50 | RhRh | 0.36 | 46 | 0.171
US50 | RhRh | 0.38 | 20 | 0.132
US50 | RhRh | 0.38 | 21 | 0.137
US50 | RhRh | 0.38 | 22 | 0.14
US50 | RhRh | 0.38 | 23 | 0.144
US50 | RhRh | 0.38 | 24 | 0.149
US50 | RhRh | 0.38 | 25 | 0.152
US50 | RhRh | 0.38 | 26 | 0.155
US50 | RhRh | 0.38 | 27 | 0.157
US50 | RhRh | 0.38 | 28 | 0.159
US50 | RhRh | 0.38 | 29 | 0.16
US50 | RhRh | 0.38 | 30 | 0.162
US50 | RhRh | 0.38 | 31 | 0.163
US50 | RhRh | 0.38 | 32 | 0.164
US50 | RhRh | 0.38 | 33 | 0.165
US50 | RhRh | 0.38 | 34 | 0.167
US50 | RhRh | 0.38 | 35 | 0.168
US50 | RhRh | 0.38 | 36 | 0.169
US50 | RhRh | 0.38 | 37 | 0.17
US50 | RhRh | 0.38 | 38 | 0.171
US50 | RhRh | 0.38 | 39 | 0.172
US50 | RhRh | 0.38 | 40 | 0.173
US50 | RhRh | 0.38 | 41 | 0.174
US50 | RhRh | 0.38 | 42 | 0.175
US50 | RhRh | 0.38 | 43 | 0.175
US50 | RhRh | 0.38 | 44 | 0.176
US50 | RhRh | 0.38 | 45 | 0.177
US50 | RhRh | 0.38 | 46 | 0.178
US50 | RhRh | 0.38 | 47 | 0.178
US50 | RhRh | 0.38 | 48 | 0.179
US50 | RhRh | 0.38 | 49 | 0.179
US50 | RhRh | 0.4 | 20 | 0.141
US50 | RhRh | 0.4 | 21 | 0.146
US50 | RhRh | 0.4 | 22 | 0.149
US50 | RhRh | 0.4 | 23 | 0.153
US50 | RhRh | 0.4 | 24 | 0.157
US50 | RhRh | 0.4 | 25 | 0.161
US50 | RhRh | 0.4 | 26 | 0.163
US50 | RhRh | 0.4 | 27 | 0.165
US50 | RhRh | 0.4 | 28 | 0.167
US50 | RhRh | 0.4 | 29 | 0.168
US50 | RhRh | 0.4 | 30 | 0.169
US50 | RhRh | 0.4 | 31 | 0.171
US50 | RhRh | 0.4 | 32 | 0.172
US50 | RhRh | 0.4 | 33 | 0.173
US50 | RhRh | 0.4 | 34 | 0.174
US50 | RhRh | 0.4 | 35 | 0.175
US50 | RhRh | 0.4 | 36 | 0.176
US50 | RhRh | 0.4 | 37 | 0.177
US50 | RhRh | 0.4 | 38 | 0.178
US50 | RhRh | 0.4 | 39 | 0.179
US50 | RhRh | 0.4 | 40 | 0.18
US50 | RhRh | 0.4 | 41 | 0.181
US50 | RhRh | 0.4 | 42 | 0.182
US50 | RhRh | 0.4 | 43 | 0.182
US50 | RhRh | 0.4 | 44 | 0.183
US50 | RhRh | 0.4 | 45 | 0.184
US50 | RhRh | 0.4 | 46 | 0.184
US50 | RhRh | 0.4 | 47 | 0.185
US50 | RhRh | 0.4 | 48 | 0.186
US50 | RhRh | 0.4 | 49 | 0.186
US50 | RhRh | 0.42 | 21 | 0.154
US50 | RhRh | 0.42 | 22 | 0.158
US50 | RhRh | 0.42 | 23 | 0.161
US50 | RhRh | 0.42 | 24 | 0.166
US50 | RhRh | 0.42 | 25 | 0.169
US50 | RhRh | 0.42 | 26 | 0.171
US50 | RhRh | 0.42 | 27 | 0.173
US50 | RhRh | 0.42 | 28 | 0.175
US50 | RhRh | 0.42 | 29 | 0.176
US50 | RhRh | 0.42 | 30 | 0.177
US50 | RhRh | 0.42 | 31 | 0.178
US50 | RhRh | 0.42 | 32 | 0.179
US50 | RhRh | 0.42 | 33 | 0.18
US50 | RhRh | 0.42 | 34 | 0.181
US50 | RhRh | 0.42 | 35 | 0.182
US50 | RhRh | 0.42 | 36 | 0.183
US50 | RhRh | 0.42 | 37 | 0.184
US50 | RhRh | 0.42 | 38 | 0.185
US50 | RhRh | 0.42 | 39 | 0.186
US50 | RhRh | 0.42 | 40 | 0.187
US50 | RhRh | 0.42 | 41 | 0.187
US50 | RhRh | 0.42 | 42 | 0.188
US50 | RhRh | 0.42 | 43 | 0.189
US50 | RhRh | 0.42 | 44 | 0.19
US50 | RhRh | 0.42 | 45 | 0.19
US50 | RhRh | 0.42 | 46 | 0.191
US50 | RhRh | 0.42 | 47 | 0.192
US50 | RhRh | 0.42 | 48 | 0.192
US50 | RhRh | 0.42 | 49 | 0.193
US50 | RhRh | 0.44 | 21 | 0.163
US50 | RhRh | 0.44 | 22 | 0.166
US50 | RhRh | 0.44 | 23 | 0.17
US50 | RhRh | 0.44 | 24 | 0.174
US50 | RhRh | 0.44 | 25 | 0.177
US50 | RhRh | 0.44 | 26 | 0.179
US50 | RhRh | 0.44 | 27 | 0.181
US50 | RhRh | 0.44 | 28 | 0.182
US50 | RhRh | 0.44 | 29 | 0.183
US50 | RhRh | 0.44 | 30 | 0.184
US50 | RhRh | 0.44 | 31 | 0.185
US50 | RhRh | 0.44 | 32 | 0.186
US50 | RhRh | 0.44 | 33 | 0.187
US50 | RhRh | 0.44 | 34 | 0.188
US50 | RhRh | 0.44 | 35 | 0.189
US50 | RhRh | 0.44 | 36 | 0.19
US50 | RhRh | 0.44 | 37 | 0.191
US50 | RhRh | 0.44 | 38 | 0.192
US50 | RhRh | 0.44 | 39 | 0.193
US50 | RhRh | 0.44 | 40 | 0.193
US50 | RhRh | 0.44 | 41 | 0.194
US50 | RhRh | 0.44 | 42 | 0.195
US50 | RhRh | 0.44 | 43 | 0.196
US50 | RhRh | 0.44 | 44 | 0.196
US50 | RhRh | 0.44 | 45 | 0.197
US50 | RhRh | 0.44 | 46 | 0.198
US50 | RhRh | 0.44 | 47 | 0.198
US50 | RhRh | 0.44 | 48 | 0.199
US50 | RhRh | 0.44 | 49 | 0.199
US50 | RhRh | 0.46 | 22 | 0.175
US50 | RhRh | 0.46 | 23 | 0.178
US50 | RhRh | 0.46 | 24 | 0.182
US50 | RhRh | 0.46 | 25 | 0.185
US50 | RhRh | 0.46 | 26 | 0.187
US50 | RhRh | 0.46 | 27 | 0.188
US50 | RhRh | 0.46 | 28 | 0.19
US50 | RhRh | 0.46 | 29 | 0.191
US50 | RhRh | 0.46 | 30 | 0.192
US50 | RhRh | 0.46 | 31 | 0.193
US50 | RhRh | 0.46 | 32 | 0.193
US50 | RhRh | 0.46 | 33 | 0.194
US50 | RhRh | 0.46 | 34 | 0.195
US50 | RhRh | 0.46 | 35 | 0.196
US50 | RhRh | 0.46 | 36 | 0.197
US50 | RhRh | 0.46 | 37 | 0.197
US50 | RhRh | 0.46 | 38 | 0.198
US50 | RhRh | 0.46 | 39 | 0.199
US50 | RhRh | 0.46 | 40 | 0.2
US50 | RhRh | 0.46 | 41 | 0.201
US50 | RhRh | 0.46 | 42 | 0.201
US50 | RhRh | 0.46 | 43 | 0.202
US50 | RhRh | 0.46 | 44 | 0.203
US50 | RhRh | 0.46 | 45 | 0.203
US50 | RhRh | 0.46 | 46 | 0.204
US50 | RhRh | 0.46 | 47 | 0.205
US50 | RhRh | 0.46 | 48 | 0.205
US50 | RhRh | 0.46 | 49 | 0.206
US50 | RhRh | 0.48 | 22 | 0.183
US50 | RhRh | 0.48 | 23 | 0.186
US50 | RhRh | 0.48 | 24 | 0.19
US50 | RhRh | 0.48 | 25 | 0.193
US50 | RhRh | 0.48 | 26 | 0.195
US50 | RhRh | 0.48 | 27 | 0.196
US50 | RhRh | 0.48 | 28 | 0.197
US50 | RhRh | 0.48 | 29 | 0.198
US50 | RhRh | 0.48 | 30 | 0.199
US50 | RhRh | 0.48 | 31 | 0.2
US50 | RhRh | 0.48 | 32 | 0.2
US50 | RhRh | 0.48 | 33 | 0.201
US50 | RhRh | 0.48 | 34 | 0.202
US50 | RhRh | 0.48 | 35 | 0.203
US50 | RhRh | 0.48 | 36 | 0.203
US50 | RhRh | 0.48 | 37 | 0.204
US50 | RhRh | 0.48 | 38 | 0.205
US50 | RhRh | 0.48 | 39 | 0.206
US50 | RhRh | 0.48 | 40 | 0.206
US50 | RhRh | 0.48 | 41 | 0.207
US50 | RhRh | 0.48 | 42 | 0.208
US50 | RhRh | 0.48 | 43 | 0.209
US50 | RhRh | 0.48 | 44 | 0.209
US50 | RhRh | 0.48 | 45 | 0.21
US50 | RhRh | 0.48 | 46 | 0.21
US50 | RhRh | 0.48 | 47 | 0.211
US50 | RhRh | 0.48 | 48 | 0.212
US50 | RhRh | 0.48 | 49 | 0.212
US50 | RhRh | 0.5 | 23 | 0.195
US50 | RhRh | 0.5 | 24 | 0.198
US50 | RhRh | 0.5 | 25 | 0.201
US50 | RhRh | 0.5 | 26 | 0.203
US50 | RhRh | 0.5 | 27 | 0.204
US50 | RhRh | 0.5 | 28 | 0.205
US50 | RhRh | 0.5 | 29 | 0.205
US50 | RhRh | 0.5 | 30 | 0.206
US50 | RhRh | 0.5 | 31 | 0.207
US50 | RhRh | 0.5 | 32 | 0.207
US50 | RhRh | 0.5 | 33 | 0.208
US50 | RhRh | 0.5 | 34 | 0.209
US50 | RhRh | 0.5 | 35 | 0.209
US50 | RhRh | 0.5 | 36 | 0.21
US50 | RhRh | 0.5 | 37 | 0.211
US50 | RhRh | 0.5 | 38 | 0.212
US50 | RhRh | 0.5 | 39 | 0.212
US50 | RhRh | 0.5 | 40 | 0.213
US50 | RhRh | 0.5 | 41 | 0.214
US50 | RhRh | 0.5 | 42 | 0.214
US50 | RhRh | 0.5 | 43 | 0.215
US50 | RhRh | 0.5 | 44 | 0.215
US50 | RhRh | 0.5 | 45 | 0.216
US50 | RhRh | 0.5 | 46 | 0.217
US50 | RhRh | 0.5 | 47 | 0.217
US50 | RhRh | 0.5 | 48 | 0.218
US50 | RhRh | 0.5 | 49 | 0.218
US50 | RhRh | 0.52 | 23 | 0.203
US50 | RhRh | 0.52 | 24 | 0.206
US50 | RhRh | 0.52 | 25 | 0.209
US50 | RhRh | 0.52 | 26 | 0.21
US50 | RhRh | 0.52 | 27 | 0.211
US50 | RhRh | 0.52 | 28 | 0.212
US50 | RhRh | 0.52 | 29 | 0.213
US50 | RhRh | 0.52 | 30 | 0.213
US50 | RhRh | 0.52 | 31 | 0.214
US50 | RhRh | 0.52 | 32 | 0.214
US50 | RhRh | 0.52 | 33 | 0.215
US50 | RhRh | 0.52 | 34 | 0.216
US50 | RhRh | 0.52 | 35 | 0.216
US50 | RhRh | 0.52 | 36 | 0.217
US50 | RhRh | 0.52 | 37 | 0.217
US50 | RhRh | 0.52 | 38 | 0.218
US50 | RhRh | 0.52 | 39 | 0.219
US50 | RhRh | 0.52 | 40 | 0.219
US50 | RhRh | 0.52 | 41 | 0.22
US50 | RhRh | 0.52 | 42 | 0.221
US50 | RhRh | 0.52 | 43 | 0.221
US50 | RhRh | 0.52 | 44 | 0.222
US50 | RhRh | 0.52 | 45 | 0.222
US50 | RhRh | 0.52 | 46 | 0.223
US50 | RhRh | 0.52 | 47 | 0.223
US50 | RhRh | 0.52 | 48 | 0.224
US50 | RhRh | 0.52 | 49 | 0.224
US50 | RhRh | 0.54 | 24 | 0.214
US50 | RhRh | 0.54 | 25 | 0.216
US50 | RhRh | 0.54 | 26 | 0.218
US50 | RhRh | 0.54 | 27 | 0.219
US50 | RhRh | 0.54 | 28 | 0.219
US50 | RhRh | 0.54 | 29 | 0.22
US50 | RhRh | 0.54 | 30 | 0.22
US50 | RhRh | 0.54 | 31 | 0.221
US50 | RhRh | 0.54 | 32 | 0.221
US50 | RhRh | 0.54 | 33 | 0.222
US50 | RhRh | 0.54 | 34 | 0.222
US50 | RhRh | 0.54 | 35 | 0.223
US50 | RhRh | 0.54 | 36 | 0.223
US50 | RhRh | 0.54 | 37 | 0.224
US50 | RhRh | 0.54 | 38 | 0.224
US50 | RhRh | 0.54 | 39 | 0.225
US50 | RhRh | 0.54 | 40 | 0.226
US50 | RhRh | 0.54 | 41 | 0.226
US50 | RhRh | 0.54 | 42 | 0.227
US50 | RhRh | 0.54 | 43 | 0.227
US50 | RhRh | 0.54 | 44 | 0.228
US50 | RhRh | 0.54 | 45 | 0.228
US50 | RhRh | 0.54 | 46 | 0.229
US50 | RhRh | 0.54 | 47 | 0.229
US50 | RhRh | 0.54 | 48 | 0.23
US50 | RhRh | 0.54 | 49 | 0.23
US50 | RhRh | 0.56 | 26 | 0.225
US50 | RhRh | 0.56 | 27 | 0.226
US50 | RhRh | 0.56 | 28 | 0.226
US50 | RhRh | 0.56 | 29 | 0.227
US50 | RhRh | 0.56 | 30 | 0.227
US50 | RhRh | 0.56 | 31 | 0.228
US50 | RhRh | 0.56 | 32 | 0.228
US50 | RhRh | 0.56 | 33 | 0.228
US50 | RhRh | 0.56 | 34 | 0.229
US50 | RhRh | 0.56 | 35 | 0.229
US50 | RhRh | 0.56 | 36 | 0.23
US50 | RhRh | 0.56 | 37 | 0.23
US50 | RhRh | 0.56 | 38 | 0.231
US50 | RhRh | 0.56 | 39 | 0.231
US50 | RhRh | 0.56 | 40 | 0.232
US50 | RhRh | 0.56 | 41 | 0.232
US50 | RhRh | 0.56 | 42 | 0.233
US50 | RhRh | 0.56 | 43 | 0.234
US50 | RhRh | 0.56 | 44 | 0.234
US50 | RhRh | 0.56 | 45 | 0.235
US50 | RhRh | 0.56 | 46 | 0.235
US50 | RhRh | 0.56 | 47 | 0.236
US50 | RhRh | 0.56 | 48 | 0.236
US50 | RhRh | 0.56 | 49 | 0.237
US50 | RhRh | 0.58 | 28 | 0.234
US50 | RhRh | 0.58 | 29 | 0.234
US50 | RhRh | 0.58 | 30 | 0.234
US50 | RhRh | 0.58 | 31 | 0.235
US50 | RhRh | 0.58 | 32 | 0.235
US50 | RhRh | 0.58 | 33 | 0.235
US50 | RhRh | 0.58 | 34 | 0.236
US50 | RhRh | 0.58 | 35 | 0.236
US50 | RhRh | 0.58 | 36 | 0.236
US50 | RhRh | 0.58 | 37 | 0.237
US50 | RhRh | 0.58 | 38 | 0.237
US50 | RhRh | 0.58 | 39 | 0.238
US50 | RhRh | 0.58 | 40 | 0.238
US50 | RhRh | 0.58 | 41 | 0.239
US50 | RhRh | 0.58 | 42 | 0.239
US50 | RhRh | 0.58 | 43 | 0.24
US50 | RhRh | 0.58 | 44 | 0.24
US50 | RhRh | 0.58 | 45 | 0.241
US50 | RhRh | 0.58 | 46 | 0.241
US50 | RhRh | 0.58 | 47 | 0.242
US50 | RhRh | 0.58 | 48 | 0.242
US50 | RhRh | 0.58 | 49 | 0.243
US50 | RhRh | 0.6 | 30 | 0.241
US50 | RhRh | 0.6 | 31 | 0.241
US50 | RhRh | 0.6 | 32 | 0.242
US50 | RhRh | 0.6 | 33 | 0.242
US50 | RhRh | 0.6 | 34 | 0.242
US50 | RhRh | 0.6 | 35 | 0.242
US50 | RhRh | 0.6 | 36 | 0.243
US50 | RhRh | 0.6 | 37 | 0.243
US50 | RhRh | 0.6 | 38 | 0.243
US50 | RhRh | 0.6 | 39 | 0.244
US50 | RhRh | 0.6 | 40 | 0.244
US50 | RhRh | 0.6 | 41 | 0.245
US50 | RhRh | 0.6 | 42 | 0.245
US50 | RhRh | 0.6 | 43 | 0.246
US50 | RhRh | 0.6 | 44 | 0.246
US50 | RhRh | 0.6 | 45 | 0.247
US50 | RhRh | 0.6 | 46 | 0.247
US50 | RhRh | 0.6 | 47 | 0.248
US50 | RhRh | 0.6 | 48 | 0.248
US50 | RhRh | 0.6 | 49 | 0.249
US50 | RhRh | 0.62 | 33 | 0.249
US50 | RhRh | 0.62 | 34 | 0.249
US50 | RhRh | 0.62 | 35 | 0.249
US50 | RhRh | 0.62 | 36 | 0.249
US50 | RhRh | 0.62 | 37 | 0.25
US50 | RhRh | 0.62 | 38 | 0.25
US50 | RhRh | 0.62 | 39 | 0.25
US50 | RhRh | 0.62 | 40 | 0.251
US50 | RhRh | 0.62 | 41 | 0.251
US50 | RhRh | 0.62 | 42 | 0.251
US50 | RhRh | 0.62 | 43 | 0.252
US50 | RhRh | 0.62 | 44 | 0.252
US50 | RhRh | 0.62 | 45 | 0.253
US50 | RhRh | 0.62 | 46 | 0.253
US50 | RhRh | 0.62 | 47 | 0.254
US50 | RhRh | 0.62 | 48 | 0.254
US50 | RhRh | 0.62 | 49 | 0.255
US50 | RhRh | 0.64 | 35 | 0.256
US50 | RhRh | 0.64 | 36 | 0.256
US50 | RhRh | 0.64 | 37 | 0.256
US50 | RhRh | 0.64 | 38 | 0.256
US50 | RhRh | 0.64 | 39 | 0.257
US50 | RhRh | 0.64 | 40 | 0.257
US50 | RhRh | 0.64 | 41 | 0.257
US50 | RhRh | 0.64 | 42 | 0.258
US50 | RhRh | 0.64 | 43 | 0.258
US50 | RhRh | 0.64 | 44 | 0.258
US50 | RhRh | 0.64 | 45 | 0.259
US50 | RhRh | 0.64 | 46 | 0.259
US50 | RhRh | 0.64 | 47 | 0.26
US50 | RhRh | 0.64 | 48 | 0.26
US50 | RhRh | 0.64 | 49 | 0.261
US50 | RhRh | 0.66 | 38 | 0.263
US50 | RhRh | 0.66 | 39 | 0.263
US50 | RhRh | 0.66 | 40 | 0.263
US50 | RhRh | 0.66 | 41 | 0.263
US50 | RhRh | 0.66 | 42 | 0.264
US50 | RhRh | 0.66 | 43 | 0.264
US50 | RhRh | 0.66 | 44 | 0.265
US50 | RhRh | 0.66 | 45 | 0.265
US50 | RhRh | 0.66 | 46 | 0.265
US50 | RhRh | 0.66 | 47 | 0.266
US50 | RhRh | 0.66 | 48 | 0.266
US50 | RhRh | 0.66 | 49 | 0.266
US50 | RhRh | 0.68 | 40 | 0.269
US50 | RhRh | 0.68 | 41 | 0.27
US50 | RhRh | 0.68 | 42 | 0.27
US50 | RhRh | 0.68 | 43 | 0.27
US50 | RhRh | 0.68 | 44 | 0.271
US50 | RhRh | 0.68 | 45 | 0.271
US50 | RhRh | 0.68 | 46 | 0.271
US50 | RhRh | 0.68 | 47 | 0.272
US50 | RhRh | 0.68 | 48 | 0.272
US50 | RhRh | 0.68 | 49 | 0.273
US50 | RhRh | 0.7 | 43 | 0.277
US50 | RhRh | 0.7 | 44 | 0.277
US50 | RhRh | 0.7 | 45 | 0.277
US50 | RhRh | 0.7 | 46 | 0.278
US50 | RhRh | 0.7 | 47 | 0.278
US50 | RhRh | 0.7 | 48 | 0.278
US50 | RhRh | 0.7 | 49 | 0.279
US50 | RhRh | 0.72 | 47 | 0.284
US50 | RhRh | 0.72 | 48 | 0.285
US50 | RhRh | 0.72 | 49 | 0.285
US50 | WAg | 0.175 | 20 | 0.046
US50 | WAg | 0.175 | 21 | 0.05
US50 | WAg | 0.175 | 22 | 0.053
US50 | WAg | 0.2 | 20 | 0.057
US50 | WAg | 0.2 | 21 | 0.061
US50 | WAg | 0.2 | 22 | 0.065
US50 | WAg | 0.2 | 23 | 0.068
US50 | WAg | 0.2 | 24 | 0.071
US50 | WAg | 0.2 | 25 | 0.074
US50 | WAg | 0.2 | 26 | 0.077
US50 | WAg | 0.2 | 27 | 0.08
US50 | WAg | 0.225 | 20 | 0.068
US50 | WAg | 0.225 | 21 | 0.072
US50 | WAg | 0.225 | 22 | 0.076
US50 | WAg | 0.225 | 23 | 0.08
US50 | WAg | 0.225 | 24 | 0.083
US50 | WAg | 0.225 | 25 | 0.086
US50 | WAg | 0.225 | 26 | 0.09
US50 | WAg | 0.225 | 27 | 0.092
US50 | WAg | 0.25 | 20 | 0.078
US50 | WAg | 0.25 | 21 | 0.083
US50 | WAg | 0.25 | 22 | 0.087
US50 | WAg | 0.25 | 23 | 0.091
US50 | WAg | 0.25 | 24 | 0.095
US50 | WAg | 0.25 | 25 | 0.098
US50 | WAg | 0.25 | 26 | 0.101
US50 | WAg | 0.25 | 27 | 0.104
US50 | WAg | 0.25 | 28 | 0.107
US50 | WAg | 0.25 | 29 | 0.109
US50 | WAg | 0.25 | 30 | 0.111
US50 | WAg | 0.25 | 31 | 0.114
US50 | WAg | 0.275 | 20 | 0.088
US50 | WAg | 0.275 | 21 | 0.093
US50 | WAg | 0.275 | 22 | 0.097
US50 | WAg | 0.275 | 23 | 0.102
US50 | WAg | 0.275 | 24 | 0.106
US50 | WAg | 0.275 | 25 | 0.109
US50 | WAg | 0.275 | 26 | 0.113
US50 | WAg | 0.275 | 27 | 0.116
US50 | WAg | 0.275 | 28 | 0.118
US50 | WAg | 0.275 | 29 | 0.121
US50 | WAg | 0.275 | 30 | 0.123
US50 | WAg | 0.275 | 31 | 0.125
US50 | WAg | 0.275 | 32 | 0.127
US50 | WAg | 0.275 | 33 | 0.129
US50 | WAg | 0.275 | 34 | 0.131
US50 | WAg | 0.3 | 20 | 0.099
US50 | WAg | 0.3 | 21 | 0.104
US50 | WAg | 0.3 | 22 | 0.108
US50 | WAg | 0.3 | 23 | 0.112
US50 | WAg | 0.3 | 24 | 0.116
US50 | WAg | 0.3 | 25 | 0.12
US50 | WAg | 0.3 | 26 | 0.123
US50 | WAg | 0.3 | 27 | 0.127
US50 | WAg | 0.3 | 28 | 0.129
US50 | WAg | 0.3 | 29 | 0.131
US50 | WAg | 0.3 | 30 | 0.134
US50 | WAg | 0.3 | 31 | 0.135
US50 | WAg | 0.3 | 32 | 0.137
US50 | WAg | 0.3 | 33 | 0.139
US50 | WAg | 0.3 | 34 | 0.141
US50 | WAg | 0.3 | 35 | 0.143
US50 | WAg | 0.3 | 36 | 0.145
US50 | WAg | 0.3 | 37 | 0.146
US50 | WAg | 0.3 | 38 | 0.148
US50 | WAg | 0.3 | 39 | 0.15
US50 | WAg | 0.3 | 40 | 0.151
US50 | WAg | 0.3 | 41 | 0.153
US50 | WAg | 0.3 | 42 | 0.154
US50 | WAg | 0.3 | 43 | 0.156
US50 | WAg | 0.325 | 20 | 0.109
US50 | WAg | 0.325 | 21 | 0.114
US50 | WAg | 0.325 | 22 | 0.118
US50 | WAg | 0.325 | 23 | 0.123
US50 | WAg | 0.325 | 24 | 0.127
US50 | WAg | 0.325 | 25 | 0.131
US50 | WAg | 0.325 | 26 | 0.134
US50 | WAg | 0.325 | 27 | 0.137
US50 | WAg | 0.325 | 28 | 0.14
US50 | WAg | 0.325 | 29 | 0.142
US50 | WAg | 0.325 | 30 | 0.144
US50 | WAg | 0.325 | 31 | 0.146
US50 | WAg | 0.325 | 32 | 0.148
US50 | WAg | 0.325 | 33 | 0.149
US50 | WAg | 0.325 | 34 | 0.151
US50 | WAg | 0.325 | 35 | 0.153
US50 | WAg | 0.325 | 36 | 0.155
US50 | WAg | 0.325 | 37 | 0.156
US50 | WAg | 0.325 | 38 | 0.158
US50 | WAg | 0.325 | 39 | 0.16
US50 | WAg | 0.325 | 40 | 0.161
US50 | WAg | 0.325 | 41 | 0.163
US50 | WAg | 0.325 | 42 | 0.164
US50 | WAg | 0.325 | 43 | 0.166
US50 | WAg | 0.325 | 44 | 0.167
US50 | WAg | 0.325 | 45 | 0.169
US50 | WAg | 0.325 | 46 | 0.17
US50 | WAg | 0.325 | 47 | 0.172
US50 | WAg | 0.325 | 48 | 0.173
US50 | WAg | 0.325 | 49 | 0.174
US50 | WAg | 0.35 | 20 | 0.12
US50 | WAg | 0.35 | 21 | 0.125
US50 | WAg | 0.35 | 22 | 0.129
US50 | WAg | 0.35 | 23 | 0.133
US50 | WAg | 0.35 | 24 | 0.137
US50 | WAg | 0.35 | 25 | 0.141
US50 | WAg | 0.35 | 26 | 0.144
US50 | WAg | 0.35 | 27 | 0.147
US50 | WAg | 0.35 | 28 | 0.15
US50 | WAg | 0.35 | 29 | 0.152
US50 | WAg | 0.35 | 30 | 0.154
US50 | WAg | 0.35 | 31 | 0.156
US50 | WAg | 0.35 | 32 | 0.158
US50 | WAg | 0.35 | 33 | 0.159
US50 | WAg | 0.35 | 34 | 0.161
US50 | WAg | 0.35 | 35 | 0.163
US50 | WAg | 0.35 | 36 | 0.164
US50 | WAg | 0.35 | 37 | 0.166
US50 | WAg | 0.35 | 38 | 0.167
US50 | WAg | 0.35 | 39 | 0.169
US50 | WAg | 0.35 | 40 | 0.171
US50 | WAg | 0.35 | 41 | 0.172
US50 | WAg | 0.35 | 42 | 0.174
US50 | WAg | 0.35 | 43 | 0.175
US50 | WAg | 0.35 | 44 | 0.177
US50 | WAg | 0.35 | 45 | 0.178
US50 | WAg | 0.35 | 46 | 0.18
US50 | WAg | 0.35 | 47 | 0.181
US50 | WAg | 0.35 | 48 | 0.182
US50 | WAg | 0.35 | 49 | 0.184
US50 | WAg | 0.375 | 20 | 0.131
US50 | WAg | 0.375 | 21 | 0.135
US50 | WAg | 0.375 | 22 | 0.14
US50 | WAg | 0.375 | 23 | 0.144
US50 | WAg | 0.375 | 24 | 0.148
US50 | WAg | 0.375 | 25 | 0.151
US50 | WAg | 0.375 | 26 | 0.155
US50 | WAg | 0.375 | 27 | 0.158
US50 | WAg | 0.375 | 28 | 0.16
US50 | WAg | 0.375 | 29 | 0.162
US50 | WAg | 0.375 | 30 | 0.164
US50 | WAg | 0.375 | 31 | 0.166
US50 | WAg | 0.375 | 32 | 0.167
US50 | WAg | 0.375 | 33 | 0.169
US50 | WAg | 0.375 | 34 | 0.17
US50 | WAg | 0.375 | 35 | 0.172
US50 | WAg | 0.375 | 36 | 0.173
US50 | WAg | 0.375 | 37 | 0.175
US50 | WAg | 0.375 | 38 | 0.177
US50 | WAg | 0.375 | 39 | 0.178
US50 | WAg | 0.375 | 40 | 0.18
US50 | WAg | 0.375 | 41 | 0.181
US50 | WAg | 0.375 | 42 | 0.183
US50 | WAg | 0.375 | 43 | 0.184
US50 | WAg | 0.375 | 44 | 0.186
US50 | WAg | 0.375 | 45 | 0.187
US50 | WAg | 0.375 | 46 | 0.189
US50 | WAg | 0.375 | 47 | 0.19
US50 | WAg | 0.375 | 48 | 0.191
US50 | WAg | 0.375 | 49 | 0.193
US50 | WAg | 0.4 | 20 | 0.141
US50 | WAg | 0.4 | 21 | 0.146
US50 | WAg | 0.4 | 22 | 0.15
US50 | WAg | 0.4 | 23 | 0.154
US50 | WAg | 0.4 | 24 | 0.158
US50 | WAg | 0.4 | 25 | 0.161
US50 | WAg | 0.4 | 26 | 0.165
US50 | WAg | 0.4 | 27 | 0.168
US50 | WAg | 0.4 | 28 | 0.17
US50 | WAg | 0.4 | 29 | 0.172
US50 | WAg | 0.4 | 30 | 0.174
US50 | WAg | 0.4 | 31 | 0.175
US50 | WAg | 0.4 | 32 | 0.177
US50 | WAg | 0.4 | 33 | 0.178
US50 | WAg | 0.4 | 34 | 0.18
US50 | WAg | 0.4 | 35 | 0.181
US50 | WAg | 0.4 | 36 | 0.183
US50 | WAg | 0.4 | 37 | 0.184
US50 | WAg | 0.4 | 38 | 0.186
US50 | WAg | 0.4 | 39 | 0.187
US50 | WAg | 0.4 | 40 | 0.188
US50 | WAg | 0.4 | 41 | 0.19
US50 | WAg | 0.4 | 42 | 0.192
US50 | WAg | 0.4 | 43 | 0.193
US50 | WAg | 0.4 | 44 | 0.194
US50 | WAg | 0.4 | 45 | 0.196
US50 | WAg | 0.4 | 46 | 0.197
US50 | WAg | 0.4 | 47 | 0.199
US50 | WAg | 0.4 | 48 | 0.2
US50 | WAg | 0.4 | 49 | 0.201
US50 | WAg | 0.425 | 20 | 0.152
US50 | WAg | 0.425 | 21 | 0.156
US50 | WAg | 0.425 | 22 | 0.161
US50 | WAg | 0.425 | 23 | 0.165
US50 | WAg | 0.425 | 24 | 0.168
US50 | WAg | 0.425 | 25 | 0.172
US50 | WAg | 0.425 | 26 | 0.175
US50 | WAg | 0.425 | 27 | 0.178
US50 | WAg | 0.425 | 28 | 0.18
US50 | WAg | 0.425 | 29 | 0.182
US50 | WAg | 0.425 | 30 | 0.183
US50 | WAg | 0.425 | 31 | 0.185
US50 | WAg | 0.425 | 32 | 0.186
US50 | WAg | 0.425 | 33 | 0.187
US50 | WAg | 0.425 | 34 | 0.189
US50 | WAg | 0.425 | 35 | 0.19
US50 | WAg | 0.425 | 36 | 0.191
US50 | WAg | 0.425 | 37 | 0.193
US50 | WAg | 0.425 | 38 | 0.194
US50 | WAg | 0.425 | 39 | 0.196
US50 | WAg | 0.425 | 40 | 0.197
US50 | WAg | 0.425 | 41 | 0.199
US50 | WAg | 0.425 | 42 | 0.2
US50 | WAg | 0.425 | 43 | 0.202
US50 | WAg | 0.425 | 44 | 0.203
US50 | WAg | 0.425 | 45 | 0.204
US50 | WAg | 0.425 | 46 | 0.206
US50 | WAg | 0.425 | 47 | 0.207
US50 | WAg | 0.425 | 48 | 0.208
US50 | WAg | 0.425 | 49 | 0.21
US50 | WAg | 0.45 | 20 | 0.163
US50 | WAg | 0.45 | 21 | 0.167
US50 | WAg | 0.45 | 22 | 0.171
US50 | WAg | 0.45 | 23 | 0.175
US50 | WAg | 0.45 | 24 | 0.179
US50 | WAg | 0.45 | 25 | 0.182
US50 | WAg | 0.45 | 26 | 0.185
US50 | WAg | 0.45 | 27 | 0.188
US50 | WAg | 0.45 | 28 | 0.19
US50 | WAg | 0.45 | 29 | 0.191
US50 | WAg | 0.45 | 30 | 0.193
US50 | WAg | 0.45 | 31 | 0.194
US50 | WAg | 0.45 | 32 | 0.195
US50 | WAg | 0.45 | 33 | 0.197
US50 | WAg | 0.45 | 34 | 0.198
US50 | WAg | 0.45 | 35 | 0.199
US50 | WAg | 0.45 | 36 | 0.2
US50 | WAg | 0.45 | 37 | 0.202
US50 | WAg | 0.45 | 38 | 0.203
US50 | WAg | 0.45 | 39 | 0.204
US50 | WAg | 0.45 | 40 | 0.206
US50 | WAg | 0.45 | 41 | 0.207
US50 | WAg | 0.45 | 42 | 0.208
US50 | WAg | 0.45 | 43 | 0.21
US50 | WAg | 0.45 | 44 | 0.211
US50 | WAg | 0.45 | 45 | 0.212
US50 | WAg | 0.45 | 46 | 0.214
US50 | WAg | 0.45 | 47 | 0.215
US50 | WAg | 0.45 | 48 | 0.217
US50 | WAg | 0.45 | 49 | 0.218
US50 | WAg | 0.475 | 20 | 0.174
US50 | WAg | 0.475 | 21 | 0.178
US50 | WAg | 0.475 | 22 | 0.181
US50 | WAg | 0.475 | 23 | 0.185
US50 | WAg | 0.475 | 24 | 0.189
US50 | WAg | 0.475 | 25 | 0.192
US50 | WAg | 0.475 | 26 | 0.195
US50 | WAg | 0.475 | 27 | 0.197
US50 | WAg | 0.475 | 28 | 0.199
US50 | WAg | 0.475 | 29 | 0.201
US50 | WAg | 0.475 | 30 | 0.202
US50 | WAg | 0.475 | 31 | 0.203
US50 | WAg | 0.475 | 32 | 0.204
US50 | WAg | 0.475 | 33 | 0.206
US50 | WAg | 0.475 | 34 | 0.207
US50 | WAg | 0.475 | 35 | 0.208
US50 | WAg | 0.475 | 36 | 0.209
US50 | WAg | 0.475 | 37 | 0.21
US50 | WAg | 0.475 | 38 | 0.211
US50 | WAg | 0.475 | 39 | 0.213
US50 | WAg | 0.475 | 40 | 0.214
US50 | WAg | 0.475 | 41 | 0.215
US50 | WAg | 0.475 | 42 | 0.217
US50 | WAg | 0.475 | 43 | 0.218
US50 | WAg | 0.475 | 44 | 0.219
US50 | WAg | 0.475 | 45 | 0.221
US50 | WAg | 0.475 | 46 | 0.222
US50 | WAg | 0.475 | 47 | 0.223
US50 | WAg | 0.475 | 48 | 0.225
US50 | WAg | 0.475 | 49 | 0.226
US50 | WAg | 0.5 | 21 | 0.188
US50 | WAg | 0.5 | 22 | 0.192
US50 | WAg | 0.5 | 23 | 0.195
US50 | WAg | 0.5 | 24 | 0.199
US50 | WAg | 0.5 | 25 | 0.202
US50 | WAg | 0.5 | 26 | 0.205
US50 | WAg | 0.5 | 27 | 0.207
US50 | WAg | 0.5 | 28 | 0.209
US50 | WAg | 0.5 | 29 | 0.21
US50 | WAg | 0.5 | 30 | 0.211
US50 | WAg | 0.5 | 31 | 0.212
US50 | WAg | 0.5 | 32 | 0.213
US50 | WAg | 0.5 | 33 | 0.215
US50 | WAg | 0.5 | 34 | 0.216
US50 | WAg | 0.5 | 35 | 0.216
US50 | WAg | 0.5 | 36 | 0.218
US50 | WAg | 0.5 | 37 | 0.219
US50 | WAg | 0.5 | 38 | 0.22
US50 | WAg | 0.5 | 39 | 0.221
US50 | WAg | 0.5 | 40 | 0.222
US50 | WAg | 0.5 | 41 | 0.224
US50 | WAg | 0.5 | 42 | 0.225
US50 | WAg | 0.5 | 43 | 0.226
US50 | WAg | 0.5 | 44 | 0.227
US50 | WAg | 0.5 | 45 | 0.229
US50 | WAg | 0.5 | 46 | 0.23
US50 | WAg | 0.5 | 47 | 0.231
US50 | WAg | 0.5 | 48 | 0.232
US50 | WAg | 0.5 | 49 | 0.234
US50 | WAg | 0.525 | 21 | 0.199
US50 | WAg | 0.525 | 22 | 0.202
US50 | WAg | 0.525 | 23 | 0.205
US50 | WAg | 0.525 | 24 | 0.209
US50 | WAg | 0.525 | 25 | 0.211
US50 | WAg | 0.525 | 26 | 0.214
US50 | WAg | 0.525 | 27 | 0.217
US50 | WAg | 0.525 | 28 | 0.218
US50 | WAg | 0.525 | 29 | 0.219
US50 | WAg | 0.525 | 30 | 0.221
US50 | WAg | 0.525 | 31 | 0.222
US50 | WAg | 0.525 | 32 | 0.222
US50 | WAg | 0.525 | 33 | 0.223
US50 | WAg | 0.525 | 34 | 0.224
US50 | WAg | 0.525 | 35 | 0.225
US50 | WAg | 0.525 | 36 | 0.226
US50 | WAg | 0.525 | 37 | 0.227
US50 | WAg | 0.525 | 38 | 0.228
US50 | WAg | 0.525 | 39 | 0.229
US50 | WAg | 0.525 | 40 | 0.23
US50 | WAg | 0.525 | 41 | 0.232
US50 | WAg | 0.525 | 42 | 0.233
US50 | WAg | 0.525 | 43 | 0.234
US50 | WAg | 0.525 | 44 | 0.235
US50 | WAg | 0.525 | 45 | 0.237
US50 | WAg | 0.525 | 46 | 0.238
US50 | WAg | 0.525 | 47 | 0.239
US50 | WAg | 0.525 | 48 | 0.24
US50 | WAg | 0.525 | 49 | 0.241
US50 | WAg | 0.55 | 22 | 0.212
US50 | WAg | 0.55 | 23 | 0.215
US50 | WAg | 0.55 | 24 | 0.218
US50 | WAg | 0.55 | 25 | 0.221
US50 | WAg | 0.55 | 26 | 0.224
US50 | WAg | 0.55 | 27 | 0.226
US50 | WAg | 0.55 | 28 | 0.227
US50 | WAg | 0.55 | 29 | 0.229
US50 | WAg | 0.55 | 30 | 0.23
US50 | WAg | 0.55 | 31 | 0.23
US50 | WAg | 0.55 | 32 | 0.231
US50 | WAg | 0.55 | 33 | 0.232
US50 | WAg | 0.55 | 34 | 0.233
US50 | WAg | 0.55 | 35 | 0.234
US50 | WAg | 0.55 | 36 | 0.235
US50 | WAg | 0.55 | 37 | 0.236
US50 | WAg | 0.55 | 38 | 0.237
US50 | WAg | 0.55 | 39 | 0.238
US50 | WAg | 0.55 | 40 | 0.239
US50 | WAg | 0.55 | 41 | 0.24
US50 | WAg | 0.55 | 42 | 0.241
US50 | WAg | 0.55 | 43 | 0.242
US50 | WAg | 0.55 | 44 | 0.243
US50 | WAg | 0.55 | 45 | 0.244
US50 | WAg | 0.55 | 46 | 0.246
US50 | WAg | 0.55 | 47 | 0.247
US50 | WAg | 0.55 | 48 | 0.248
US50 | WAg | 0.55 | 49 | 0.249
US50 | WAg | 0.575 | 22 | 0.222
US50 | WAg | 0.575 | 23 | 0.225
US50 | WAg | 0.575 | 24 | 0.228
US50 | WAg | 0.575 | 25 | 0.23
US50 | WAg | 0.575 | 26 | 0.233
US50 | WAg | 0.575 | 27 | 0.235
US50 | WAg | 0.575 | 28 | 0.237
US50 | WAg | 0.575 | 29 | 0.238
US50 | WAg | 0.575 | 30 | 0.239
US50 | WAg | 0.575 | 31 | 0.239
US50 | WAg | 0.575 | 32 | 0.24
US50 | WAg | 0.575 | 33 | 0.241
US50 | WAg | 0.575 | 34 | 0.242
US50 | WAg | 0.575 | 35 | 0.242
US50 | WAg | 0.575 | 36 | 0.243
US50 | WAg | 0.575 | 37 | 0.244
US50 | WAg | 0.575 | 38 | 0.245
US50 | WAg | 0.575 | 39 | 0.246
US50 | WAg | 0.575 | 40 | 0.247
US50 | WAg | 0.575 | 41 | 0.248
US50 | WAg | 0.575 | 42 | 0.249
US50 | WAg | 0.575 | 43 | 0.25
US50 | WAg | 0.575 | 44 | 0.251
US50 | WAg | 0.575 | 45 | 0.252
US50 | WAg | 0.575 | 46 | 0.253
US50 | WAg | 0.575 | 47 | 0.254
US50 | WAg | 0.575 | 48 | 0.256
US50 | WAg | 0.575 | 49 | 0.257
US50 | WAg | 0.6 | 23 | 0.235
US50 | WAg | 0.6 | 24 | 0.237
US50 | WAg | 0.6 | 25 | 0.24
US50 | WAg | 0.6 | 26 | 0.242
US50 | WAg | 0.6 | 27 | 0.244
US50 | WAg | 0.6 | 28 | 0.246
US50 | WAg | 0.6 | 29 | 0.247
US50 | WAg | 0.6 | 30 | 0.247
US50 | WAg | 0.6 | 31 | 0.248
US50 | WAg | 0.6 | 32 | 0.249
US50 | WAg | 0.6 | 33 | 0.249
US50 | WAg | 0.6 | 34 | 0.25
US50 | WAg | 0.6 | 35 | 0.251
US50 | WAg | 0.6 | 36 | 0.251
US50 | WAg | 0.6 | 37 | 0.252
US50 | WAg | 0.6 | 38 | 0.253
US50 | WAg | 0.6 | 39 | 0.254
US50 | WAg | 0.6 | 40 | 0.255
US50 | WAg | 0.6 | 41 | 0.256
US50 | WAg | 0.6 | 42 | 0.257
US50 | WAg | 0.6 | 43 | 0.258
US50 | WAg | 0.6 | 44 | 0.259
US50 | WAg | 0.6 | 45 | 0.26
US50 | WAg | 0.6 | 46 | 0.261
US50 | WAg | 0.6 | 47 | 0.262
US50 | WAg | 0.6 | 48 | 0.263
US50 | WAg | 0.6 | 49 | 0.264
US50 | WAg | 0.625 | 23 | 0.244
US50 | WAg | 0.625 | 24 | 0.247
US50 | WAg | 0.625 | 25 | 0.249
US50 | WAg | 0.625 | 26 | 0.251
US50 | WAg | 0.625 | 27 | 0.253
US50 | WAg | 0.625 | 28 | 0.254
US50 | WAg | 0.625 | 29 | 0.255
US50 | WAg | 0.625 | 30 | 0.256
US50 | WAg | 0.625 | 31 | 0.257
US50 | WAg | 0.625 | 32 | 0.257
US50 | WAg | 0.625 | 33 | 0.258
US50 | WAg | 0.625 | 34 | 0.258
US50 | WAg | 0.625 | 35 | 0.259
US50 | WAg | 0.625 | 36 | 0.259
US50 | WAg | 0.625 | 37 | 0.26
US50 | WAg | 0.625 | 38 | 0.261
US50 | WAg | 0.625 | 39 | 0.262
US50 | WAg | 0.625 | 40 | 0.263
US50 | WAg | 0.625 | 41 | 0.263
US50 | WAg | 0.625 | 42 | 0.264
US50 | WAg | 0.625 | 43 | 0.265
US50 | WAg | 0.625 | 44 | 0.266
US50 | WAg | 0.625 | 45 | 0.267
US50 | WAg | 0.625 | 46 | 0.268
US50 | WAg | 0.625 | 47 | 0.269
US50 | WAg | 0.625 | 48 | 0.271
US50 | WAg | 0.625 | 49 | 0.272
US50 | WAg | 0.65 | 28 | 0.263
US50 | WAg | 0.65 | 29 | 0.264
US50 | WAg | 0.65 | 30 | 0.265
US50 | WAg | 0.65 | 31 | 0.265
US50 | WAg | 0.65 | 32 | 0.266
US50 | WAg | 0.65 | 33 | 0.266
US50 | WAg | 0.65 | 34 | 0.267
US50 | WAg | 0.65 | 35 | 0.267
US50 | WAg | 0.65 | 36 | 0.268
US50 | WAg | 0.65 | 37 | 0.268
US50 | WAg | 0.65 | 38 | 0.269
US50 | WAg | 0.65 | 39 | 0.27
US50 | WAg | 0.65 | 40 | 0.27
US50 | WAg | 0.65 | 41 | 0.271
US50 | WAg | 0.65 | 42 | 0.272
US50 | WAg | 0.65 | 43 | 0.273
US50 | WAg | 0.65 | 44 | 0.274
US50 | WAg | 0.65 | 45 | 0.275
US50 | WAg | 0.65 | 46 | 0.276
US50 | WAg | 0.65 | 47 | 0.277
US50 | WAg | 0.65 | 48 | 0.278
US50 | WAg | 0.65 | 49 | 0.279
US50 | WAg | 0.675 | 28 | 0.272
US50 | WAg | 0.675 | 29 | 0.273
US50 | WAg | 0.675 | 30 | 0.273
US50 | WAg | 0.675 | 31 | 0.274
US50 | WAg | 0.675 | 32 | 0.274
US50 | WAg | 0.675 | 33 | 0.274
US50 | WAg | 0.675 | 34 | 0.275
US50 | WAg | 0.675 | 35 | 0.275
US50 | WAg | 0.675 | 36 | 0.276
US50 | WAg | 0.675 | 37 | 0.276
US50 | WAg | 0.675 | 38 | 0.277
US50 | WAg | 0.675 | 39 | 0.278
US50 | WAg | 0.675 | 40 | 0.278
US50 | WAg | 0.675 | 41 | 0.279
US50 | WAg | 0.675 | 42 | 0.28
US50 | WAg | 0.675 | 43 | 0.281
US50 | WAg | 0.675 | 44 | 0.282
US50 | WAg | 0.675 | 45 | 0.282
US50 | WAg | 0.675 | 46 | 0.283
US50 | WAg | 0.675 | 47 | 0.284
US50 | WAg | 0.675 | 48 | 0.285
US50 | WAg | 0.675 | 49 | 0.286
US50 | WAg | 0.7 | 32 | 0.282
US50 | WAg | 0.7 | 33 | 0.282
US50 | WAg | 0.7 | 34 | 0.283
US50 | WAg | 0.7 | 35 | 0.283
US50 | WAg | 0.7 | 36 | 0.284
US50 | WAg | 0.7 | 37 | 0.284
US50 | WAg | 0.7 | 38 | 0.285
US50 | WAg | 0.7 | 39 | 0.285
US50 | WAg | 0.7 | 40 | 0.286
US50 | WAg | 0.7 | 41 | 0.286
US50 | WAg | 0.7 | 42 | 0.287
US50 | WAg | 0.7 | 43 | 0.288
US50 | WAg | 0.7 | 44 | 0.289
US50 | WAg | 0.7 | 45 | 0.29
US50 | WAg | 0.7 | 46 | 0.291
US50 | WAg | 0.7 | 47 | 0.292
US50 | WAg | 0.7 | 48 | 0.293
US50 | WAg | 0.7 | 49 | 0.293
US50 | WAg | 0.725 | 35 | 0.291
US50 | WAg | 0.725 | 36 | 0.291
US50 | WAg | 0.725 | 37 | 0.292
US50 | WAg | 0.725 | 38 | 0.292
US50 | WAg | 0.725 | 39 | 0.293
US50 | WAg | 0.725 | 40 | 0.293
US50 | WAg | 0.725 | 41 | 0.294
US50 | WAg | 0.725 | 42 | 0.295
US50 | WAg | 0.725 | 43 | 0.295
US50 | WAg | 0.725 | 44 | 0.296
US50 | WAg | 0.725 | 45 | 0.297
US50 | WAg | 0.725 | 46 | 0.298
US50 | WAg | 0.725 | 47 | 0.299
US50 | WAg | 0.725 | 48 | 0.3
US50 | WAg | 0.725 | 49 | 0.301
US50 | WAg | 0.75 | 44 | 0.304
US50 | WAg | 0.75 | 45 | 0.304
US50 | WAg | 0.75 | 46 | 0.305
US50 | WAg | 0.75 | 47 | 0.306
US50 | WAg | 0.75 | 48 | 0.307
US50 | WAg | 0.75 | 49 | 0.308
US50 | WAl | 0.2 | 20 | 0.057
US50 | WAl | 0.2 | 21 | 0.061
US50 | WAl | 0.25 | 20 | 0.078
US50 | WAl | 0.25 | 21 | 0.082
US50 | WAl | 0.25 | 22 | 0.086
US50 | WAl | 0.25 | 23 | 0.09
US50 | WAl | 0.25 | 24 | 0.094
US50 | WAl | 0.25 | 25 | 0.097
US50 | WAl | 0.3 | 20 | 0.098
US50 | WAl | 0.3 | 21 | 0.103
US50 | WAl | 0.3 | 22 | 0.107
US50 | WAl | 0.3 | 23 | 0.111
US50 | WAl | 0.3 | 24 | 0.115
US50 | WAl | 0.3 | 25 | 0.119
US50 | WAl | 0.3 | 26 | 0.122
US50 | WAl | 0.3 | 27 | 0.126
US50 | WAl | 0.3 | 28 | 0.129
US50 | WAl | 0.3 | 29 | 0.132
US50 | WAl | 0.3 | 30 | 0.135
US50 | WAl | 0.35 | 20 | 0.12
US50 | WAl | 0.35 | 21 | 0.124
US50 | WAl | 0.35 | 22 | 0.128
US50 | WAl | 0.35 | 23 | 0.132
US50 | WAl | 0.35 | 24 | 0.136
US50 | WAl | 0.35 | 25 | 0.14
US50 | WAl | 0.35 | 26 | 0.143
US50 | WAl | 0.35 | 27 | 0.147
US50 | WAl | 0.35 | 28 | 0.15
US50 | WAl | 0.35 | 29 | 0.153
US50 | WAl | 0.35 | 30 | 0.156
US50 | WAl | 0.35 | 31 | 0.158
US50 | WAl | 0.35 | 32 | 0.161
US50 | WAl | 0.35 | 33 | 0.164
US50 | WAl | 0.35 | 34 | 0.166
US50 | WAl | 0.35 | 35 | 0.168
US50 | WAl | 0.4 | 20 | 0.141
US50 | WAl | 0.4 | 21 | 0.145
US50 | WAl | 0.4 | 22 | 0.15
US50 | WAl | 0.4 | 23 | 0.153
US50 | WAl | 0.4 | 24 | 0.157
US50 | WAl | 0.4 | 25 | 0.161
US50 | WAl | 0.4 | 26 | 0.164
US50 | WAl | 0.4 | 27 | 0.167
US50 | WAl | 0.4 | 28 | 0.17
US50 | WAl | 0.4 | 29 | 0.173
US50 | WAl | 0.4 | 30 | 0.176
US50 | WAl | 0.4 | 31 | 0.178
US50 | WAl | 0.4 | 32 | 0.181
US50 | WAl | 0.4 | 33 | 0.183
US50 | WAl | 0.4 | 34 | 0.186
US50 | WAl | 0.4 | 35 | 0.188
US50 | WAl | 0.4 | 36 | 0.19
US50 | WAl | 0.4 | 37 | 0.192
US50 | WAl | 0.4 | 38 | 0.194
US50 | WAl | 0.4 | 39 | 0.196
US50 | WAl | 0.4 | 40 | 0.197
US50 | WAl | 0.45 | 20 | 0.163
US50 | WAl | 0.45 | 21 | 0.167
US50 | WAl | 0.45 | 22 | 0.171
US50 | WAl | 0.45 | 23 | 0.174
US50 | WAl | 0.45 | 24 | 0.178
US50 | WAl | 0.45 | 25 | 0.181
US50 | WAl | 0.45 | 26 | 0.184
US50 | WAl | 0.45 | 27 | 0.187
US50 | WAl | 0.45 | 28 | 0.19
US50 | WAl | 0.45 | 29 | 0.193
US50 | WAl | 0.45 | 30 | 0.195
US50 | WAl | 0.45 | 31 | 0.198
US50 | WAl | 0.45 | 32 | 0.2
US50 | WAl | 0.45 | 33 | 0.202
US50 | WAl | 0.45 | 34 | 0.205
US50 | WAl | 0.45 | 35 | 0.207
US50 | WAl | 0.45 | 36 | 0.209
US50 | WAl | 0.45 | 37 | 0.21
US50 | WAl | 0.45 | 38 | 0.212
US50 | WAl | 0.45 | 39 | 0.214
US50 | WAl | 0.45 | 40 | 0.216
US50 | WAl | 0.45 | 41 | 0.217
US50 | WAl | 0.45 | 42 | 0.219
US50 | WAl | 0.45 | 43 | 0.221
US50 | WAl | 0.45 | 44 | 0.222
US50 | WAl | 0.5 | 20 | 0.184
US50 | WAl | 0.5 | 21 | 0.188
US50 | WAl | 0.5 | 22 | 0.191
US50 | WAl | 0.5 | 23 | 0.195
US50 | WAl | 0.5 | 24 | 0.198
US50 | WAl | 0.5 | 25 | 0.201
US50 | WAl | 0.5 | 26 | 0.204
US50 | WAl | 0.5 | 27 | 0.207
US50 | WAl | 0.5 | 28 | 0.209
US50 | WAl | 0.5 | 29 | 0.212
US50 | WAl | 0.5 | 30 | 0.214
US50 | WAl | 0.5 | 31 | 0.217
US50 | WAl | 0.5 | 32 | 0.219
US50 | WAl | 0.5 | 33 | 0.221
US50 | WAl | 0.5 | 34 | 0.223
US50 | WAl | 0.5 | 35 | 0.225
US50 | WAl | 0.5 | 36 | 0.227
US50 | WAl | 0.5 | 37 | 0.228
US50 | WAl | 0.5 | 38 | 0.23
US50 | WAl | 0.5 | 39 | 0.232
US50 | WAl | 0.5 | 40 | 0.233
US50 | WAl | 0.5 | 41 | 0.235
US50 | WAl | 0.5 | 42 | 0.236
US50 | WAl | 0.5 | 43 | 0.238
US50 | WAl | 0.5 | 44 | 0.239
US50 | WAl | 0.5 | 45 | 0.24
US50 | WAl | 0.5 | 46 | 0.242
US50 | WAl | 0.5 | 47 | 0.243
US50 | WAl | 0.5 | 48 | 0.244
US50 | WAl | 0.5 | 49 | 0.245
US50 | WAl | 0.55 | 21 | 0.209
US50 | WAl | 0.55 | 22 | 0.212
US50 | WAl | 0.55 | 23 | 0.215
US50 | WAl | 0.55 | 24 | 0.218
US50 | WAl | 0.55 | 25 | 0.22
US50 | WAl | 0.55 | 26 | 0.223
US50 | WAl | 0.55 | 27 | 0.226
US50 | WAl | 0.55 | 28 | 0.228
US50 | WAl | 0.55 | 29 | 0.231
US50 | WAl | 0.55 | 30 | 0.233
US50 | WAl | 0.55 | 31 | 0.235
US50 | WAl | 0.55 | 32 | 0.237
US50 | WAl | 0.55 | 33 | 0.239
US50 | WAl | 0.55 | 34 | 0.241
US50 | WAl | 0.55 | 35 | 0.243
US50 | WAl | 0.55 | 36 | 0.244
US50 | WAl | 0.55 | 37 | 0.246
US50 | WAl | 0.55 | 38 | 0.247
US50 | WAl | 0.55 | 39 | 0.249
US50 | WAl | 0.55 | 40 | 0.25
US50 | WAl | 0.55 | 41 | 0.252
US50 | WAl | 0.55 | 42 | 0.253
US50 | WAl | 0.55 | 43 | 0.254
US50 | WAl | 0.55 | 44 | 0.256
US50 | WAl | 0.55 | 45 | 0.257
US50 | WAl | 0.55 | 46 | 0.258
US50 | WAl | 0.55 | 47 | 0.259
US50 | WAl | 0.55 | 48 | 0.26
US50 | WAl | 0.55 | 49 | 0.261
US50 | WAl | 0.6 | 22 | 0.232
US50 | WAl | 0.6 | 23 | 0.235
US50 | WAl | 0.6 | 24 | 0.237
US50 | WAl | 0.6 | 25 | 0.239
US50 | WAl | 0.6 | 26 | 0.242
US50 | WAl | 0.6 | 27 | 0.244
US50 | WAl | 0.6 | 28 | 0.247
US50 | WAl | 0.6 | 29 | 0.249
US50 | WAl | 0.6 | 30 | 0.251
US50 | WAl | 0.6 | 31 | 0.253
US50 | WAl | 0.6 | 32 | 0.255
US50 | WAl | 0.6 | 33 | 0.256
US50 | WAl | 0.6 | 34 | 0.258
US50 | WAl | 0.6 | 35 | 0.26
US50 | WAl | 0.6 | 36 | 0.261
US50 | WAl | 0.6 | 37 | 0.263
US50 | WAl | 0.6 | 38 | 0.264
US50 | WAl | 0.6 | 39 | 0.265
US50 | WAl | 0.6 | 40 | 0.267
US50 | WAl | 0.6 | 41 | 0.268
US50 | WAl | 0.6 | 42 | 0.269
US50 | WAl | 0.6 | 43 | 0.27
US50 | WAl | 0.6 | 44 | 0.272
US50 | WAl | 0.6 | 45 | 0.273
US50 | WAl | 0.6 | 46 | 0.274
US50 | WAl | 0.6 | 47 | 0.275
US50 | WAl | 0.6 | 48 | 0.276
US50 | WAl | 0.6 | 49 | 0.277
US50 | WAl | 0.65 | 23 | 0.253
US50 | WAl | 0.65 | 24 | 0.256
US50 | WAl | 0.65 | 25 | 0.258
US50 | WAl | 0.65 | 26 | 0.26
US50 | WAl | 0.65 | 27 | 0.262
US50 | WAl | 0.65 | 28 | 0.264
US50 | WAl | 0.65 | 29 | 0.266
US50 | WAl | 0.65 | 30 | 0.268
US50 | WAl | 0.65 | 31 | 0.27
US50 | WAl | 0.65 | 32 | 0.272
US50 | WAl | 0.65 | 33 | 0.273
US50 | WAl | 0.65 | 34 | 0.275
US50 | WAl | 0.65 | 35 | 0.276
US50 | WAl | 0.65 | 36 | 0.278
US50 | WAl | 0.65 | 37 | 0.279
US50 | WAl | 0.65 | 38 | 0.28
US50 | WAl | 0.65 | 39 | 0.281
US50 | WAl | 0.65 | 40 | 0.283
US50 | WAl | 0.65 | 41 | 0.284
US50 | WAl | 0.65 | 42 | 0.285
US50 | WAl | 0.65 | 43 | 0.286
US50 | WAl | 0.65 | 44 | 0.287
US50 | WAl | 0.65 | 45 | 0.288
US50 | WAl | 0.65 | 46 | 0.289
US50 | WAl | 0.65 | 47 | 0.29
US50 | WAl | 0.65 | 48 | 0.291
US50 | WAl | 0.65 | 49 | 0.292
US50 | WAl | 0.7 | 23 | 0.272
US50 | WAl | 0.7 | 24 | 0.274
US50 | WAl | 0.7 | 25 | 0.275
US50 | WAl | 0.7 | 26 | 0.277
US50 | WAl | 0.7 | 27 | 0.279
US50 | WAl | 0.7 | 28 | 0.281
US50 | WAl | 0.7 | 29 | 0.283
US50 | WAl | 0.7 | 30 | 0.285
US50 | WAl | 0.7 | 31 | 0.286
US50 | WAl | 0.7 | 32 | 0.288
US50 | WAl | 0.7 | 33 | 0.289
US50 | WAl | 0.7 | 34 | 0.291
US50 | WAl | 0.7 | 35 | 0.292
US50 | WAl | 0.7 | 36 | 0.293
US50 | WAl | 0.7 | 37 | 0.295
US50 | WAl | 0.7 | 38 | 0.296
US50 | WAl | 0.7 | 39 | 0.297
US50 | WAl | 0.7 | 40 | 0.298
US50 | WAl | 0.7 | 41 | 0.299
US50 | WAl | 0.7 | 42 | 0.3
US50 | WAl | 0.7 | 43 | 0.301
US50 | WAl | 0.7 | 44 | 0.302
US50 | WAl | 0.7 | 45 | 0.303
US50 | WAl | 0.7 | 46 | 0.304
US50 | WAl | 0.7 | 47 | 0.305
US50 | WAl | 0.7 | 48 | 0.305
US50 | WAl | 0.7 | 49 | 0.306
US50 | WAl | 0.75 | 24 | 0.291
US50 | WAl | 0.75 | 25 | 0.292
US50 | WAl | 0.75 | 26 | 0.294
US50 | WAl | 0.75 | 27 | 0.296
US50 | WAl | 0.75 | 28 | 0.298
US50 | WAl | 0.75 | 29 | 0.299
US50 | WAl | 0.75 | 30 | 0.301
US50 | WAl | 0.75 | 31 | 0.302
US50 | WAl | 0.75 | 32 | 0.304
US50 | WAl | 0.75 | 33 | 0.305
US50 | WAl | 0.75 | 34 | 0.306
US50 | WAl | 0.75 | 35 | 0.307
US50 | WAl | 0.75 | 36 | 0.309
US50 | WAl | 0.75 | 37 | 0.31
US50 | WAl | 0.75 | 38 | 0.311
US50 | WAl | 0.75 | 39 | 0.312
US50 | WAl | 0.75 | 40 | 0.313
US50 | WAl | 0.75 | 41 | 0.314
US50 | WAl | 0.75 | 42 | 0.315
US50 | WAl | 0.75 | 43 | 0.316
US50 | WAl | 0.75 | 44 | 0.316
US50 | WAl | 0.75 | 45 | 0.317
US50 | WAl | 0.75 | 46 | 0.318
US50 | WAl | 0.75 | 47 | 0.319
US50 | WAl | 0.75 | 48 | 0.319
US50 | WAl | 0.75 | 49 | 0.32
US50 | WAl | 0.8 | 25 | 0.309
US50 | WAl | 0.8 | 26 | 0.31
US50 | WAl | 0.8 | 27 | 0.312
US50 | WAl | 0.8 | 28 | 0.313
US50 | WAl | 0.8 | 29 | 0.315
US50 | WAl | 0.8 | 30 | 0.316
US50 | WAl | 0.8 | 31 | 0.317
US50 | WAl | 0.8 | 32 | 0.319
US50 | WAl | 0.8 | 33 | 0.32
US50 | WAl | 0.8 | 34 | 0.321
US50 | WAl | 0.8 | 35 | 0.322
US50 | WAl | 0.8 | 36 | 0.323
US50 | WAl | 0.8 | 37 | 0.324
US50 | WAl | 0.8 | 38 | 0.325
US50 | WAl | 0.8 | 39 | 0.326
US50 | WAl | 0.8 | 40 | 0.327
US50 | WAl | 0.8 | 41 | 0.328
US50 | WAl | 0.8 | 42 | 0.329
US50 | WAl | 0.8 | 43 | 0.33
US50 | WAl | 0.8 | 44 | 0.33
US50 | WAl | 0.8 | 45 | 0.331
US50 | WAl | 0.8 | 46 | 0.332
US50 | WAl | 0.8 | 47 | 0.332
US50 | WAl | 0.8 | 48 | 0.333
US50 | WAl | 0.8 | 49 | 0.334
US50 | WAl | 0.85 | 26 | 0.326
US50 | WAl | 0.85 | 27 | 0.327
US50 | WAl | 0.85 | 28 | 0.328
US50 | WAl | 0.85 | 29 | 0.33
US50 | WAl | 0.85 | 30 | 0.331
US50 | WAl | 0.85 | 31 | 0.332
US50 | WAl | 0.85 | 32 | 0.333
US50 | WAl | 0.85 | 33 | 0.334
US50 | WAl | 0.85 | 34 | 0.335
US50 | WAl | 0.85 | 35 | 0.336
US50 | WAl | 0.85 | 36 | 0.337
US50 | WAl | 0.85 | 37 | 0.338
US50 | WAl | 0.85 | 38 | 0.339
US50 | WAl | 0.85 | 39 | 0.34
US50 | WAl | 0.85 | 40 | 0.341
US50 | WAl | 0.85 | 41 | 0.342
US50 | WAl | 0.85 | 42 | 0.342
US50 | WAl | 0.85 | 43 | 0.343
US50 | WAl | 0.85 | 44 | 0.344
US50 | WAl | 0.85 | 45 | 0.344
US50 | WAl | 0.85 | 46 | 0.345
US50 | WAl | 0.85 | 47 | 0.346
US50 | WAl | 0.85 | 48 | 0.346
US50 | WAl | 0.85 | 49 | 0.347
US50 | WAl | 0.9 | 27 | 0.342
US50 | WAl | 0.9 | 28 | 0.343
US50 | WAl | 0.9 | 29 | 0.344
US50 | WAl | 0.9 | 30 | 0.345
US50 | WAl | 0.9 | 31 | 0.346
US50 | WAl | 0.9 | 32 | 0.347
US50 | WAl | 0.9 | 33 | 0.348
US50 | WAl | 0.9 | 34 | 0.349
US50 | WAl | 0.9 | 35 | 0.35
US50 | WAl | 0.9 | 36 | 0.351
US50 | WAl | 0.9 | 37 | 0.352
US50 | WAl | 0.9 | 38 | 0.353
US50 | WAl | 0.9 | 39 | 0.353
US50 | WAl | 0.9 | 40 | 0.354
US50 | WAl | 0.9 | 41 | 0.355
US50 | WAl | 0.9 | 42 | 0.355
US50 | WAl | 0.9 | 43 | 0.356
US50 | WAl | 0.9 | 44 | 0.357
US50 | WAl | 0.9 | 45 | 0.357
US50 | WAl | 0.9 | 46 | 0.358
US50 | WAl | 0.9 | 47 | 0.358
US50 | WAl | 0.9 | 48 | 0.359
US50 | WAl | 0.9 | 49 | 0.359
US50 | WAl | 0.95 | 27 | 0.356
US50 | WAl | 0.95 | 28 | 0.357
US50 | WAl | 0.95 | 29 | 0.358
US50 | WAl | 0.95 | 30 | 0.359
US50 | WAl | 0.95 | 31 | 0.36
US50 | WAl | 0.95 | 32 | 0.361
US50 | WAl | 0.95 | 33 | 0.361
US50 | WAl | 0.95 | 34 | 0.362
US50 | WAl | 0.95 | 35 | 0.363
US50 | WAl | 0.95 | 36 | 0.364
US50 | WAl | 0.95 | 37 | 0.365
US50 | WAl | 0.95 | 38 | 0.365
US50 | WAl | 0.95 | 39 | 0.366
US50 | WAl | 0.95 | 40 | 0.367
US50 | WAl | 0.95 | 41 | 0.367
US50 | WAl | 0.95 | 42 | 0.368
US50 | WAl | 0.95 | 43 | 0.368
US50 | WAl | 0.95 | 44 | 0.369
US50 | WAl | 0.95 | 45 | 0.369
US50 | WAl | 0.95 | 46 | 0.37
US50 | WAl | 0.95 | 47 | 0.37
US50 | WAl | 0.95 | 48 | 0.371
US50 | WAl | 0.95 | 49 | 0.371
US50 | WAl | 1 | 28 | 0.37
US50 | WAl | 1 | 29 | 0.371
US50 | WAl | 1 | 30 | 0.372
US50 | WAl | 1 | 31 | 0.372
US50 | WAl | 1 | 32 | 0.373
US50 | WAl | 1 | 33 | 0.374
US50 | WAl | 1 | 34 | 0.375
US50 | WAl | 1 | 35 | 0.376
US50 | WAl | 1 | 36 | 0.376
US50 | WAl | 1 | 37 | 0.377
US50 | WAl | 1 | 38 | 0.378
US50 | WAl | 1 | 39 | 0.378
US50 | WAl | 1 | 40 | 0.379
US50 | WAl | 1 | 41 | 0.379
US50 | WAl | 1 | 42 | 0.38
US50 | WAl | 1 | 43 | 0.38
US50 | WAl | 1 | 44 | 0.381
US50 | WAl | 1 | 45 | 0.381
US50 | WAl | 1 | 46 | 0.382
US50 | WAl | 1 | 47 | 0.382
US50 | WAl | 1 | 48 | 0.382
US50 | WAl | 1 | 49 | 0.383
US50 | WAl | 1.05 | 29 | 0.383
US50 | WAl | 1.05 | 30 | 0.384
US50 | WAl | 1.05 | 31 | 0.385
US50 | WAl | 1.05 | 32 | 0.386
US50 | WAl | 1.05 | 33 | 0.386
US50 | WAl | 1.05 | 34 | 0.387
US50 | WAl | 1.05 | 35 | 0.388
US50 | WAl | 1.05 | 36 | 0.388
US50 | WAl | 1.05 | 37 | 0.389
US50 | WAl | 1.05 | 38 | 0.389
US50 | WAl | 1.05 | 39 | 0.39
US50 | WAl | 1.05 | 40 | 0.39
US50 | WAl | 1.05 | 41 | 0.391
US50 | WAl | 1.05 | 42 | 0.391
US50 | WAl | 1.05 | 43 | 0.392
US50 | WAl | 1.05 | 44 | 0.392
US50 | WAl | 1.05 | 45 | 0.393
US50 | WAl | 1.05 | 46 | 0.393
US50 | WAl | 1.05 | 47 | 0.393
US50 | WAl | 1.05 | 48 | 0.394
US50 | WAl | 1.05 | 49 | 0.394
US50 | WAl | 1.1 | 29 | 0.395
US50 | WAl | 1.1 | 30 | 0.396
US50 | WAl | 1.1 | 31 | 0.397
US50 | WAl | 1.1 | 32 | 0.397
US50 | WAl | 1.1 | 33 | 0.398
US50 | WAl | 1.1 | 34 | 0.399
US50 | WAl | 1.1 | 35 | 0.399
US50 | WAl | 1.1 | 36 | 0.4
US50 | WAl | 1.1 | 37 | 0.4
US50 | WAl | 1.1 | 38 | 0.401
US50 | WAl | 1.1 | 39 | 0.401
US50 | WAl | 1.1 | 40 | 0.402
US50 | WAl | 1.1 | 41 | 0.402
US50 | WAl | 1.1 | 42 | 0.402
US50 | WAl | 1.1 | 43 | 0.403
US50 | WAl | 1.1 | 44 | 0.403
US50 | WAl | 1.1 | 45 | 0.404
US50 | WAl | 1.1 | 46 | 0.404
US50 | WAl | 1.1 | 47 | 0.404
US50 | WAl | 1.1 | 48 | 0.404
US50 | WAl | 1.1 | 49 | 0.405
US50 | WAl | 1.15 | 30 | 0.408
US50 | WAl | 1.15 | 31 | 0.408
US50 | WAl | 1.15 | 32 | 0.409
US50 | WAl | 1.15 | 33 | 0.409
US50 | WAl | 1.15 | 34 | 0.41
US50 | WAl | 1.15 | 35 | 0.41
US50 | WAl | 1.15 | 36 | 0.411
US50 | WAl | 1.15 | 37 | 0.411
US50 | WAl | 1.15 | 38 | 0.412
US50 | WAl | 1.15 | 39 | 0.412
US50 | WAl | 1.15 | 40 | 0.412
US50 | WAl | 1.15 | 41 | 0.413
US50 | WAl | 1.15 | 42 | 0.413
US50 | WAl | 1.15 | 43 | 0.413
US50 | WAl | 1.15 | 44 | 0.414
US50 | WAl | 1.15 | 45 | 0.414
US50 | WAl | 1.15 | 46 | 0.414
US50 | WAl | 1.15 | 47 | 0.415
US50 | WAl | 1.15 | 48 | 0.415
US50 | WAl | 1.15 | 49 | 0.415
US50 | WAl | 1.2 | 31 | 0.419
US50 | WAl | 1.2 | 32 | 0.42
US50 | WAl | 1.2 | 33 | 0.42
US50 | WAl | 1.2 | 34 | 0.421
US50 | WAl | 1.2 | 35 | 0.421
US50 | WAl | 1.2 | 36 | 0.421
US50 | WAl | 1.2 | 37 | 0.422
US50 | WAl | 1.2 | 38 | 0.422
US50 | WAl | 1.2 | 39 | 0.422
US50 | WAl | 1.2 | 40 | 0.423
US50 | WAl | 1.2 | 41 | 0.423
US50 | WAl | 1.2 | 42 | 0.423
US50 | WAl | 1.2 | 43 | 0.424
US50 | WAl | 1.2 | 44 | 0.424
US50 | WAl | 1.2 | 45 | 0.424
US50 | WAl | 1.2 | 46 | 0.424
US50 | WAl | 1.2 | 47 | 0.425
US50 | WAl | 1.2 | 48 | 0.425
US50 | WAl | 1.2 | 49 | 0.425
US50 | WAl | 1.25 | 36 | 0.432
US50 | WAl | 1.25 | 37 | 0.432
US50 | WAl | 1.25 | 38 | 0.432
US50 | WAl | 1.25 | 39 | 0.432
US50 | WAl | 1.25 | 40 | 0.433
US50 | WAl | 1.25 | 41 | 0.433
US50 | WAl | 1.25 | 42 | 0.433
US50 | WAl | 1.25 | 43 | 0.433
US50 | WAl | 1.25 | 44 | 0.434
US50 | WAl | 1.25 | 45 | 0.434
US50 | WAl | 1.25 | 46 | 0.434
US50 | WAl | 1.25 | 47 | 0.434
US50 | WAl | 1.25 | 48 | 0.434
US50 | WAl | 1.25 | 49 | 0.435
US50 | WAl | 1.3 | 41 | 0.442
US50 | WAl | 1.3 | 42 | 0.443
US50 | WAl | 1.3 | 43 | 0.443
US50 | WAl | 1.3 | 44 | 0.443
US50 | WAl | 1.3 | 45 | 0.443
US50 | WAl | 1.3 | 46 | 0.443
US50 | WAl | 1.3 | 47 | 0.444
US50 | WAl | 1.3 | 48 | 0.444
US50 | WAl | 1.3 | 49 | 0.444
US50 | WAl | 1.35 | 45 | 0.452
US50 | WAl | 1.35 | 46 | 0.452
US50 | WAl | 1.35 | 47 | 0.453
US50 | WAl | 1.35 | 48 | 0.453
US50 | WAl | 1.35 | 49 | 0.453
US50 | WRh | 0.2 | 20 | 0.057
US50 | WRh | 0.2 | 21 | 0.061
US50 | WRh | 0.225 | 20 | 0.068
US50 | WRh | 0.225 | 21 | 0.072
US50 | WRh | 0.225 | 22 | 0.076
US50 | WRh | 0.225 | 23 | 0.08
US50 | WRh | 0.225 | 24 | 0.083
US50 | WRh | 0.25 | 20 | 0.078
US50 | WRh | 0.25 | 21 | 0.083
US50 | WRh | 0.25 | 22 | 0.087
US50 | WRh | 0.25 | 23 | 0.091
US50 | WRh | 0.25 | 24 | 0.095
US50 | WRh | 0.25 | 25 | 0.098
US50 | WRh | 0.25 | 26 | 0.1
US50 | WRh | 0.275 | 20 | 0.088
US50 | WRh | 0.275 | 21 | 0.093
US50 | WRh | 0.275 | 22 | 0.097
US50 | WRh | 0.275 | 23 | 0.102
US50 | WRh | 0.275 | 24 | 0.106
US50 | WRh | 0.275 | 25 | 0.109
US50 | WRh | 0.275 | 26 | 0.112
US50 | WRh | 0.275 | 27 | 0.114
US50 | WRh | 0.275 | 28 | 0.116
US50 | WRh | 0.275 | 29 | 0.118
US50 | WRh | 0.275 | 30 | 0.119
US50 | WRh | 0.275 | 31 | 0.121
US50 | WRh | 0.275 | 32 | 0.123
US50 | WRh | 0.3 | 20 | 0.099
US50 | WRh | 0.3 | 21 | 0.103
US50 | WRh | 0.3 | 22 | 0.108
US50 | WRh | 0.3 | 23 | 0.112
US50 | WRh | 0.3 | 24 | 0.116
US50 | WRh | 0.3 | 25 | 0.12
US50 | WRh | 0.3 | 26 | 0.122
US50 | WRh | 0.3 | 27 | 0.124
US50 | WRh | 0.3 | 28 | 0.126
US50 | WRh | 0.3 | 29 | 0.128
US50 | WRh | 0.3 | 30 | 0.13
US50 | WRh | 0.3 | 31 | 0.131
US50 | WRh | 0.3 | 32 | 0.133
US50 | WRh | 0.3 | 33 | 0.135
US50 | WRh | 0.3 | 34 | 0.137
US50 | WRh | 0.3 | 35 | 0.139
US50 | WRh | 0.3 | 36 | 0.14
US50 | WRh | 0.325 | 20 | 0.109
US50 | WRh | 0.325 | 21 | 0.114
US50 | WRh | 0.325 | 22 | 0.118
US50 | WRh | 0.325 | 23 | 0.123
US50 | WRh | 0.325 | 24 | 0.127
US50 | WRh | 0.325 | 25 | 0.13
US50 | WRh | 0.325 | 26 | 0.133
US50 | WRh | 0.325 | 27 | 0.135
US50 | WRh | 0.325 | 28 | 0.136
US50 | WRh | 0.325 | 29 | 0.138
US50 | WRh | 0.325 | 30 | 0.14
US50 | WRh | 0.325 | 31 | 0.141
US50 | WRh | 0.325 | 32 | 0.143
US50 | WRh | 0.325 | 33 | 0.145
US50 | WRh | 0.325 | 34 | 0.146
US50 | WRh | 0.325 | 35 | 0.148
US50 | WRh | 0.325 | 36 | 0.15
US50 | WRh | 0.325 | 37 | 0.152
US50 | WRh | 0.325 | 38 | 0.153
US50 | WRh | 0.325 | 39 | 0.155
US50 | WRh | 0.325 | 40 | 0.157
US50 | WRh | 0.325 | 41 | 0.158
US50 | WRh | 0.325 | 42 | 0.16
US50 | WRh | 0.35 | 20 | 0.12
US50 | WRh | 0.35 | 21 | 0.125
US50 | WRh | 0.35 | 22 | 0.129
US50 | WRh | 0.35 | 23 | 0.133
US50 | WRh | 0.35 | 24 | 0.137
US50 | WRh | 0.35 | 25 | 0.141
US50 | WRh | 0.35 | 26 | 0.143
US50 | WRh | 0.35 | 27 | 0.145
US50 | WRh | 0.35 | 28 | 0.146
US50 | WRh | 0.35 | 29 | 0.148
US50 | WRh | 0.35 | 30 | 0.149
US50 | WRh | 0.35 | 31 | 0.151
US50 | WRh | 0.35 | 32 | 0.153
US50 | WRh | 0.35 | 33 | 0.154
US50 | WRh | 0.35 | 34 | 0.156
US50 | WRh | 0.35 | 35 | 0.157
US50 | WRh | 0.35 | 36 | 0.159
US50 | WRh | 0.35 | 37 | 0.161
US50 | WRh | 0.35 | 38 | 0.162
US50 | WRh | 0.35 | 39 | 0.164
US50 | WRh | 0.35 | 40 | 0.166
US50 | WRh | 0.35 | 41 | 0.167
US50 | WRh | 0.35 | 42 | 0.169
US50 | WRh | 0.35 | 43 | 0.171
US50 | WRh | 0.35 | 44 | 0.172
US50 | WRh | 0.35 | 45 | 0.174
US50 | WRh | 0.35 | 46 | 0.175
US50 | WRh | 0.35 | 47 | 0.177
US50 | WRh | 0.375 | 20 | 0.131
US50 | WRh | 0.375 | 21 | 0.135
US50 | WRh | 0.375 | 22 | 0.139
US50 | WRh | 0.375 | 23 | 0.144
US50 | WRh | 0.375 | 24 | 0.148
US50 | WRh | 0.375 | 25 | 0.151
US50 | WRh | 0.375 | 26 | 0.153
US50 | WRh | 0.375 | 27 | 0.155
US50 | WRh | 0.375 | 28 | 0.156
US50 | WRh | 0.375 | 29 | 0.158
US50 | WRh | 0.375 | 30 | 0.159
US50 | WRh | 0.375 | 31 | 0.16
US50 | WRh | 0.375 | 32 | 0.162
US50 | WRh | 0.375 | 33 | 0.163
US50 | WRh | 0.375 | 34 | 0.165
US50 | WRh | 0.375 | 35 | 0.166
US50 | WRh | 0.375 | 36 | 0.168
US50 | WRh | 0.375 | 37 | 0.17
US50 | WRh | 0.375 | 38 | 0.171
US50 | WRh | 0.375 | 39 | 0.173
US50 | WRh | 0.375 | 40 | 0.175
US50 | WRh | 0.375 | 41 | 0.176
US50 | WRh | 0.375 | 42 | 0.178
US50 | WRh | 0.375 | 43 | 0.179
US50 | WRh | 0.375 | 44 | 0.181
US50 | WRh | 0.375 | 45 | 0.182
US50 | WRh | 0.375 | 46 | 0.184
US50 | WRh | 0.375 | 47 | 0.186
US50 | WRh | 0.375 | 48 | 0.187
US50 | WRh | 0.375 | 49 | 0.189
US50 | WRh | 0.4 | 20 | 0.141
US50 | WRh | 0.4 | 21 | 0.146
US50 | WRh | 0.4 | 22 | 0.15
US50 | WRh | 0.4 | 23 | 0.154
US50 | WRh | 0.4 | 24 | 0.158
US50 | WRh | 0.4 | 25 | 0.161
US50 | WRh | 0.4 | 26 | 0.163
US50 | WRh | 0.4 | 27 | 0.165
US50 | WRh | 0.4 | 28 | 0.166
US50 | WRh | 0.4 | 29 | 0.167
US50 | WRh | 0.4 | 30 | 0.169
US50 | WRh | 0.4 | 31 | 0.17
US50 | WRh | 0.4 | 32 | 0.171
US50 | WRh | 0.4 | 33 | 0.173
US50 | WRh | 0.4 | 34 | 0.174
US50 | WRh | 0.4 | 35 | 0.175
US50 | WRh | 0.4 | 36 | 0.177
US50 | WRh | 0.4 | 37 | 0.178
US50 | WRh | 0.4 | 38 | 0.18
US50 | WRh | 0.4 | 39 | 0.181
US50 | WRh | 0.4 | 40 | 0.183
US50 | WRh | 0.4 | 41 | 0.185
US50 | WRh | 0.4 | 42 | 0.186
US50 | WRh | 0.4 | 43 | 0.188
US50 | WRh | 0.4 | 44 | 0.189
US50 | WRh | 0.4 | 45 | 0.191
US50 | WRh | 0.4 | 46 | 0.192
US50 | WRh | 0.4 | 47 | 0.194
US50 | WRh | 0.4 | 48 | 0.196
US50 | WRh | 0.4 | 49 | 0.197
US50 | WRh | 0.425 | 20 | 0.152
US50 | WRh | 0.425 | 21 | 0.156
US50 | WRh | 0.425 | 22 | 0.161
US50 | WRh | 0.425 | 23 | 0.165
US50 | WRh | 0.425 | 24 | 0.168
US50 | WRh | 0.425 | 25 | 0.171
US50 | WRh | 0.425 | 26 | 0.173
US50 | WRh | 0.425 | 27 | 0.175
US50 | WRh | 0.425 | 28 | 0.176
US50 | WRh | 0.425 | 29 | 0.177
US50 | WRh | 0.425 | 30 | 0.178
US50 | WRh | 0.425 | 31 | 0.179
US50 | WRh | 0.425 | 32 | 0.18
US50 | WRh | 0.425 | 33 | 0.182
US50 | WRh | 0.425 | 34 | 0.183
US50 | WRh | 0.425 | 35 | 0.184
US50 | WRh | 0.425 | 36 | 0.186
US50 | WRh | 0.425 | 37 | 0.187
US50 | WRh | 0.425 | 38 | 0.189
US50 | WRh | 0.425 | 39 | 0.19
US50 | WRh | 0.425 | 40 | 0.192
US50 | WRh | 0.425 | 41 | 0.193
US50 | WRh | 0.425 | 42 | 0.195
US50 | WRh | 0.425 | 43 | 0.196
US50 | WRh | 0.425 | 44 | 0.198
US50 | WRh | 0.425 | 45 | 0.199
US50 | WRh | 0.425 | 46 | 0.201
US50 | WRh | 0.425 | 47 | 0.202
US50 | WRh | 0.425 | 48 | 0.204
US50 | WRh | 0.425 | 49 | 0.205
US50 | WRh | 0.45 | 20 | 0.163
US50 | WRh | 0.45 | 21 | 0.167
US50 | WRh | 0.45 | 22 | 0.171
US50 | WRh | 0.45 | 23 | 0.175
US50 | WRh | 0.45 | 24 | 0.179
US50 | WRh | 0.45 | 25 | 0.181
US50 | WRh | 0.45 | 26 | 0.183
US50 | WRh | 0.45 | 27 | 0.184
US50 | WRh | 0.45 | 28 | 0.186
US50 | WRh | 0.45 | 29 | 0.187
US50 | WRh | 0.45 | 30 | 0.187
US50 | WRh | 0.45 | 31 | 0.188
US50 | WRh | 0.45 | 32 | 0.189
US50 | WRh | 0.45 | 33 | 0.19
US50 | WRh | 0.45 | 34 | 0.192
US50 | WRh | 0.45 | 35 | 0.193
US50 | WRh | 0.45 | 36 | 0.194
US50 | WRh | 0.45 | 37 | 0.196
US50 | WRh | 0.45 | 38 | 0.197
US50 | WRh | 0.45 | 39 | 0.198
US50 | WRh | 0.45 | 40 | 0.2
US50 | WRh | 0.45 | 41 | 0.201
US50 | WRh | 0.45 | 42 | 0.203
US50 | WRh | 0.45 | 43 | 0.204
US50 | WRh | 0.45 | 44 | 0.206
US50 | WRh | 0.45 | 45 | 0.207
US50 | WRh | 0.45 | 46 | 0.209
US50 | WRh | 0.45 | 47 | 0.21
US50 | WRh | 0.45 | 48 | 0.212
US50 | WRh | 0.45 | 49 | 0.213
US50 | WRh | 0.475 | 20 | 0.174
US50 | WRh | 0.475 | 21 | 0.178
US50 | WRh | 0.475 | 22 | 0.181
US50 | WRh | 0.475 | 23 | 0.185
US50 | WRh | 0.475 | 24 | 0.189
US50 | WRh | 0.475 | 25 | 0.191
US50 | WRh | 0.475 | 26 | 0.193
US50 | WRh | 0.475 | 27 | 0.194
US50 | WRh | 0.475 | 28 | 0.195
US50 | WRh | 0.475 | 29 | 0.196
US50 | WRh | 0.475 | 30 | 0.197
US50 | WRh | 0.475 | 31 | 0.198
US50 | WRh | 0.475 | 32 | 0.198
US50 | WRh | 0.475 | 33 | 0.199
US50 | WRh | 0.475 | 34 | 0.201
US50 | WRh | 0.475 | 35 | 0.202
US50 | WRh | 0.475 | 36 | 0.203
US50 | WRh | 0.475 | 37 | 0.204
US50 | WRh | 0.475 | 38 | 0.205
US50 | WRh | 0.475 | 39 | 0.207
US50 | WRh | 0.475 | 40 | 0.208
US50 | WRh | 0.475 | 41 | 0.21
US50 | WRh | 0.475 | 42 | 0.211
US50 | WRh | 0.475 | 43 | 0.212
US50 | WRh | 0.475 | 44 | 0.214
US50 | WRh | 0.475 | 45 | 0.215
US50 | WRh | 0.475 | 46 | 0.217
US50 | WRh | 0.475 | 47 | 0.218
US50 | WRh | 0.475 | 48 | 0.22
US50 | WRh | 0.475 | 49 | 0.221
US50 | WRh | 0.5 | 20 | 0.185
US50 | WRh | 0.5 | 21 | 0.188
US50 | WRh | 0.5 | 22 | 0.192
US50 | WRh | 0.5 | 23 | 0.195
US50 | WRh | 0.5 | 24 | 0.199
US50 | WRh | 0.5 | 25 | 0.201
US50 | WRh | 0.5 | 26 | 0.203
US50 | WRh | 0.5 | 27 | 0.204
US50 | WRh | 0.5 | 28 | 0.205
US50 | WRh | 0.5 | 29 | 0.205
US50 | WRh | 0.5 | 30 | 0.206
US50 | WRh | 0.5 | 31 | 0.207
US50 | WRh | 0.5 | 32 | 0.207
US50 | WRh | 0.5 | 33 | 0.208
US50 | WRh | 0.5 | 34 | 0.209
US50 | WRh | 0.5 | 35 | 0.21
US50 | WRh | 0.5 | 36 | 0.211
US50 | WRh | 0.5 | 37 | 0.213
US50 | WRh | 0.5 | 38 | 0.214
US50 | WRh | 0.5 | 39 | 0.215
US50 | WRh | 0.5 | 40 | 0.216
US50 | WRh | 0.5 | 41 | 0.218
US50 | WRh | 0.5 | 42 | 0.219
US50 | WRh | 0.5 | 43 | 0.22
US50 | WRh | 0.5 | 44 | 0.222
US50 | WRh | 0.5 | 45 | 0.223
US50 | WRh | 0.5 | 46 | 0.225
US50 | WRh | 0.5 | 47 | 0.226
US50 | WRh | 0.5 | 48 | 0.228
US50 | WRh | 0.5 | 49 | 0.229
US50 | WRh | 0.525 | 21 | 0.199
US50 | WRh | 0.525 | 22 | 0.202
US50 | WRh | 0.525 | 23 | 0.205
US50 | WRh | 0.525 | 24 | 0.209
US50 | WRh | 0.525 | 25 | 0.211
US50 | WRh | 0.525 | 26 | 0.212
US50 | WRh | 0.525 | 27 | 0.213
US50 | WRh | 0.525 | 28 | 0.214
US50 | WRh | 0.525 | 29 | 0.215
US50 | WRh | 0.525 | 30 | 0.215
US50 | WRh | 0.525 | 31 | 0.216
US50 | WRh | 0.525 | 32 | 0.216
US50 | WRh | 0.525 | 33 | 0.217
US50 | WRh | 0.525 | 34 | 0.218
US50 | WRh | 0.525 | 35 | 0.219
US50 | WRh | 0.525 | 36 | 0.22
US50 | WRh | 0.525 | 37 | 0.221
US50 | WRh | 0.525 | 38 | 0.222
US50 | WRh | 0.525 | 39 | 0.223
US50 | WRh | 0.525 | 40 | 0.224
US50 | WRh | 0.525 | 41 | 0.226
US50 | WRh | 0.525 | 42 | 0.227
US50 | WRh | 0.525 | 43 | 0.228
US50 | WRh | 0.525 | 44 | 0.23
US50 | WRh | 0.525 | 45 | 0.231
US50 | WRh | 0.525 | 46 | 0.233
US50 | WRh | 0.525 | 47 | 0.234
US50 | WRh | 0.525 | 48 | 0.235
US50 | WRh | 0.525 | 49 | 0.237
US50 | WRh | 0.55 | 21 | 0.209
US50 | WRh | 0.55 | 22 | 0.212
US50 | WRh | 0.55 | 23 | 0.215
US50 | WRh | 0.55 | 24 | 0.218
US50 | WRh | 0.55 | 25 | 0.221
US50 | WRh | 0.55 | 26 | 0.222
US50 | WRh | 0.55 | 27 | 0.223
US50 | WRh | 0.55 | 28 | 0.223
US50 | WRh | 0.55 | 29 | 0.224
US50 | WRh | 0.55 | 30 | 0.224
US50 | WRh | 0.55 | 31 | 0.225
US50 | WRh | 0.55 | 32 | 0.226
US50 | WRh | 0.55 | 33 | 0.226
US50 | WRh | 0.55 | 34 | 0.227
US50 | WRh | 0.55 | 35 | 0.227
US50 | WRh | 0.55 | 36 | 0.228
US50 | WRh | 0.55 | 37 | 0.229
US50 | WRh | 0.55 | 38 | 0.23
US50 | WRh | 0.55 | 39 | 0.232
US50 | WRh | 0.55 | 40 | 0.233
US50 | WRh | 0.55 | 41 | 0.234
US50 | WRh | 0.55 | 42 | 0.235
US50 | WRh | 0.55 | 43 | 0.236
US50 | WRh | 0.55 | 44 | 0.238
US50 | WRh | 0.55 | 45 | 0.239
US50 | WRh | 0.55 | 46 | 0.24
US50 | WRh | 0.55 | 47 | 0.242
US50 | WRh | 0.55 | 48 | 0.243
US50 | WRh | 0.55 | 49 | 0.245
US50 | WRh | 0.575 | 22 | 0.222
US50 | WRh | 0.575 | 23 | 0.225
US50 | WRh | 0.575 | 24 | 0.228
US50 | WRh | 0.575 | 25 | 0.23
US50 | WRh | 0.575 | 26 | 0.231
US50 | WRh | 0.575 | 27 | 0.232
US50 | WRh | 0.575 | 28 | 0.233
US50 | WRh | 0.575 | 29 | 0.233
US50 | WRh | 0.575 | 30 | 0.233
US50 | WRh | 0.575 | 31 | 0.234
US50 | WRh | 0.575 | 32 | 0.234
US50 | WRh | 0.575 | 33 | 0.235
US50 | WRh | 0.575 | 34 | 0.236
US50 | WRh | 0.575 | 35 | 0.236
US50 | WRh | 0.575 | 36 | 0.237
US50 | WRh | 0.575 | 37 | 0.238
US50 | WRh | 0.575 | 38 | 0.239
US50 | WRh | 0.575 | 39 | 0.24
US50 | WRh | 0.575 | 40 | 0.241
US50 | WRh | 0.575 | 41 | 0.242
US50 | WRh | 0.575 | 42 | 0.243
US50 | WRh | 0.575 | 43 | 0.244
US50 | WRh | 0.575 | 44 | 0.246
US50 | WRh | 0.575 | 45 | 0.247
US50 | WRh | 0.575 | 46 | 0.248
US50 | WRh | 0.575 | 47 | 0.249
US50 | WRh | 0.575 | 48 | 0.251
US50 | WRh | 0.575 | 49 | 0.252
US50 | WRh | 0.6 | 22 | 0.232
US50 | WRh | 0.6 | 23 | 0.235
US50 | WRh | 0.6 | 24 | 0.237
US50 | WRh | 0.6 | 25 | 0.239
US50 | WRh | 0.6 | 26 | 0.241
US50 | WRh | 0.6 | 27 | 0.241
US50 | WRh | 0.6 | 28 | 0.242
US50 | WRh | 0.6 | 29 | 0.242
US50 | WRh | 0.6 | 30 | 0.242
US50 | WRh | 0.6 | 31 | 0.243
US50 | WRh | 0.6 | 32 | 0.243
US50 | WRh | 0.6 | 33 | 0.244
US50 | WRh | 0.6 | 34 | 0.244
US50 | WRh | 0.6 | 35 | 0.245
US50 | WRh | 0.6 | 36 | 0.245
US50 | WRh | 0.6 | 37 | 0.246
US50 | WRh | 0.6 | 38 | 0.247
US50 | WRh | 0.6 | 39 | 0.248
US50 | WRh | 0.6 | 40 | 0.249
US50 | WRh | 0.6 | 41 | 0.25
US50 | WRh | 0.6 | 42 | 0.251
US50 | WRh | 0.6 | 43 | 0.252
US50 | WRh | 0.6 | 44 | 0.253
US50 | WRh | 0.6 | 45 | 0.255
US50 | WRh | 0.6 | 46 | 0.256
US50 | WRh | 0.6 | 47 | 0.257
US50 | WRh | 0.6 | 48 | 0.258
US50 | WRh | 0.6 | 49 | 0.26
US50 | WRh | 0.625 | 22 | 0.242
US50 | WRh | 0.625 | 23 | 0.244
US50 | WRh | 0.625 | 24 | 0.247
US50 | WRh | 0.625 | 25 | 0.249
US50 | WRh | 0.625 | 26 | 0.25
US50 | WRh | 0.625 | 27 | 0.25
US50 | WRh | 0.625 | 28 | 0.251
US50 | WRh | 0.625 | 29 | 0.251
US50 | WRh | 0.625 | 30 | 0.251
US50 | WRh | 0.625 | 31 | 0.251
US50 | WRh | 0.625 | 32 | 0.252
US50 | WRh | 0.625 | 33 | 0.252
US50 | WRh | 0.625 | 34 | 0.253
US50 | WRh | 0.625 | 35 | 0.253
US50 | WRh | 0.625 | 36 | 0.254
US50 | WRh | 0.625 | 37 | 0.254
US50 | WRh | 0.625 | 38 | 0.255
US50 | WRh | 0.625 | 39 | 0.256
US50 | WRh | 0.625 | 40 | 0.257
US50 | WRh | 0.625 | 41 | 0.258
US50 | WRh | 0.625 | 42 | 0.259
US50 | WRh | 0.625 | 43 | 0.26
US50 | WRh | 0.625 | 44 | 0.261
US50 | WRh | 0.625 | 45 | 0.262
US50 | WRh | 0.625 | 46 | 0.264
US50 | WRh | 0.625 | 47 | 0.265
US50 | WRh | 0.625 | 48 | 0.266
US50 | WRh | 0.625 | 49 | 0.267
US50 | WRh | 0.65 | 23 | 0.254
US50 | WRh | 0.65 | 24 | 0.256
US50 | WRh | 0.65 | 25 | 0.258
US50 | WRh | 0.65 | 26 | 0.259
US50 | WRh | 0.65 | 27 | 0.259
US50 | WRh | 0.65 | 28 | 0.259
US50 | WRh | 0.65 | 29 | 0.26
US50 | WRh | 0.65 | 30 | 0.26
US50 | WRh | 0.65 | 31 | 0.26
US50 | WRh | 0.65 | 32 | 0.26
US50 | WRh | 0.65 | 33 | 0.261
US50 | WRh | 0.65 | 34 | 0.261
US50 | WRh | 0.65 | 35 | 0.262
US50 | WRh | 0.65 | 36 | 0.262
US50 | WRh | 0.65 | 37 | 0.262
US50 | WRh | 0.65 | 38 | 0.263
US50 | WRh | 0.65 | 39 | 0.264
US50 | WRh | 0.65 | 40 | 0.265
US50 | WRh | 0.65 | 41 | 0.266
US50 | WRh | 0.65 | 42 | 0.267
US50 | WRh | 0.65 | 43 | 0.268
US50 | WRh | 0.65 | 44 | 0.269
US50 | WRh | 0.65 | 45 | 0.27
US50 | WRh | 0.65 | 46 | 0.271
US50 | WRh | 0.65 | 47 | 0.272
US50 | WRh | 0.65 | 48 | 0.274
US50 | WRh | 0.65 | 49 | 0.275
US50 | WRh | 0.675 | 25 | 0.267
US50 | WRh | 0.675 | 26 | 0.267
US50 | WRh | 0.675 | 27 | 0.268
US50 | WRh | 0.675 | 28 | 0.268
US50 | WRh | 0.675 | 29 | 0.268
US50 | WRh | 0.675 | 30 | 0.269
US50 | WRh | 0.675 | 31 | 0.269
US50 | WRh | 0.675 | 32 | 0.269
US50 | WRh | 0.675 | 33 | 0.269
US50 | WRh | 0.675 | 34 | 0.27
US50 | WRh | 0.675 | 35 | 0.27
US50 | WRh | 0.675 | 36 | 0.27
US50 | WRh | 0.675 | 37 | 0.271
US50 | WRh | 0.675 | 38 | 0.271
US50 | WRh | 0.675 | 39 | 0.272
US50 | WRh | 0.675 | 40 | 0.273
US50 | WRh | 0.675 | 41 | 0.274
US50 | WRh | 0.675 | 42 | 0.274
US50 | WRh | 0.675 | 43 | 0.275
US50 | WRh | 0.675 | 44 | 0.276
US50 | WRh | 0.675 | 45 | 0.278
US50 | WRh | 0.675 | 46 | 0.279
US50 | WRh | 0.675 | 47 | 0.28
US50 | WRh | 0.675 | 48 | 0.281
US50 | WRh | 0.675 | 49 | 0.282
US50 | WRh | 0.7 | 27 | 0.276
US50 | WRh | 0.7 | 28 | 0.277
US50 | WRh | 0.7 | 29 | 0.277
US50 | WRh | 0.7 | 30 | 0.277
US50 | WRh | 0.7 | 31 | 0.277
US50 | WRh | 0.7 | 32 | 0.277
US50 | WRh | 0.7 | 33 | 0.278
US50 | WRh | 0.7 | 34 | 0.278
US50 | WRh | 0.7 | 35 | 0.278
US50 | WRh | 0.7 | 36 | 0.278
US50 | WRh | 0.7 | 37 | 0.279
US50 | WRh | 0.7 | 38 | 0.279
US50 | WRh | 0.7 | 39 | 0.28
US50 | WRh | 0.7 | 40 | 0.28
US50 | WRh | 0.7 | 41 | 0.281
US50 | WRh | 0.7 | 42 | 0.282
US50 | WRh | 0.7 | 43 | 0.283
US50 | WRh | 0.7 | 44 | 0.284
US50 | WRh | 0.7 | 45 | 0.285
US50 | WRh | 0.7 | 46 | 0.286
US50 | WRh | 0.7 | 47 | 0.287
US50 | WRh | 0.7 | 48 | 0.289
US50 | WRh | 0.7 | 49 | 0.29
US50 | WRh | 0.725 | 33 | 0.286
US50 | WRh | 0.725 | 34 | 0.286
US50 | WRh | 0.725 | 35 | 0.286
US50 | WRh | 0.725 | 36 | 0.286
US50 | WRh | 0.725 | 37 | 0.287
US50 | WRh | 0.725 | 38 | 0.287
US50 | WRh | 0.725 | 39 | 0.288
US50 | WRh | 0.725 | 40 | 0.288
US50 | WRh | 0.725 | 41 | 0.289
US50 | WRh | 0.725 | 42 | 0.29
US50 | WRh | 0.725 | 43 | 0.291
US50 | WRh | 0.725 | 44 | 0.292
US50 | WRh | 0.725 | 45 | 0.293
US50 | WRh | 0.725 | 46 | 0.294
US50 | WRh | 0.725 | 47 | 0.295
US50 | WRh | 0.725 | 48 | 0.296
US50 | WRh | 0.725 | 49 | 0.297
US50 | WRh | 0.75 | 37 | 0.295
US50 | WRh | 0.75 | 38 | 0.295
US50 | WRh | 0.75 | 39 | 0.295
US50 | WRh | 0.75 | 40 | 0.296
US50 | WRh | 0.75 | 41 | 0.297
US50 | WRh | 0.75 | 42 | 0.297
US50 | WRh | 0.75 | 43 | 0.298
US50 | WRh | 0.75 | 44 | 0.299
US50 | WRh | 0.75 | 45 | 0.3
US50 | WRh | 0.75 | 46 | 0.301
US50 | WRh | 0.75 | 47 | 0.302
US50 | WRh | 0.75 | 48 | 0.303
US50 | WRh | 0.75 | 49 | 0.304
US50 | WRh | 0.775 | 43 | 0.305
US50 | WRh | 0.775 | 44 | 0.306
US50 | WRh | 0.775 | 45 | 0.307
US50 | WRh | 0.775 | 46 | 0.308
US50 | WRh | 0.775 | 47 | 0.309
US50 | WRh | 0.775 | 48 | 0.31
US50 | WRh | 0.775 | 49 | 0.312
US50 | WRh | 0.8 | 48 | 0.318
US50 | WRh | 0.8 | 49 | 0.319

## GAMMA_DBT_42mm_CC
densityMode | TargetFilter | ProjAngle | HVL | kVp | Gamma
FDA | MoMo | 09_25 | 0.2 | 20 | 0.046
FDA | MoMo | 09_25 | 0.2 | 21 | 0.049
FDA | MoMo | 09_25 | 0.2 | 22 | 0.051
FDA | MoMo | 09_25 | 0.2 | 23 | 0.053
FDA | MoMo | 09_25 | 0.2 | 24 | 0.055
FDA | MoMo | 09_25 | 0.2 | 25 | 0.056
FDA | MoMo | 09_25 | 0.21 | 20 | 0.049
FDA | MoMo | 09_25 | 0.21 | 21 | 0.053
FDA | MoMo | 09_25 | 0.21 | 22 | 0.055
FDA | MoMo | 09_25 | 0.21 | 23 | 0.057
FDA | MoMo | 09_25 | 0.21 | 24 | 0.059
FDA | MoMo | 09_25 | 0.21 | 25 | 0.06
FDA | MoMo | 09_25 | 0.22 | 20 | 0.053
FDA | MoMo | 09_25 | 0.22 | 21 | 0.056
FDA | MoMo | 09_25 | 0.22 | 22 | 0.059
FDA | MoMo | 09_25 | 0.22 | 23 | 0.061
FDA | MoMo | 09_25 | 0.22 | 24 | 0.062
FDA | MoMo | 09_25 | 0.22 | 25 | 0.064
FDA | MoMo | 09_25 | 0.22 | 26 | 0.065
FDA | MoMo | 09_25 | 0.23 | 20 | 0.056
FDA | MoMo | 09_25 | 0.23 | 21 | 0.06
FDA | MoMo | 09_25 | 0.23 | 22 | 0.063
FDA | MoMo | 09_25 | 0.23 | 23 | 0.064
FDA | MoMo | 09_25 | 0.23 | 24 | 0.066
FDA | MoMo | 09_25 | 0.23 | 25 | 0.067
FDA | MoMo | 09_25 | 0.23 | 26 | 0.069
FDA | MoMo | 09_25 | 0.23 | 27 | 0.07
FDA | MoMo | 09_25 | 0.23 | 28 | 0.071
FDA | MoMo | 09_25 | 0.24 | 20 | 0.06
FDA | MoMo | 09_25 | 0.24 | 21 | 0.064
FDA | MoMo | 09_25 | 0.24 | 22 | 0.066
FDA | MoMo | 09_25 | 0.24 | 23 | 0.068
FDA | MoMo | 09_25 | 0.24 | 24 | 0.07
FDA | MoMo | 09_25 | 0.24 | 25 | 0.071
FDA | MoMo | 09_25 | 0.24 | 26 | 0.072
FDA | MoMo | 09_25 | 0.24 | 27 | 0.073
FDA | MoMo | 09_25 | 0.24 | 28 | 0.075
FDA | MoMo | 09_25 | 0.24 | 29 | 0.076
FDA | MoMo | 09_25 | 0.25 | 20 | 0.064
FDA | MoMo | 09_25 | 0.25 | 21 | 0.067
FDA | MoMo | 09_25 | 0.25 | 22 | 0.07
FDA | MoMo | 09_25 | 0.25 | 23 | 0.072
FDA | MoMo | 09_25 | 0.25 | 24 | 0.073
FDA | MoMo | 09_25 | 0.25 | 25 | 0.074
FDA | MoMo | 09_25 | 0.25 | 26 | 0.076
FDA | MoMo | 09_25 | 0.25 | 27 | 0.077
FDA | MoMo | 09_25 | 0.25 | 28 | 0.078
FDA | MoMo | 09_25 | 0.25 | 29 | 0.079
FDA | MoMo | 09_25 | 0.25 | 30 | 0.08
FDA | MoMo | 09_25 | 0.25 | 31 | 0.081
FDA | MoMo | 09_25 | 0.26 | 20 | 0.067
FDA | MoMo | 09_25 | 0.26 | 21 | 0.071
FDA | MoMo | 09_25 | 0.26 | 22 | 0.073
FDA | MoMo | 09_25 | 0.26 | 23 | 0.075
FDA | MoMo | 09_25 | 0.26 | 24 | 0.077
FDA | MoMo | 09_25 | 0.26 | 25 | 0.078
FDA | MoMo | 09_25 | 0.26 | 26 | 0.079
FDA | MoMo | 09_25 | 0.26 | 27 | 0.08
FDA | MoMo | 09_25 | 0.26 | 28 | 0.081
FDA | MoMo | 09_25 | 0.26 | 29 | 0.082
FDA | MoMo | 09_25 | 0.26 | 30 | 0.083
FDA | MoMo | 09_25 | 0.26 | 31 | 0.084
FDA | MoMo | 09_25 | 0.26 | 32 | 0.085
FDA | MoMo | 09_25 | 0.27 | 20 | 0.071
FDA | MoMo | 09_25 | 0.27 | 21 | 0.075
FDA | MoMo | 09_25 | 0.27 | 22 | 0.077
FDA | MoMo | 09_25 | 0.27 | 23 | 0.079
FDA | MoMo | 09_25 | 0.27 | 24 | 0.08
FDA | MoMo | 09_25 | 0.27 | 25 | 0.081
FDA | MoMo | 09_25 | 0.27 | 26 | 0.082
FDA | MoMo | 09_25 | 0.27 | 27 | 0.083
FDA | MoMo | 09_25 | 0.27 | 28 | 0.084
FDA | MoMo | 09_25 | 0.27 | 29 | 0.085
FDA | MoMo | 09_25 | 0.27 | 30 | 0.086
FDA | MoMo | 09_25 | 0.27 | 31 | 0.087
FDA | MoMo | 09_25 | 0.27 | 32 | 0.088
FDA | MoMo | 09_25 | 0.27 | 33 | 0.089
FDA | MoMo | 09_25 | 0.27 | 34 | 0.09
FDA | MoMo | 09_25 | 0.28 | 20 | 0.074
FDA | MoMo | 09_25 | 0.28 | 21 | 0.078
FDA | MoMo | 09_25 | 0.28 | 22 | 0.08
FDA | MoMo | 09_25 | 0.28 | 23 | 0.082
FDA | MoMo | 09_25 | 0.28 | 24 | 0.083
FDA | MoMo | 09_25 | 0.28 | 25 | 0.084
FDA | MoMo | 09_25 | 0.28 | 26 | 0.085
FDA | MoMo | 09_25 | 0.28 | 27 | 0.086
FDA | MoMo | 09_25 | 0.28 | 28 | 0.087
FDA | MoMo | 09_25 | 0.28 | 29 | 0.088
FDA | MoMo | 09_25 | 0.28 | 30 | 0.089
FDA | MoMo | 09_25 | 0.28 | 31 | 0.09
FDA | MoMo | 09_25 | 0.28 | 32 | 0.091
FDA | MoMo | 09_25 | 0.28 | 33 | 0.092
FDA | MoMo | 09_25 | 0.28 | 34 | 0.093
FDA | MoMo | 09_25 | 0.28 | 35 | 0.094
FDA | MoMo | 09_25 | 0.28 | 36 | 0.095
FDA | MoMo | 09_25 | 0.29 | 20 | 0.078
FDA | MoMo | 09_25 | 0.29 | 21 | 0.082
FDA | MoMo | 09_25 | 0.29 | 22 | 0.084
FDA | MoMo | 09_25 | 0.29 | 23 | 0.085
FDA | MoMo | 09_25 | 0.29 | 24 | 0.087
FDA | MoMo | 09_25 | 0.29 | 25 | 0.088
FDA | MoMo | 09_25 | 0.29 | 26 | 0.089
FDA | MoMo | 09_25 | 0.29 | 27 | 0.09
FDA | MoMo | 09_25 | 0.29 | 28 | 0.09
FDA | MoMo | 09_25 | 0.29 | 29 | 0.091
FDA | MoMo | 09_25 | 0.29 | 30 | 0.092
FDA | MoMo | 09_25 | 0.29 | 31 | 0.093
FDA | MoMo | 09_25 | 0.29 | 32 | 0.094
FDA | MoMo | 09_25 | 0.29 | 33 | 0.095
FDA | MoMo | 09_25 | 0.29 | 34 | 0.096
FDA | MoMo | 09_25 | 0.29 | 35 | 0.097
FDA | MoMo | 09_25 | 0.29 | 36 | 0.098
FDA | MoMo | 09_25 | 0.29 | 37 | 0.099
FDA | MoMo | 09_25 | 0.29 | 38 | 0.099
FDA | MoMo | 09_25 | 0.3 | 20 | 0.081
FDA | MoMo | 09_25 | 0.3 | 21 | 0.085
FDA | MoMo | 09_25 | 0.3 | 22 | 0.088
FDA | MoMo | 09_25 | 0.3 | 23 | 0.089
FDA | MoMo | 09_25 | 0.3 | 24 | 0.09
FDA | MoMo | 09_25 | 0.3 | 25 | 0.091
FDA | MoMo | 09_25 | 0.3 | 26 | 0.092
FDA | MoMo | 09_25 | 0.3 | 27 | 0.093
FDA | MoMo | 09_25 | 0.3 | 28 | 0.093
FDA | MoMo | 09_25 | 0.3 | 29 | 0.094
FDA | MoMo | 09_25 | 0.3 | 30 | 0.095
FDA | MoMo | 09_25 | 0.3 | 31 | 0.096
FDA | MoMo | 09_25 | 0.3 | 32 | 0.097
FDA | MoMo | 09_25 | 0.3 | 33 | 0.098
FDA | MoMo | 09_25 | 0.3 | 34 | 0.099
FDA | MoMo | 09_25 | 0.3 | 35 | 0.1
FDA | MoMo | 09_25 | 0.3 | 36 | 0.101
FDA | MoMo | 09_25 | 0.3 | 37 | 0.101
FDA | MoMo | 09_25 | 0.3 | 38 | 0.102
FDA | MoMo | 09_25 | 0.3 | 39 | 0.103
FDA | MoMo | 09_25 | 0.3 | 40 | 0.104
FDA | MoMo | 09_25 | 0.3 | 41 | 0.104
FDA | MoMo | 09_25 | 0.3 | 42 | 0.105
FDA | MoMo | 09_25 | 0.31 | 20 | 0.085
FDA | MoMo | 09_25 | 0.31 | 21 | 0.089
FDA | MoMo | 09_25 | 0.31 | 22 | 0.091
FDA | MoMo | 09_25 | 0.31 | 23 | 0.092
FDA | MoMo | 09_25 | 0.31 | 24 | 0.093
FDA | MoMo | 09_25 | 0.31 | 25 | 0.094
FDA | MoMo | 09_25 | 0.31 | 26 | 0.095
FDA | MoMo | 09_25 | 0.31 | 27 | 0.096
FDA | MoMo | 09_25 | 0.31 | 28 | 0.096
FDA | MoMo | 09_25 | 0.31 | 29 | 0.097
FDA | MoMo | 09_25 | 0.31 | 30 | 0.098
FDA | MoMo | 09_25 | 0.31 | 31 | 0.099
FDA | MoMo | 09_25 | 0.31 | 32 | 0.1
FDA | MoMo | 09_25 | 0.31 | 33 | 0.101
FDA | MoMo | 09_25 | 0.31 | 34 | 0.102
FDA | MoMo | 09_25 | 0.31 | 35 | 0.102
FDA | MoMo | 09_25 | 0.31 | 36 | 0.103
FDA | MoMo | 09_25 | 0.31 | 37 | 0.104
FDA | MoMo | 09_25 | 0.31 | 38 | 0.105
FDA | MoMo | 09_25 | 0.31 | 39 | 0.106
FDA | MoMo | 09_25 | 0.31 | 40 | 0.106
FDA | MoMo | 09_25 | 0.31 | 41 | 0.107
FDA | MoMo | 09_25 | 0.31 | 42 | 0.108
FDA | MoMo | 09_25 | 0.31 | 43 | 0.108
FDA | MoMo | 09_25 | 0.32 | 20 | 0.088
FDA | MoMo | 09_25 | 0.32 | 21 | 0.092
FDA | MoMo | 09_25 | 0.32 | 22 | 0.095
FDA | MoMo | 09_25 | 0.32 | 23 | 0.096
FDA | MoMo | 09_25 | 0.32 | 24 | 0.097
FDA | MoMo | 09_25 | 0.32 | 25 | 0.097
FDA | MoMo | 09_25 | 0.32 | 26 | 0.098
FDA | MoMo | 09_25 | 0.32 | 27 | 0.099
FDA | MoMo | 09_25 | 0.32 | 28 | 0.1
FDA | MoMo | 09_25 | 0.32 | 29 | 0.1
FDA | MoMo | 09_25 | 0.32 | 30 | 0.101
FDA | MoMo | 09_25 | 0.32 | 31 | 0.102
FDA | MoMo | 09_25 | 0.32 | 32 | 0.103
FDA | MoMo | 09_25 | 0.32 | 33 | 0.104
FDA | MoMo | 09_25 | 0.32 | 34 | 0.104
FDA | MoMo | 09_25 | 0.32 | 35 | 0.105
FDA | MoMo | 09_25 | 0.32 | 36 | 0.106
FDA | MoMo | 09_25 | 0.32 | 37 | 0.107
FDA | MoMo | 09_25 | 0.32 | 38 | 0.108
FDA | MoMo | 09_25 | 0.32 | 39 | 0.108
FDA | MoMo | 09_25 | 0.32 | 40 | 0.109
FDA | MoMo | 09_25 | 0.32 | 41 | 0.11
FDA | MoMo | 09_25 | 0.32 | 42 | 0.11
FDA | MoMo | 09_25 | 0.32 | 43 | 0.111
FDA | MoMo | 09_25 | 0.32 | 44 | 0.111
FDA | MoMo | 09_25 | 0.32 | 45 | 0.112
FDA | MoMo | 09_25 | 0.32 | 46 | 0.112
FDA | MoMo | 09_25 | 0.33 | 22 | 0.098
FDA | MoMo | 09_25 | 0.33 | 23 | 0.099
FDA | MoMo | 09_25 | 0.33 | 24 | 0.1
FDA | MoMo | 09_25 | 0.33 | 25 | 0.101
FDA | MoMo | 09_25 | 0.33 | 26 | 0.101
FDA | MoMo | 09_25 | 0.33 | 27 | 0.102
FDA | MoMo | 09_25 | 0.33 | 28 | 0.102
FDA | MoMo | 09_25 | 0.33 | 29 | 0.103
FDA | MoMo | 09_25 | 0.33 | 30 | 0.104
FDA | MoMo | 09_25 | 0.33 | 31 | 0.105
FDA | MoMo | 09_25 | 0.33 | 32 | 0.105
FDA | MoMo | 09_25 | 0.33 | 33 | 0.106
FDA | MoMo | 09_25 | 0.33 | 34 | 0.107
FDA | MoMo | 09_25 | 0.33 | 35 | 0.108
FDA | MoMo | 09_25 | 0.33 | 36 | 0.109
FDA | MoMo | 09_25 | 0.33 | 37 | 0.109
FDA | MoMo | 09_25 | 0.33 | 38 | 0.11
FDA | MoMo | 09_25 | 0.33 | 39 | 0.111
FDA | MoMo | 09_25 | 0.33 | 40 | 0.112
FDA | MoMo | 09_25 | 0.33 | 41 | 0.112
FDA | MoMo | 09_25 | 0.33 | 42 | 0.113
FDA | MoMo | 09_25 | 0.33 | 43 | 0.113
FDA | MoMo | 09_25 | 0.33 | 44 | 0.114
FDA | MoMo | 09_25 | 0.33 | 45 | 0.115
FDA | MoMo | 09_25 | 0.33 | 46 | 0.115
FDA | MoMo | 09_25 | 0.33 | 47 | 0.116
FDA | MoMo | 09_25 | 0.33 | 48 | 0.116
FDA | MoMo | 09_25 | 0.33 | 49 | 0.117
FDA | MoMo | 09_25 | 0.34 | 22 | 0.101
FDA | MoMo | 09_25 | 0.34 | 23 | 0.103
FDA | MoMo | 09_25 | 0.34 | 24 | 0.103
FDA | MoMo | 09_25 | 0.34 | 25 | 0.104
FDA | MoMo | 09_25 | 0.34 | 26 | 0.104
FDA | MoMo | 09_25 | 0.34 | 27 | 0.105
FDA | MoMo | 09_25 | 0.34 | 28 | 0.105
FDA | MoMo | 09_25 | 0.34 | 29 | 0.106
FDA | MoMo | 09_25 | 0.34 | 30 | 0.107
FDA | MoMo | 09_25 | 0.34 | 31 | 0.107
FDA | MoMo | 09_25 | 0.34 | 32 | 0.108
FDA | MoMo | 09_25 | 0.34 | 33 | 0.109
FDA | MoMo | 09_25 | 0.34 | 34 | 0.11
FDA | MoMo | 09_25 | 0.34 | 35 | 0.11
FDA | MoMo | 09_25 | 0.34 | 36 | 0.111
FDA | MoMo | 09_25 | 0.34 | 37 | 0.112
FDA | MoMo | 09_25 | 0.34 | 38 | 0.113
FDA | MoMo | 09_25 | 0.34 | 39 | 0.114
FDA | MoMo | 09_25 | 0.34 | 40 | 0.114
FDA | MoMo | 09_25 | 0.34 | 41 | 0.115
FDA | MoMo | 09_25 | 0.34 | 42 | 0.116
FDA | MoMo | 09_25 | 0.34 | 43 | 0.116
FDA | MoMo | 09_25 | 0.34 | 44 | 0.117
FDA | MoMo | 09_25 | 0.34 | 45 | 0.117
FDA | MoMo | 09_25 | 0.34 | 46 | 0.118
FDA | MoMo | 09_25 | 0.34 | 47 | 0.118
FDA | MoMo | 09_25 | 0.34 | 48 | 0.119
FDA | MoMo | 09_25 | 0.34 | 49 | 0.119
FDA | MoMo | 09_25 | 0.35 | 22 | 0.105
FDA | MoMo | 09_25 | 0.35 | 23 | 0.106
FDA | MoMo | 09_25 | 0.35 | 24 | 0.107
FDA | MoMo | 09_25 | 0.35 | 25 | 0.107
FDA | MoMo | 09_25 | 0.35 | 26 | 0.108
FDA | MoMo | 09_25 | 0.35 | 27 | 0.108
FDA | MoMo | 09_25 | 0.35 | 28 | 0.108
FDA | MoMo | 09_25 | 0.35 | 29 | 0.109
FDA | MoMo | 09_25 | 0.35 | 30 | 0.11
FDA | MoMo | 09_25 | 0.35 | 31 | 0.11
FDA | MoMo | 09_25 | 0.35 | 32 | 0.111
FDA | MoMo | 09_25 | 0.35 | 33 | 0.112
FDA | MoMo | 09_25 | 0.35 | 34 | 0.112
FDA | MoMo | 09_25 | 0.35 | 35 | 0.113
FDA | MoMo | 09_25 | 0.35 | 36 | 0.114
FDA | MoMo | 09_25 | 0.35 | 37 | 0.115
FDA | MoMo | 09_25 | 0.35 | 38 | 0.115
FDA | MoMo | 09_25 | 0.35 | 39 | 0.116
FDA | MoMo | 09_25 | 0.35 | 40 | 0.117
FDA | MoMo | 09_25 | 0.35 | 41 | 0.117
FDA | MoMo | 09_25 | 0.35 | 42 | 0.118
FDA | MoMo | 09_25 | 0.35 | 43 | 0.119
FDA | MoMo | 09_25 | 0.35 | 44 | 0.119
FDA | MoMo | 09_25 | 0.35 | 45 | 0.12
FDA | MoMo | 09_25 | 0.35 | 46 | 0.12
FDA | MoMo | 09_25 | 0.35 | 47 | 0.121
FDA | MoMo | 09_25 | 0.35 | 48 | 0.121
FDA | MoMo | 09_25 | 0.35 | 49 | 0.122
FDA | MoMo | 09_25 | 0.36 | 23 | 0.109
FDA | MoMo | 09_25 | 0.36 | 24 | 0.11
FDA | MoMo | 09_25 | 0.36 | 25 | 0.11
FDA | MoMo | 09_25 | 0.36 | 26 | 0.111
FDA | MoMo | 09_25 | 0.36 | 27 | 0.111
FDA | MoMo | 09_25 | 0.36 | 28 | 0.111
FDA | MoMo | 09_25 | 0.36 | 29 | 0.112
FDA | MoMo | 09_25 | 0.36 | 30 | 0.112
FDA | MoMo | 09_25 | 0.36 | 31 | 0.113
FDA | MoMo | 09_25 | 0.36 | 32 | 0.114
FDA | MoMo | 09_25 | 0.36 | 33 | 0.114
FDA | MoMo | 09_25 | 0.36 | 34 | 0.115
FDA | MoMo | 09_25 | 0.36 | 35 | 0.116
FDA | MoMo | 09_25 | 0.36 | 36 | 0.117
FDA | MoMo | 09_25 | 0.36 | 37 | 0.117
FDA | MoMo | 09_25 | 0.36 | 38 | 0.118
FDA | MoMo | 09_25 | 0.36 | 39 | 0.119
FDA | MoMo | 09_25 | 0.36 | 40 | 0.119
FDA | MoMo | 09_25 | 0.36 | 41 | 0.12
FDA | MoMo | 09_25 | 0.36 | 42 | 0.121
FDA | MoMo | 09_25 | 0.36 | 43 | 0.121
FDA | MoMo | 09_25 | 0.36 | 44 | 0.122
FDA | MoMo | 09_25 | 0.36 | 45 | 0.122
FDA | MoMo | 09_25 | 0.36 | 46 | 0.123
FDA | MoMo | 09_25 | 0.36 | 47 | 0.123
FDA | MoMo | 09_25 | 0.36 | 48 | 0.124
FDA | MoMo | 09_25 | 0.36 | 49 | 0.124
FDA | MoMo | 09_25 | 0.37 | 24 | 0.113
FDA | MoMo | 09_25 | 0.37 | 25 | 0.114
FDA | MoMo | 09_25 | 0.37 | 26 | 0.114
FDA | MoMo | 09_25 | 0.37 | 27 | 0.114
FDA | MoMo | 09_25 | 0.37 | 28 | 0.115
FDA | MoMo | 09_25 | 0.37 | 29 | 0.115
FDA | MoMo | 09_25 | 0.37 | 30 | 0.115
FDA | MoMo | 09_25 | 0.37 | 31 | 0.116
FDA | MoMo | 09_25 | 0.37 | 32 | 0.116
FDA | MoMo | 09_25 | 0.37 | 33 | 0.117
FDA | MoMo | 09_25 | 0.37 | 34 | 0.118
FDA | MoMo | 09_25 | 0.37 | 35 | 0.118
FDA | MoMo | 09_25 | 0.37 | 36 | 0.119
FDA | MoMo | 09_25 | 0.37 | 37 | 0.12
FDA | MoMo | 09_25 | 0.37 | 38 | 0.121
FDA | MoMo | 09_25 | 0.37 | 39 | 0.121
FDA | MoMo | 09_25 | 0.37 | 40 | 0.122
FDA | MoMo | 09_25 | 0.37 | 41 | 0.123
FDA | MoMo | 09_25 | 0.37 | 42 | 0.123
FDA | MoMo | 09_25 | 0.37 | 43 | 0.124
FDA | MoMo | 09_25 | 0.37 | 44 | 0.124
FDA | MoMo | 09_25 | 0.37 | 45 | 0.125
FDA | MoMo | 09_25 | 0.37 | 46 | 0.126
FDA | MoMo | 09_25 | 0.37 | 47 | 0.126
FDA | MoMo | 09_25 | 0.37 | 48 | 0.127
FDA | MoMo | 09_25 | 0.37 | 49 | 0.127
FDA | MoMo | 09_25 | 0.38 | 26 | 0.117
FDA | MoMo | 09_25 | 0.38 | 27 | 0.117
FDA | MoMo | 09_25 | 0.38 | 28 | 0.118
FDA | MoMo | 09_25 | 0.38 | 29 | 0.118
FDA | MoMo | 09_25 | 0.38 | 30 | 0.118
FDA | MoMo | 09_25 | 0.38 | 31 | 0.119
FDA | MoMo | 09_25 | 0.38 | 32 | 0.119
FDA | MoMo | 09_25 | 0.38 | 33 | 0.12
FDA | MoMo | 09_25 | 0.38 | 34 | 0.121
FDA | MoMo | 09_25 | 0.38 | 35 | 0.121
FDA | MoMo | 09_25 | 0.38 | 36 | 0.122
FDA | MoMo | 09_25 | 0.38 | 37 | 0.123
FDA | MoMo | 09_25 | 0.38 | 38 | 0.123
FDA | MoMo | 09_25 | 0.38 | 39 | 0.124
FDA | MoMo | 09_25 | 0.38 | 40 | 0.125
FDA | MoMo | 09_25 | 0.38 | 41 | 0.125
FDA | MoMo | 09_25 | 0.38 | 42 | 0.126
FDA | MoMo | 09_25 | 0.38 | 43 | 0.126
FDA | MoMo | 09_25 | 0.38 | 44 | 0.127
FDA | MoMo | 09_25 | 0.38 | 45 | 0.128
FDA | MoMo | 09_25 | 0.38 | 46 | 0.128
FDA | MoMo | 09_25 | 0.38 | 47 | 0.129
FDA | MoMo | 09_25 | 0.38 | 48 | 0.129
FDA | MoMo | 09_25 | 0.38 | 49 | 0.13
FDA | MoMo | 09_25 | 0.39 | 26 | 0.12
FDA | MoMo | 09_25 | 0.39 | 27 | 0.121
FDA | MoMo | 09_25 | 0.39 | 28 | 0.121
FDA | MoMo | 09_25 | 0.39 | 29 | 0.121
FDA | MoMo | 09_25 | 0.39 | 30 | 0.121
FDA | MoMo | 09_25 | 0.39 | 31 | 0.122
FDA | MoMo | 09_25 | 0.39 | 32 | 0.122
FDA | MoMo | 09_25 | 0.39 | 33 | 0.123
FDA | MoMo | 09_25 | 0.39 | 34 | 0.123
FDA | MoMo | 09_25 | 0.39 | 35 | 0.124
FDA | MoMo | 09_25 | 0.39 | 36 | 0.124
FDA | MoMo | 09_25 | 0.39 | 37 | 0.125
FDA | MoMo | 09_25 | 0.39 | 38 | 0.126
FDA | MoMo | 09_25 | 0.39 | 39 | 0.127
FDA | MoMo | 09_25 | 0.39 | 40 | 0.127
FDA | MoMo | 09_25 | 0.39 | 41 | 0.128
FDA | MoMo | 09_25 | 0.39 | 42 | 0.128
FDA | MoMo | 09_25 | 0.39 | 43 | 0.129
FDA | MoMo | 09_25 | 0.39 | 44 | 0.13
FDA | MoMo | 09_25 | 0.39 | 45 | 0.13
FDA | MoMo | 09_25 | 0.39 | 46 | 0.131
FDA | MoMo | 09_25 | 0.39 | 47 | 0.131
FDA | MoMo | 09_25 | 0.39 | 48 | 0.132
FDA | MoMo | 09_25 | 0.39 | 49 | 0.132
FDA | MoMo | 09_25 | 0.4 | 27 | 0.124
FDA | MoMo | 09_25 | 0.4 | 28 | 0.124
FDA | MoMo | 09_25 | 0.4 | 29 | 0.124
FDA | MoMo | 09_25 | 0.4 | 30 | 0.124
FDA | MoMo | 09_25 | 0.4 | 31 | 0.125
FDA | MoMo | 09_25 | 0.4 | 32 | 0.125
FDA | MoMo | 09_25 | 0.4 | 33 | 0.125
FDA | MoMo | 09_25 | 0.4 | 34 | 0.126
FDA | MoMo | 09_25 | 0.4 | 35 | 0.127
FDA | MoMo | 09_25 | 0.4 | 36 | 0.127
FDA | MoMo | 09_25 | 0.4 | 37 | 0.128
FDA | MoMo | 09_25 | 0.4 | 38 | 0.129
FDA | MoMo | 09_25 | 0.4 | 39 | 0.129
FDA | MoMo | 09_25 | 0.4 | 40 | 0.13
FDA | MoMo | 09_25 | 0.4 | 41 | 0.13
FDA | MoMo | 09_25 | 0.4 | 42 | 0.131
FDA | MoMo | 09_25 | 0.4 | 43 | 0.132
FDA | MoMo | 09_25 | 0.4 | 44 | 0.132
FDA | MoMo | 09_25 | 0.4 | 45 | 0.133
FDA | MoMo | 09_25 | 0.4 | 46 | 0.133
FDA | MoMo | 09_25 | 0.4 | 47 | 0.134
FDA | MoMo | 09_25 | 0.4 | 48 | 0.134
FDA | MoMo | 09_25 | 0.4 | 49 | 0.135
FDA | MoMo | 09_25 | 0.41 | 29 | 0.127
FDA | MoMo | 09_25 | 0.41 | 30 | 0.127
FDA | MoMo | 09_25 | 0.41 | 31 | 0.128
FDA | MoMo | 09_25 | 0.41 | 32 | 0.128
FDA | MoMo | 09_25 | 0.41 | 33 | 0.128
FDA | MoMo | 09_25 | 0.41 | 34 | 0.129
FDA | MoMo | 09_25 | 0.41 | 35 | 0.129
FDA | MoMo | 09_25 | 0.41 | 36 | 0.13
FDA | MoMo | 09_25 | 0.41 | 37 | 0.131
FDA | MoMo | 09_25 | 0.41 | 38 | 0.131
FDA | MoMo | 09_25 | 0.41 | 39 | 0.132
FDA | MoMo | 09_25 | 0.41 | 40 | 0.132
FDA | MoMo | 09_25 | 0.41 | 41 | 0.133
FDA | MoMo | 09_25 | 0.41 | 42 | 0.134
FDA | MoMo | 09_25 | 0.41 | 43 | 0.134
FDA | MoMo | 09_25 | 0.41 | 44 | 0.135
FDA | MoMo | 09_25 | 0.41 | 45 | 0.135
FDA | MoMo | 09_25 | 0.41 | 46 | 0.136
FDA | MoMo | 09_25 | 0.41 | 47 | 0.137
FDA | MoMo | 09_25 | 0.41 | 48 | 0.137
FDA | MoMo | 09_25 | 0.41 | 49 | 0.138
FDA | MoMo | 09_25 | 0.42 | 30 | 0.13
FDA | MoMo | 09_25 | 0.42 | 31 | 0.131
FDA | MoMo | 09_25 | 0.42 | 32 | 0.131
FDA | MoMo | 09_25 | 0.42 | 33 | 0.131
FDA | MoMo | 09_25 | 0.42 | 34 | 0.132
FDA | MoMo | 09_25 | 0.42 | 35 | 0.132
FDA | MoMo | 09_25 | 0.42 | 36 | 0.133
FDA | MoMo | 09_25 | 0.42 | 37 | 0.133
FDA | MoMo | 09_25 | 0.42 | 38 | 0.134
FDA | MoMo | 09_25 | 0.42 | 39 | 0.134
FDA | MoMo | 09_25 | 0.42 | 40 | 0.135
FDA | MoMo | 09_25 | 0.42 | 41 | 0.136
FDA | MoMo | 09_25 | 0.42 | 42 | 0.136
FDA | MoMo | 09_25 | 0.42 | 43 | 0.137
FDA | MoMo | 09_25 | 0.42 | 44 | 0.138
FDA | MoMo | 09_25 | 0.42 | 45 | 0.138
FDA | MoMo | 09_25 | 0.42 | 46 | 0.139
FDA | MoMo | 09_25 | 0.42 | 47 | 0.139
FDA | MoMo | 09_25 | 0.42 | 48 | 0.14
FDA | MoMo | 09_25 | 0.42 | 49 | 0.14
FDA | MoMo | 09_25 | 0.43 | 32 | 0.134
FDA | MoMo | 09_25 | 0.43 | 33 | 0.134
FDA | MoMo | 09_25 | 0.43 | 34 | 0.135
FDA | MoMo | 09_25 | 0.43 | 35 | 0.135
FDA | MoMo | 09_25 | 0.43 | 36 | 0.136
FDA | MoMo | 09_25 | 0.43 | 37 | 0.136
FDA | MoMo | 09_25 | 0.43 | 38 | 0.137
FDA | MoMo | 09_25 | 0.43 | 39 | 0.137
FDA | MoMo | 09_25 | 0.43 | 40 | 0.138
FDA | MoMo | 09_25 | 0.43 | 41 | 0.139
FDA | MoMo | 09_25 | 0.43 | 42 | 0.139
FDA | MoMo | 09_25 | 0.43 | 43 | 0.14
FDA | MoMo | 09_25 | 0.43 | 44 | 0.14
FDA | MoMo | 09_25 | 0.43 | 45 | 0.141
FDA | MoMo | 09_25 | 0.43 | 46 | 0.141
FDA | MoMo | 09_25 | 0.43 | 47 | 0.142
FDA | MoMo | 09_25 | 0.43 | 48 | 0.143
FDA | MoMo | 09_25 | 0.43 | 49 | 0.143
FDA | MoMo | 09_25 | 0.44 | 33 | 0.137
FDA | MoMo | 09_25 | 0.44 | 34 | 0.138
FDA | MoMo | 09_25 | 0.44 | 35 | 0.138
FDA | MoMo | 09_25 | 0.44 | 36 | 0.138
FDA | MoMo | 09_25 | 0.44 | 37 | 0.139
FDA | MoMo | 09_25 | 0.44 | 38 | 0.14
FDA | MoMo | 09_25 | 0.44 | 39 | 0.14
FDA | MoMo | 09_25 | 0.44 | 40 | 0.141
FDA | MoMo | 09_25 | 0.44 | 41 | 0.141
FDA | MoMo | 09_25 | 0.44 | 42 | 0.142
FDA | MoMo | 09_25 | 0.44 | 43 | 0.142
FDA | MoMo | 09_25 | 0.44 | 44 | 0.143
FDA | MoMo | 09_25 | 0.44 | 45 | 0.144
FDA | MoMo | 09_25 | 0.44 | 46 | 0.144
FDA | MoMo | 09_25 | 0.44 | 47 | 0.145
FDA | MoMo | 09_25 | 0.44 | 48 | 0.145
FDA | MoMo | 09_25 | 0.44 | 49 | 0.146
FDA | MoMo | 09_25 | 0.45 | 35 | 0.141
FDA | MoMo | 09_25 | 0.45 | 36 | 0.141
FDA | MoMo | 09_25 | 0.45 | 37 | 0.142
FDA | MoMo | 09_25 | 0.45 | 38 | 0.143
FDA | MoMo | 09_25 | 0.45 | 39 | 0.143
FDA | MoMo | 09_25 | 0.45 | 40 | 0.144
FDA | MoMo | 09_25 | 0.45 | 41 | 0.144
FDA | MoMo | 09_25 | 0.45 | 42 | 0.145
FDA | MoMo | 09_25 | 0.45 | 43 | 0.145
FDA | MoMo | 09_25 | 0.45 | 44 | 0.146
FDA | MoMo | 09_25 | 0.45 | 45 | 0.147
FDA | MoMo | 09_25 | 0.45 | 46 | 0.147
FDA | MoMo | 09_25 | 0.45 | 47 | 0.148
FDA | MoMo | 09_25 | 0.45 | 48 | 0.148
FDA | MoMo | 09_25 | 0.45 | 49 | 0.149
FDA | MoMo | 09_25 | 0.46 | 37 | 0.145
FDA | MoMo | 09_25 | 0.46 | 38 | 0.146
FDA | MoMo | 09_25 | 0.46 | 39 | 0.146
FDA | MoMo | 09_25 | 0.46 | 40 | 0.147
FDA | MoMo | 09_25 | 0.46 | 41 | 0.147
FDA | MoMo | 09_25 | 0.46 | 42 | 0.148
FDA | MoMo | 09_25 | 0.46 | 43 | 0.148
FDA | MoMo | 09_25 | 0.46 | 44 | 0.149
FDA | MoMo | 09_25 | 0.46 | 45 | 0.15
FDA | MoMo | 09_25 | 0.46 | 46 | 0.15
FDA | MoMo | 09_25 | 0.46 | 47 | 0.151
FDA | MoMo | 09_25 | 0.46 | 48 | 0.151
FDA | MoMo | 09_25 | 0.46 | 49 | 0.152
FDA | MoMo | 09_25 | 0.47 | 39 | 0.149
FDA | MoMo | 09_25 | 0.47 | 40 | 0.15
FDA | MoMo | 09_25 | 0.47 | 41 | 0.15
FDA | MoMo | 09_25 | 0.47 | 42 | 0.151
FDA | MoMo | 09_25 | 0.47 | 43 | 0.152
FDA | MoMo | 09_25 | 0.47 | 44 | 0.152
FDA | MoMo | 09_25 | 0.47 | 45 | 0.153
FDA | MoMo | 09_25 | 0.47 | 46 | 0.153
FDA | MoMo | 09_25 | 0.47 | 47 | 0.154
FDA | MoMo | 09_25 | 0.47 | 48 | 0.155
FDA | MoMo | 09_25 | 0.47 | 49 | 0.155
FDA | MoMo | 09_25 | 0.48 | 43 | 0.155
FDA | MoMo | 09_25 | 0.48 | 44 | 0.155
FDA | MoMo | 09_25 | 0.48 | 45 | 0.156
FDA | MoMo | 09_25 | 0.48 | 46 | 0.157
FDA | MoMo | 09_25 | 0.48 | 47 | 0.157
FDA | MoMo | 09_25 | 0.48 | 48 | 0.158
FDA | MoMo | 09_25 | 0.48 | 49 | 0.159
FDA | MoMo | 09_25 | 0.49 | 44 | 0.159
FDA | MoMo | 09_25 | 0.49 | 45 | 0.16
FDA | MoMo | 09_25 | 0.49 | 46 | 0.16
FDA | MoMo | 09_25 | 0.49 | 47 | 0.161
FDA | MoMo | 09_25 | 0.49 | 48 | 0.162
FDA | MoMo | 09_25 | 0.49 | 49 | 0.162
FDA | MoMo | 09_25 | 0.5 | 47 | 0.165
FDA | MoMo | 09_25 | 0.5 | 48 | 0.166
FDA | MoMo | 09_25 | 0.5 | 49 | 0.166
FDA | MoRh | 09_25 | 0.195 | 20 | 0.044
FDA | MoRh | 09_25 | 0.195 | 21 | 0.047
FDA | MoRh | 09_25 | 0.195 | 22 | 0.05
FDA | MoRh | 09_25 | 0.21 | 20 | 0.049
FDA | MoRh | 09_25 | 0.21 | 21 | 0.053
FDA | MoRh | 09_25 | 0.21 | 22 | 0.056
FDA | MoRh | 09_25 | 0.21 | 23 | 0.058
FDA | MoRh | 09_25 | 0.21 | 24 | 0.061
FDA | MoRh | 09_25 | 0.225 | 20 | 0.055
FDA | MoRh | 09_25 | 0.225 | 21 | 0.058
FDA | MoRh | 09_25 | 0.225 | 22 | 0.061
FDA | MoRh | 09_25 | 0.225 | 23 | 0.064
FDA | MoRh | 09_25 | 0.225 | 24 | 0.067
FDA | MoRh | 09_25 | 0.225 | 25 | 0.068
FDA | MoRh | 09_25 | 0.225 | 26 | 0.07
FDA | MoRh | 09_25 | 0.225 | 27 | 0.071
FDA | MoRh | 09_25 | 0.24 | 20 | 0.06
FDA | MoRh | 09_25 | 0.24 | 21 | 0.064
FDA | MoRh | 09_25 | 0.24 | 22 | 0.067
FDA | MoRh | 09_25 | 0.24 | 23 | 0.07
FDA | MoRh | 09_25 | 0.24 | 24 | 0.072
FDA | MoRh | 09_25 | 0.24 | 25 | 0.074
FDA | MoRh | 09_25 | 0.24 | 26 | 0.076
FDA | MoRh | 09_25 | 0.24 | 27 | 0.077
FDA | MoRh | 09_25 | 0.255 | 20 | 0.065
FDA | MoRh | 09_25 | 0.255 | 21 | 0.069
FDA | MoRh | 09_25 | 0.255 | 22 | 0.073
FDA | MoRh | 09_25 | 0.255 | 23 | 0.076
FDA | MoRh | 09_25 | 0.255 | 24 | 0.078
FDA | MoRh | 09_25 | 0.255 | 25 | 0.08
FDA | MoRh | 09_25 | 0.255 | 26 | 0.081
FDA | MoRh | 09_25 | 0.255 | 27 | 0.082
FDA | MoRh | 09_25 | 0.255 | 28 | 0.083
FDA | MoRh | 09_25 | 0.27 | 20 | 0.071
FDA | MoRh | 09_25 | 0.27 | 21 | 0.075
FDA | MoRh | 09_25 | 0.27 | 22 | 0.078
FDA | MoRh | 09_25 | 0.27 | 23 | 0.081
FDA | MoRh | 09_25 | 0.27 | 24 | 0.084
FDA | MoRh | 09_25 | 0.27 | 25 | 0.086
FDA | MoRh | 09_25 | 0.27 | 26 | 0.087
FDA | MoRh | 09_25 | 0.27 | 27 | 0.088
FDA | MoRh | 09_25 | 0.27 | 28 | 0.089
FDA | MoRh | 09_25 | 0.27 | 29 | 0.09
FDA | MoRh | 09_25 | 0.27 | 30 | 0.09
FDA | MoRh | 09_25 | 0.27 | 31 | 0.091
FDA | MoRh | 09_25 | 0.285 | 20 | 0.076
FDA | MoRh | 09_25 | 0.285 | 21 | 0.08
FDA | MoRh | 09_25 | 0.285 | 22 | 0.083
FDA | MoRh | 09_25 | 0.285 | 23 | 0.086
FDA | MoRh | 09_25 | 0.285 | 24 | 0.089
FDA | MoRh | 09_25 | 0.285 | 25 | 0.091
FDA | MoRh | 09_25 | 0.285 | 26 | 0.092
FDA | MoRh | 09_25 | 0.285 | 27 | 0.093
FDA | MoRh | 09_25 | 0.285 | 28 | 0.094
FDA | MoRh | 09_25 | 0.285 | 29 | 0.095
FDA | MoRh | 09_25 | 0.285 | 30 | 0.096
FDA | MoRh | 09_25 | 0.285 | 31 | 0.096
FDA | MoRh | 09_25 | 0.285 | 32 | 0.097
FDA | MoRh | 09_25 | 0.285 | 33 | 0.098
FDA | MoRh | 09_25 | 0.3 | 20 | 0.081
FDA | MoRh | 09_25 | 0.3 | 21 | 0.085
FDA | MoRh | 09_25 | 0.3 | 22 | 0.089
FDA | MoRh | 09_25 | 0.3 | 23 | 0.092
FDA | MoRh | 09_25 | 0.3 | 24 | 0.094
FDA | MoRh | 09_25 | 0.3 | 25 | 0.096
FDA | MoRh | 09_25 | 0.3 | 26 | 0.098
FDA | MoRh | 09_25 | 0.3 | 27 | 0.098
FDA | MoRh | 09_25 | 0.3 | 28 | 0.099
FDA | MoRh | 09_25 | 0.3 | 29 | 0.1
FDA | MoRh | 09_25 | 0.3 | 30 | 0.101
FDA | MoRh | 09_25 | 0.3 | 31 | 0.101
FDA | MoRh | 09_25 | 0.3 | 32 | 0.102
FDA | MoRh | 09_25 | 0.3 | 33 | 0.102
FDA | MoRh | 09_25 | 0.3 | 34 | 0.103
FDA | MoRh | 09_25 | 0.3 | 35 | 0.104
FDA | MoRh | 09_25 | 0.3 | 36 | 0.104
FDA | MoRh | 09_25 | 0.3 | 37 | 0.105
FDA | MoRh | 09_25 | 0.3 | 38 | 0.105
FDA | MoRh | 09_25 | 0.3 | 39 | 0.106
FDA | MoRh | 09_25 | 0.3 | 40 | 0.106
FDA | MoRh | 09_25 | 0.3 | 41 | 0.107
FDA | MoRh | 09_25 | 0.3 | 42 | 0.107
FDA | MoRh | 09_25 | 0.315 | 20 | 0.087
FDA | MoRh | 09_25 | 0.315 | 21 | 0.091
FDA | MoRh | 09_25 | 0.315 | 22 | 0.094
FDA | MoRh | 09_25 | 0.315 | 23 | 0.097
FDA | MoRh | 09_25 | 0.315 | 24 | 0.1
FDA | MoRh | 09_25 | 0.315 | 25 | 0.102
FDA | MoRh | 09_25 | 0.315 | 26 | 0.103
FDA | MoRh | 09_25 | 0.315 | 27 | 0.104
FDA | MoRh | 09_25 | 0.315 | 28 | 0.104
FDA | MoRh | 09_25 | 0.315 | 29 | 0.105
FDA | MoRh | 09_25 | 0.315 | 30 | 0.105
FDA | MoRh | 09_25 | 0.315 | 31 | 0.106
FDA | MoRh | 09_25 | 0.315 | 32 | 0.107
FDA | MoRh | 09_25 | 0.315 | 33 | 0.107
FDA | MoRh | 09_25 | 0.315 | 34 | 0.108
FDA | MoRh | 09_25 | 0.315 | 35 | 0.108
FDA | MoRh | 09_25 | 0.315 | 36 | 0.109
FDA | MoRh | 09_25 | 0.315 | 37 | 0.109
FDA | MoRh | 09_25 | 0.315 | 38 | 0.11
FDA | MoRh | 09_25 | 0.315 | 39 | 0.111
FDA | MoRh | 09_25 | 0.315 | 40 | 0.111
FDA | MoRh | 09_25 | 0.315 | 41 | 0.112
FDA | MoRh | 09_25 | 0.315 | 42 | 0.112
FDA | MoRh | 09_25 | 0.33 | 20 | 0.092
FDA | MoRh | 09_25 | 0.33 | 21 | 0.096
FDA | MoRh | 09_25 | 0.33 | 22 | 0.099
FDA | MoRh | 09_25 | 0.33 | 23 | 0.102
FDA | MoRh | 09_25 | 0.33 | 24 | 0.105
FDA | MoRh | 09_25 | 0.33 | 25 | 0.107
FDA | MoRh | 09_25 | 0.33 | 26 | 0.108
FDA | MoRh | 09_25 | 0.33 | 27 | 0.109
FDA | MoRh | 09_25 | 0.33 | 28 | 0.109
FDA | MoRh | 09_25 | 0.33 | 29 | 0.11
FDA | MoRh | 09_25 | 0.33 | 30 | 0.11
FDA | MoRh | 09_25 | 0.33 | 31 | 0.111
FDA | MoRh | 09_25 | 0.33 | 32 | 0.111
FDA | MoRh | 09_25 | 0.33 | 33 | 0.112
FDA | MoRh | 09_25 | 0.33 | 34 | 0.113
FDA | MoRh | 09_25 | 0.33 | 35 | 0.113
FDA | MoRh | 09_25 | 0.33 | 36 | 0.114
FDA | MoRh | 09_25 | 0.33 | 37 | 0.114
FDA | MoRh | 09_25 | 0.33 | 38 | 0.115
FDA | MoRh | 09_25 | 0.33 | 39 | 0.115
FDA | MoRh | 09_25 | 0.33 | 40 | 0.116
FDA | MoRh | 09_25 | 0.33 | 41 | 0.116
FDA | MoRh | 09_25 | 0.33 | 42 | 0.116
FDA | MoRh | 09_25 | 0.345 | 20 | 0.097
FDA | MoRh | 09_25 | 0.345 | 21 | 0.101
FDA | MoRh | 09_25 | 0.345 | 22 | 0.105
FDA | MoRh | 09_25 | 0.345 | 23 | 0.108
FDA | MoRh | 09_25 | 0.345 | 24 | 0.11
FDA | MoRh | 09_25 | 0.345 | 25 | 0.112
FDA | MoRh | 09_25 | 0.345 | 26 | 0.113
FDA | MoRh | 09_25 | 0.345 | 27 | 0.114
FDA | MoRh | 09_25 | 0.345 | 28 | 0.114
FDA | MoRh | 09_25 | 0.345 | 29 | 0.115
FDA | MoRh | 09_25 | 0.345 | 30 | 0.115
FDA | MoRh | 09_25 | 0.345 | 31 | 0.116
FDA | MoRh | 09_25 | 0.345 | 32 | 0.116
FDA | MoRh | 09_25 | 0.345 | 33 | 0.117
FDA | MoRh | 09_25 | 0.345 | 34 | 0.117
FDA | MoRh | 09_25 | 0.345 | 35 | 0.118
FDA | MoRh | 09_25 | 0.345 | 36 | 0.118
FDA | MoRh | 09_25 | 0.345 | 37 | 0.119
FDA | MoRh | 09_25 | 0.345 | 38 | 0.119
FDA | MoRh | 09_25 | 0.345 | 39 | 0.12
FDA | MoRh | 09_25 | 0.345 | 40 | 0.12
FDA | MoRh | 09_25 | 0.345 | 41 | 0.121
FDA | MoRh | 09_25 | 0.345 | 42 | 0.121
FDA | MoRh | 09_25 | 0.345 | 43 | 0.121
FDA | MoRh | 09_25 | 0.345 | 44 | 0.122
FDA | MoRh | 09_25 | 0.345 | 45 | 0.122
FDA | MoRh | 09_25 | 0.345 | 46 | 0.123
FDA | MoRh | 09_25 | 0.36 | 20 | 0.103
FDA | MoRh | 09_25 | 0.36 | 21 | 0.107
FDA | MoRh | 09_25 | 0.36 | 22 | 0.11
FDA | MoRh | 09_25 | 0.36 | 23 | 0.113
FDA | MoRh | 09_25 | 0.36 | 24 | 0.116
FDA | MoRh | 09_25 | 0.36 | 25 | 0.117
FDA | MoRh | 09_25 | 0.36 | 26 | 0.118
FDA | MoRh | 09_25 | 0.36 | 27 | 0.119
FDA | MoRh | 09_25 | 0.36 | 28 | 0.119
FDA | MoRh | 09_25 | 0.36 | 29 | 0.12
FDA | MoRh | 09_25 | 0.36 | 30 | 0.12
FDA | MoRh | 09_25 | 0.36 | 31 | 0.12
FDA | MoRh | 09_25 | 0.36 | 32 | 0.121
FDA | MoRh | 09_25 | 0.36 | 33 | 0.121
FDA | MoRh | 09_25 | 0.36 | 34 | 0.122
FDA | MoRh | 09_25 | 0.36 | 35 | 0.122
FDA | MoRh | 09_25 | 0.36 | 36 | 0.123
FDA | MoRh | 09_25 | 0.36 | 37 | 0.123
FDA | MoRh | 09_25 | 0.36 | 38 | 0.124
FDA | MoRh | 09_25 | 0.36 | 39 | 0.124
FDA | MoRh | 09_25 | 0.36 | 40 | 0.125
FDA | MoRh | 09_25 | 0.36 | 41 | 0.125
FDA | MoRh | 09_25 | 0.36 | 42 | 0.125
FDA | MoRh | 09_25 | 0.36 | 43 | 0.126
FDA | MoRh | 09_25 | 0.36 | 44 | 0.126
FDA | MoRh | 09_25 | 0.36 | 45 | 0.127
FDA | MoRh | 09_25 | 0.36 | 46 | 0.127
FDA | MoRh | 09_25 | 0.36 | 47 | 0.127
FDA | MoRh | 09_25 | 0.36 | 48 | 0.128
FDA | MoRh | 09_25 | 0.36 | 49 | 0.128
FDA | MoRh | 09_25 | 0.375 | 20 | 0.108
FDA | MoRh | 09_25 | 0.375 | 21 | 0.112
FDA | MoRh | 09_25 | 0.375 | 22 | 0.115
FDA | MoRh | 09_25 | 0.375 | 23 | 0.118
FDA | MoRh | 09_25 | 0.375 | 24 | 0.121
FDA | MoRh | 09_25 | 0.375 | 25 | 0.123
FDA | MoRh | 09_25 | 0.375 | 26 | 0.123
FDA | MoRh | 09_25 | 0.375 | 27 | 0.124
FDA | MoRh | 09_25 | 0.375 | 28 | 0.124
FDA | MoRh | 09_25 | 0.375 | 29 | 0.125
FDA | MoRh | 09_25 | 0.375 | 30 | 0.125
FDA | MoRh | 09_25 | 0.375 | 31 | 0.125
FDA | MoRh | 09_25 | 0.375 | 32 | 0.126
FDA | MoRh | 09_25 | 0.375 | 33 | 0.126
FDA | MoRh | 09_25 | 0.375 | 34 | 0.126
FDA | MoRh | 09_25 | 0.375 | 35 | 0.127
FDA | MoRh | 09_25 | 0.375 | 36 | 0.127
FDA | MoRh | 09_25 | 0.375 | 37 | 0.128
FDA | MoRh | 09_25 | 0.375 | 38 | 0.128
FDA | MoRh | 09_25 | 0.375 | 39 | 0.129
FDA | MoRh | 09_25 | 0.375 | 40 | 0.129
FDA | MoRh | 09_25 | 0.375 | 41 | 0.129
FDA | MoRh | 09_25 | 0.375 | 42 | 0.13
FDA | MoRh | 09_25 | 0.375 | 43 | 0.13
FDA | MoRh | 09_25 | 0.375 | 44 | 0.131
FDA | MoRh | 09_25 | 0.375 | 45 | 0.131
FDA | MoRh | 09_25 | 0.375 | 46 | 0.131
FDA | MoRh | 09_25 | 0.375 | 47 | 0.132
FDA | MoRh | 09_25 | 0.375 | 48 | 0.132
FDA | MoRh | 09_25 | 0.375 | 49 | 0.132
FDA | MoRh | 09_25 | 0.39 | 20 | 0.114
FDA | MoRh | 09_25 | 0.39 | 21 | 0.118
FDA | MoRh | 09_25 | 0.39 | 22 | 0.121
FDA | MoRh | 09_25 | 0.39 | 23 | 0.124
FDA | MoRh | 09_25 | 0.39 | 24 | 0.126
FDA | MoRh | 09_25 | 0.39 | 25 | 0.128
FDA | MoRh | 09_25 | 0.39 | 26 | 0.129
FDA | MoRh | 09_25 | 0.39 | 27 | 0.129
FDA | MoRh | 09_25 | 0.39 | 28 | 0.129
FDA | MoRh | 09_25 | 0.39 | 29 | 0.13
FDA | MoRh | 09_25 | 0.39 | 30 | 0.13
FDA | MoRh | 09_25 | 0.39 | 31 | 0.13
FDA | MoRh | 09_25 | 0.39 | 32 | 0.13
FDA | MoRh | 09_25 | 0.39 | 33 | 0.131
FDA | MoRh | 09_25 | 0.39 | 34 | 0.131
FDA | MoRh | 09_25 | 0.39 | 35 | 0.131
FDA | MoRh | 09_25 | 0.39 | 36 | 0.132
FDA | MoRh | 09_25 | 0.39 | 37 | 0.132
FDA | MoRh | 09_25 | 0.39 | 38 | 0.133
FDA | MoRh | 09_25 | 0.39 | 39 | 0.133
FDA | MoRh | 09_25 | 0.39 | 40 | 0.133
FDA | MoRh | 09_25 | 0.39 | 41 | 0.134
FDA | MoRh | 09_25 | 0.39 | 42 | 0.134
FDA | MoRh | 09_25 | 0.39 | 43 | 0.135
FDA | MoRh | 09_25 | 0.39 | 44 | 0.135
FDA | MoRh | 09_25 | 0.39 | 45 | 0.135
FDA | MoRh | 09_25 | 0.39 | 46 | 0.136
FDA | MoRh | 09_25 | 0.39 | 47 | 0.136
FDA | MoRh | 09_25 | 0.39 | 48 | 0.136
FDA | MoRh | 09_25 | 0.39 | 49 | 0.137
FDA | MoRh | 09_25 | 0.405 | 20 | 0.119
FDA | MoRh | 09_25 | 0.405 | 21 | 0.123
FDA | MoRh | 09_25 | 0.405 | 22 | 0.126
FDA | MoRh | 09_25 | 0.405 | 23 | 0.129
FDA | MoRh | 09_25 | 0.405 | 24 | 0.131
FDA | MoRh | 09_25 | 0.405 | 25 | 0.133
FDA | MoRh | 09_25 | 0.405 | 26 | 0.134
FDA | MoRh | 09_25 | 0.405 | 27 | 0.134
FDA | MoRh | 09_25 | 0.405 | 28 | 0.134
FDA | MoRh | 09_25 | 0.405 | 29 | 0.135
FDA | MoRh | 09_25 | 0.405 | 30 | 0.135
FDA | MoRh | 09_25 | 0.405 | 31 | 0.135
FDA | MoRh | 09_25 | 0.405 | 32 | 0.135
FDA | MoRh | 09_25 | 0.405 | 33 | 0.135
FDA | MoRh | 09_25 | 0.405 | 34 | 0.136
FDA | MoRh | 09_25 | 0.405 | 35 | 0.136
FDA | MoRh | 09_25 | 0.405 | 36 | 0.136
FDA | MoRh | 09_25 | 0.405 | 37 | 0.137
FDA | MoRh | 09_25 | 0.405 | 38 | 0.137
FDA | MoRh | 09_25 | 0.405 | 39 | 0.137
FDA | MoRh | 09_25 | 0.405 | 40 | 0.138
FDA | MoRh | 09_25 | 0.405 | 41 | 0.138
FDA | MoRh | 09_25 | 0.405 | 42 | 0.139
FDA | MoRh | 09_25 | 0.405 | 43 | 0.139
FDA | MoRh | 09_25 | 0.405 | 44 | 0.139
FDA | MoRh | 09_25 | 0.405 | 45 | 0.14
FDA | MoRh | 09_25 | 0.405 | 46 | 0.14
FDA | MoRh | 09_25 | 0.405 | 47 | 0.14
FDA | MoRh | 09_25 | 0.405 | 48 | 0.141
FDA | MoRh | 09_25 | 0.405 | 49 | 0.141
FDA | MoRh | 09_25 | 0.42 | 23 | 0.134
FDA | MoRh | 09_25 | 0.42 | 24 | 0.137
FDA | MoRh | 09_25 | 0.42 | 25 | 0.138
FDA | MoRh | 09_25 | 0.42 | 26 | 0.139
FDA | MoRh | 09_25 | 0.42 | 27 | 0.139
FDA | MoRh | 09_25 | 0.42 | 28 | 0.14
FDA | MoRh | 09_25 | 0.42 | 29 | 0.14
FDA | MoRh | 09_25 | 0.42 | 30 | 0.14
FDA | MoRh | 09_25 | 0.42 | 31 | 0.14
FDA | MoRh | 09_25 | 0.42 | 32 | 0.14
FDA | MoRh | 09_25 | 0.42 | 33 | 0.14
FDA | MoRh | 09_25 | 0.42 | 34 | 0.14
FDA | MoRh | 09_25 | 0.42 | 35 | 0.141
FDA | MoRh | 09_25 | 0.42 | 36 | 0.141
FDA | MoRh | 09_25 | 0.42 | 37 | 0.141
FDA | MoRh | 09_25 | 0.42 | 38 | 0.142
FDA | MoRh | 09_25 | 0.42 | 39 | 0.142
FDA | MoRh | 09_25 | 0.42 | 40 | 0.142
FDA | MoRh | 09_25 | 0.42 | 41 | 0.143
FDA | MoRh | 09_25 | 0.42 | 42 | 0.143
FDA | MoRh | 09_25 | 0.42 | 43 | 0.143
FDA | MoRh | 09_25 | 0.42 | 44 | 0.144
FDA | MoRh | 09_25 | 0.42 | 45 | 0.144
FDA | MoRh | 09_25 | 0.42 | 46 | 0.145
FDA | MoRh | 09_25 | 0.42 | 47 | 0.145
FDA | MoRh | 09_25 | 0.42 | 48 | 0.145
FDA | MoRh | 09_25 | 0.42 | 49 | 0.146
FDA | MoRh | 09_25 | 0.435 | 23 | 0.14
FDA | MoRh | 09_25 | 0.435 | 24 | 0.142
FDA | MoRh | 09_25 | 0.435 | 25 | 0.144
FDA | MoRh | 09_25 | 0.435 | 26 | 0.144
FDA | MoRh | 09_25 | 0.435 | 27 | 0.145
FDA | MoRh | 09_25 | 0.435 | 28 | 0.145
FDA | MoRh | 09_25 | 0.435 | 29 | 0.145
FDA | MoRh | 09_25 | 0.435 | 30 | 0.145
FDA | MoRh | 09_25 | 0.435 | 31 | 0.145
FDA | MoRh | 09_25 | 0.435 | 32 | 0.145
FDA | MoRh | 09_25 | 0.435 | 33 | 0.145
FDA | MoRh | 09_25 | 0.435 | 34 | 0.145
FDA | MoRh | 09_25 | 0.435 | 35 | 0.145
FDA | MoRh | 09_25 | 0.435 | 36 | 0.146
FDA | MoRh | 09_25 | 0.435 | 37 | 0.146
FDA | MoRh | 09_25 | 0.435 | 38 | 0.146
FDA | MoRh | 09_25 | 0.435 | 39 | 0.147
FDA | MoRh | 09_25 | 0.435 | 40 | 0.147
FDA | MoRh | 09_25 | 0.435 | 41 | 0.147
FDA | MoRh | 09_25 | 0.435 | 42 | 0.148
FDA | MoRh | 09_25 | 0.435 | 43 | 0.148
FDA | MoRh | 09_25 | 0.435 | 44 | 0.148
FDA | MoRh | 09_25 | 0.435 | 45 | 0.149
FDA | MoRh | 09_25 | 0.435 | 46 | 0.149
FDA | MoRh | 09_25 | 0.435 | 47 | 0.149
FDA | MoRh | 09_25 | 0.435 | 48 | 0.15
FDA | MoRh | 09_25 | 0.435 | 49 | 0.15
FDA | MoRh | 09_25 | 0.45 | 23 | 0.145
FDA | MoRh | 09_25 | 0.45 | 24 | 0.147
FDA | MoRh | 09_25 | 0.45 | 25 | 0.149
FDA | MoRh | 09_25 | 0.45 | 26 | 0.15
FDA | MoRh | 09_25 | 0.45 | 27 | 0.15
FDA | MoRh | 09_25 | 0.45 | 28 | 0.15
FDA | MoRh | 09_25 | 0.45 | 29 | 0.15
FDA | MoRh | 09_25 | 0.45 | 30 | 0.15
FDA | MoRh | 09_25 | 0.45 | 31 | 0.15
FDA | MoRh | 09_25 | 0.45 | 32 | 0.15
FDA | MoRh | 09_25 | 0.45 | 33 | 0.15
FDA | MoRh | 09_25 | 0.45 | 34 | 0.15
FDA | MoRh | 09_25 | 0.45 | 35 | 0.15
FDA | MoRh | 09_25 | 0.45 | 36 | 0.15
FDA | MoRh | 09_25 | 0.45 | 37 | 0.151
FDA | MoRh | 09_25 | 0.45 | 38 | 0.151
FDA | MoRh | 09_25 | 0.45 | 39 | 0.151
FDA | MoRh | 09_25 | 0.45 | 40 | 0.152
FDA | MoRh | 09_25 | 0.45 | 41 | 0.152
FDA | MoRh | 09_25 | 0.45 | 42 | 0.152
FDA | MoRh | 09_25 | 0.45 | 43 | 0.153
FDA | MoRh | 09_25 | 0.45 | 44 | 0.153
FDA | MoRh | 09_25 | 0.45 | 45 | 0.153
FDA | MoRh | 09_25 | 0.45 | 46 | 0.154
FDA | MoRh | 09_25 | 0.45 | 47 | 0.154
FDA | MoRh | 09_25 | 0.45 | 48 | 0.154
FDA | MoRh | 09_25 | 0.45 | 49 | 0.155
FDA | MoRh | 09_25 | 0.465 | 23 | 0.151
FDA | MoRh | 09_25 | 0.465 | 24 | 0.153
FDA | MoRh | 09_25 | 0.465 | 25 | 0.154
FDA | MoRh | 09_25 | 0.465 | 26 | 0.155
FDA | MoRh | 09_25 | 0.465 | 27 | 0.155
FDA | MoRh | 09_25 | 0.465 | 28 | 0.155
FDA | MoRh | 09_25 | 0.465 | 29 | 0.155
FDA | MoRh | 09_25 | 0.465 | 30 | 0.155
FDA | MoRh | 09_25 | 0.465 | 31 | 0.155
FDA | MoRh | 09_25 | 0.465 | 32 | 0.155
FDA | MoRh | 09_25 | 0.465 | 33 | 0.155
FDA | MoRh | 09_25 | 0.465 | 34 | 0.155
FDA | MoRh | 09_25 | 0.465 | 35 | 0.155
FDA | MoRh | 09_25 | 0.465 | 36 | 0.155
FDA | MoRh | 09_25 | 0.465 | 37 | 0.156
FDA | MoRh | 09_25 | 0.465 | 38 | 0.156
FDA | MoRh | 09_25 | 0.465 | 39 | 0.156
FDA | MoRh | 09_25 | 0.465 | 40 | 0.156
FDA | MoRh | 09_25 | 0.465 | 41 | 0.157
FDA | MoRh | 09_25 | 0.465 | 42 | 0.157
FDA | MoRh | 09_25 | 0.465 | 43 | 0.157
FDA | MoRh | 09_25 | 0.465 | 44 | 0.158
FDA | MoRh | 09_25 | 0.465 | 45 | 0.158
FDA | MoRh | 09_25 | 0.465 | 46 | 0.159
FDA | MoRh | 09_25 | 0.465 | 47 | 0.159
FDA | MoRh | 09_25 | 0.465 | 48 | 0.159
FDA | MoRh | 09_25 | 0.465 | 49 | 0.16
FDA | MoRh | 09_25 | 0.48 | 25 | 0.16
FDA | MoRh | 09_25 | 0.48 | 26 | 0.16
FDA | MoRh | 09_25 | 0.48 | 27 | 0.16
FDA | MoRh | 09_25 | 0.48 | 28 | 0.16
FDA | MoRh | 09_25 | 0.48 | 29 | 0.16
FDA | MoRh | 09_25 | 0.48 | 30 | 0.16
FDA | MoRh | 09_25 | 0.48 | 31 | 0.16
FDA | MoRh | 09_25 | 0.48 | 32 | 0.16
FDA | MoRh | 09_25 | 0.48 | 33 | 0.16
FDA | MoRh | 09_25 | 0.48 | 34 | 0.16
FDA | MoRh | 09_25 | 0.48 | 35 | 0.16
FDA | MoRh | 09_25 | 0.48 | 36 | 0.16
FDA | MoRh | 09_25 | 0.48 | 37 | 0.161
FDA | MoRh | 09_25 | 0.48 | 38 | 0.161
FDA | MoRh | 09_25 | 0.48 | 39 | 0.161
FDA | MoRh | 09_25 | 0.48 | 40 | 0.161
FDA | MoRh | 09_25 | 0.48 | 41 | 0.162
FDA | MoRh | 09_25 | 0.48 | 42 | 0.162
FDA | MoRh | 09_25 | 0.48 | 43 | 0.162
FDA | MoRh | 09_25 | 0.48 | 44 | 0.163
FDA | MoRh | 09_25 | 0.48 | 45 | 0.163
FDA | MoRh | 09_25 | 0.48 | 46 | 0.163
FDA | MoRh | 09_25 | 0.48 | 47 | 0.164
FDA | MoRh | 09_25 | 0.48 | 48 | 0.164
FDA | MoRh | 09_25 | 0.48 | 49 | 0.165
FDA | MoRh | 09_25 | 0.495 | 28 | 0.166
FDA | MoRh | 09_25 | 0.495 | 29 | 0.166
FDA | MoRh | 09_25 | 0.495 | 30 | 0.166
FDA | MoRh | 09_25 | 0.495 | 31 | 0.166
FDA | MoRh | 09_25 | 0.495 | 32 | 0.166
FDA | MoRh | 09_25 | 0.495 | 33 | 0.166
FDA | MoRh | 09_25 | 0.495 | 34 | 0.166
FDA | MoRh | 09_25 | 0.495 | 35 | 0.166
FDA | MoRh | 09_25 | 0.495 | 36 | 0.166
FDA | MoRh | 09_25 | 0.495 | 37 | 0.166
FDA | MoRh | 09_25 | 0.495 | 38 | 0.166
FDA | MoRh | 09_25 | 0.495 | 39 | 0.166
FDA | MoRh | 09_25 | 0.495 | 40 | 0.167
FDA | MoRh | 09_25 | 0.495 | 41 | 0.167
FDA | MoRh | 09_25 | 0.495 | 42 | 0.167
FDA | MoRh | 09_25 | 0.495 | 43 | 0.168
FDA | MoRh | 09_25 | 0.495 | 44 | 0.168
FDA | MoRh | 09_25 | 0.495 | 45 | 0.168
FDA | MoRh | 09_25 | 0.495 | 46 | 0.169
FDA | MoRh | 09_25 | 0.495 | 47 | 0.169
FDA | MoRh | 09_25 | 0.495 | 48 | 0.169
FDA | MoRh | 09_25 | 0.495 | 49 | 0.17
FDA | MoRh | 09_25 | 0.51 | 28 | 0.171
FDA | MoRh | 09_25 | 0.51 | 29 | 0.171
FDA | MoRh | 09_25 | 0.51 | 30 | 0.171
FDA | MoRh | 09_25 | 0.51 | 31 | 0.171
FDA | MoRh | 09_25 | 0.51 | 32 | 0.171
FDA | MoRh | 09_25 | 0.51 | 33 | 0.171
FDA | MoRh | 09_25 | 0.51 | 34 | 0.171
FDA | MoRh | 09_25 | 0.51 | 35 | 0.171
FDA | MoRh | 09_25 | 0.51 | 36 | 0.171
FDA | MoRh | 09_25 | 0.51 | 37 | 0.171
FDA | MoRh | 09_25 | 0.51 | 38 | 0.171
FDA | MoRh | 09_25 | 0.51 | 39 | 0.172
FDA | MoRh | 09_25 | 0.51 | 40 | 0.172
FDA | MoRh | 09_25 | 0.51 | 41 | 0.172
FDA | MoRh | 09_25 | 0.51 | 42 | 0.172
FDA | MoRh | 09_25 | 0.51 | 43 | 0.173
FDA | MoRh | 09_25 | 0.51 | 44 | 0.173
FDA | MoRh | 09_25 | 0.51 | 45 | 0.174
FDA | MoRh | 09_25 | 0.51 | 46 | 0.174
FDA | MoRh | 09_25 | 0.51 | 47 | 0.174
FDA | MoRh | 09_25 | 0.51 | 48 | 0.175
FDA | MoRh | 09_25 | 0.51 | 49 | 0.175
FDA | MoRh | 09_25 | 0.525 | 29 | 0.177
FDA | MoRh | 09_25 | 0.525 | 30 | 0.177
FDA | MoRh | 09_25 | 0.525 | 31 | 0.177
FDA | MoRh | 09_25 | 0.525 | 32 | 0.177
FDA | MoRh | 09_25 | 0.525 | 33 | 0.177
FDA | MoRh | 09_25 | 0.525 | 34 | 0.177
FDA | MoRh | 09_25 | 0.525 | 35 | 0.177
FDA | MoRh | 09_25 | 0.525 | 36 | 0.177
FDA | MoRh | 09_25 | 0.525 | 37 | 0.177
FDA | MoRh | 09_25 | 0.525 | 38 | 0.177
FDA | MoRh | 09_25 | 0.525 | 39 | 0.177
FDA | MoRh | 09_25 | 0.525 | 40 | 0.177
FDA | MoRh | 09_25 | 0.525 | 41 | 0.178
FDA | MoRh | 09_25 | 0.525 | 42 | 0.178
FDA | MoRh | 09_25 | 0.525 | 43 | 0.178
FDA | MoRh | 09_25 | 0.525 | 44 | 0.179
FDA | MoRh | 09_25 | 0.525 | 45 | 0.179
FDA | MoRh | 09_25 | 0.525 | 46 | 0.179
FDA | MoRh | 09_25 | 0.525 | 47 | 0.18
FDA | MoRh | 09_25 | 0.525 | 48 | 0.18
FDA | MoRh | 09_25 | 0.525 | 49 | 0.181
FDA | MoRh | 09_25 | 0.54 | 32 | 0.182
FDA | MoRh | 09_25 | 0.54 | 33 | 0.182
FDA | MoRh | 09_25 | 0.54 | 34 | 0.182
FDA | MoRh | 09_25 | 0.54 | 35 | 0.182
FDA | MoRh | 09_25 | 0.54 | 36 | 0.182
FDA | MoRh | 09_25 | 0.54 | 37 | 0.182
FDA | MoRh | 09_25 | 0.54 | 38 | 0.183
FDA | MoRh | 09_25 | 0.54 | 39 | 0.183
FDA | MoRh | 09_25 | 0.54 | 40 | 0.183
FDA | MoRh | 09_25 | 0.54 | 41 | 0.183
FDA | MoRh | 09_25 | 0.54 | 42 | 0.184
FDA | MoRh | 09_25 | 0.54 | 43 | 0.184
FDA | MoRh | 09_25 | 0.54 | 44 | 0.184
FDA | MoRh | 09_25 | 0.54 | 45 | 0.185
FDA | MoRh | 09_25 | 0.54 | 46 | 0.185
FDA | MoRh | 09_25 | 0.54 | 47 | 0.186
FDA | MoRh | 09_25 | 0.54 | 48 | 0.186
FDA | MoRh | 09_25 | 0.54 | 49 | 0.187
FDA | MoRh | 09_25 | 0.555 | 34 | 0.188
FDA | MoRh | 09_25 | 0.555 | 35 | 0.188
FDA | MoRh | 09_25 | 0.555 | 36 | 0.188
FDA | MoRh | 09_25 | 0.555 | 37 | 0.188
FDA | MoRh | 09_25 | 0.555 | 38 | 0.188
FDA | MoRh | 09_25 | 0.555 | 39 | 0.188
FDA | MoRh | 09_25 | 0.555 | 40 | 0.189
FDA | MoRh | 09_25 | 0.555 | 41 | 0.189
FDA | MoRh | 09_25 | 0.555 | 42 | 0.189
FDA | MoRh | 09_25 | 0.555 | 43 | 0.19
FDA | MoRh | 09_25 | 0.555 | 44 | 0.19
FDA | MoRh | 09_25 | 0.555 | 45 | 0.19
FDA | MoRh | 09_25 | 0.555 | 46 | 0.191
FDA | MoRh | 09_25 | 0.555 | 47 | 0.191
FDA | MoRh | 09_25 | 0.555 | 48 | 0.192
FDA | MoRh | 09_25 | 0.555 | 49 | 0.193
FDA | MoRh | 09_25 | 0.57 | 43 | 0.195
FDA | MoRh | 09_25 | 0.57 | 44 | 0.196
FDA | MoRh | 09_25 | 0.57 | 45 | 0.196
FDA | MoRh | 09_25 | 0.57 | 46 | 0.197
FDA | MoRh | 09_25 | 0.57 | 47 | 0.197
FDA | MoRh | 09_25 | 0.57 | 48 | 0.198
FDA | MoRh | 09_25 | 0.57 | 49 | 0.198
FDA | MoRh | 09_25 | 0.585 | 43 | 0.201
FDA | MoRh | 09_25 | 0.585 | 44 | 0.202
FDA | MoRh | 09_25 | 0.585 | 45 | 0.202
FDA | MoRh | 09_25 | 0.585 | 46 | 0.203
FDA | MoRh | 09_25 | 0.585 | 47 | 0.203
FDA | MoRh | 09_25 | 0.585 | 48 | 0.204
FDA | MoRh | 09_25 | 0.585 | 49 | 0.204
FDA | MoRh | 09_25 | 0.6 | 43 | 0.206
FDA | MoRh | 09_25 | 0.6 | 44 | 0.207
FDA | MoRh | 09_25 | 0.6 | 45 | 0.208
FDA | MoRh | 09_25 | 0.6 | 46 | 0.208
FDA | MoRh | 09_25 | 0.6 | 47 | 0.209
FDA | MoRh | 09_25 | 0.6 | 48 | 0.209
FDA | MoRh | 09_25 | 0.6 | 49 | 0.21
FDA | RhAg | 09_25 | 0.2 | 20 | 0.046
FDA | RhAg | 09_25 | 0.2 | 21 | 0.048
FDA | RhAg | 09_25 | 0.2 | 22 | 0.05
FDA | RhAg | 09_25 | 0.2 | 23 | 0.053
FDA | RhAg | 09_25 | 0.2 | 24 | 0.056
FDA | RhAg | 09_25 | 0.2 | 25 | 0.058
FDA | RhAg | 09_25 | 0.22 | 20 | 0.053
FDA | RhAg | 09_25 | 0.22 | 21 | 0.056
FDA | RhAg | 09_25 | 0.22 | 22 | 0.058
FDA | RhAg | 09_25 | 0.22 | 23 | 0.061
FDA | RhAg | 09_25 | 0.22 | 24 | 0.064
FDA | RhAg | 09_25 | 0.22 | 25 | 0.067
FDA | RhAg | 09_25 | 0.22 | 26 | 0.07
FDA | RhAg | 09_25 | 0.22 | 27 | 0.073
FDA | RhAg | 09_25 | 0.24 | 20 | 0.06
FDA | RhAg | 09_25 | 0.24 | 21 | 0.063
FDA | RhAg | 09_25 | 0.24 | 22 | 0.066
FDA | RhAg | 09_25 | 0.24 | 23 | 0.069
FDA | RhAg | 09_25 | 0.24 | 24 | 0.072
FDA | RhAg | 09_25 | 0.24 | 25 | 0.076
FDA | RhAg | 09_25 | 0.24 | 26 | 0.079
FDA | RhAg | 09_25 | 0.24 | 27 | 0.082
FDA | RhAg | 09_25 | 0.24 | 28 | 0.084
FDA | RhAg | 09_25 | 0.24 | 29 | 0.086
FDA | RhAg | 09_25 | 0.26 | 20 | 0.067
FDA | RhAg | 09_25 | 0.26 | 21 | 0.071
FDA | RhAg | 09_25 | 0.26 | 22 | 0.074
FDA | RhAg | 09_25 | 0.26 | 23 | 0.077
FDA | RhAg | 09_25 | 0.26 | 24 | 0.08
FDA | RhAg | 09_25 | 0.26 | 25 | 0.084
FDA | RhAg | 09_25 | 0.26 | 26 | 0.087
FDA | RhAg | 09_25 | 0.26 | 27 | 0.09
FDA | RhAg | 09_25 | 0.26 | 28 | 0.092
FDA | RhAg | 09_25 | 0.26 | 29 | 0.095
FDA | RhAg | 09_25 | 0.26 | 30 | 0.096
FDA | RhAg | 09_25 | 0.26 | 31 | 0.098
FDA | RhAg | 09_25 | 0.26 | 32 | 0.1
FDA | RhAg | 09_25 | 0.28 | 20 | 0.074
FDA | RhAg | 09_25 | 0.28 | 21 | 0.078
FDA | RhAg | 09_25 | 0.28 | 22 | 0.081
FDA | RhAg | 09_25 | 0.28 | 23 | 0.084
FDA | RhAg | 09_25 | 0.28 | 24 | 0.088
FDA | RhAg | 09_25 | 0.28 | 25 | 0.092
FDA | RhAg | 09_25 | 0.28 | 26 | 0.095
FDA | RhAg | 09_25 | 0.28 | 27 | 0.098
FDA | RhAg | 09_25 | 0.28 | 28 | 0.101
FDA | RhAg | 09_25 | 0.28 | 29 | 0.103
FDA | RhAg | 09_25 | 0.28 | 30 | 0.105
FDA | RhAg | 09_25 | 0.28 | 31 | 0.106
FDA | RhAg | 09_25 | 0.28 | 32 | 0.108
FDA | RhAg | 09_25 | 0.28 | 33 | 0.109
FDA | RhAg | 09_25 | 0.28 | 34 | 0.11
FDA | RhAg | 09_25 | 0.3 | 20 | 0.081
FDA | RhAg | 09_25 | 0.3 | 21 | 0.085
FDA | RhAg | 09_25 | 0.3 | 22 | 0.088
FDA | RhAg | 09_25 | 0.3 | 23 | 0.092
FDA | RhAg | 09_25 | 0.3 | 24 | 0.096
FDA | RhAg | 09_25 | 0.3 | 25 | 0.1
FDA | RhAg | 09_25 | 0.3 | 26 | 0.103
FDA | RhAg | 09_25 | 0.3 | 27 | 0.106
FDA | RhAg | 09_25 | 0.3 | 28 | 0.109
FDA | RhAg | 09_25 | 0.3 | 29 | 0.111
FDA | RhAg | 09_25 | 0.3 | 30 | 0.112
FDA | RhAg | 09_25 | 0.3 | 31 | 0.114
FDA | RhAg | 09_25 | 0.3 | 32 | 0.115
FDA | RhAg | 09_25 | 0.3 | 33 | 0.117
FDA | RhAg | 09_25 | 0.3 | 34 | 0.118
FDA | RhAg | 09_25 | 0.3 | 35 | 0.119
FDA | RhAg | 09_25 | 0.3 | 36 | 0.12
FDA | RhAg | 09_25 | 0.3 | 37 | 0.121
FDA | RhAg | 09_25 | 0.32 | 20 | 0.088
FDA | RhAg | 09_25 | 0.32 | 21 | 0.093
FDA | RhAg | 09_25 | 0.32 | 22 | 0.096
FDA | RhAg | 09_25 | 0.32 | 23 | 0.099
FDA | RhAg | 09_25 | 0.32 | 24 | 0.104
FDA | RhAg | 09_25 | 0.32 | 25 | 0.107
FDA | RhAg | 09_25 | 0.32 | 26 | 0.111
FDA | RhAg | 09_25 | 0.32 | 27 | 0.114
FDA | RhAg | 09_25 | 0.32 | 28 | 0.116
FDA | RhAg | 09_25 | 0.32 | 29 | 0.118
FDA | RhAg | 09_25 | 0.32 | 30 | 0.12
FDA | RhAg | 09_25 | 0.32 | 31 | 0.121
FDA | RhAg | 09_25 | 0.32 | 32 | 0.123
FDA | RhAg | 09_25 | 0.32 | 33 | 0.124
FDA | RhAg | 09_25 | 0.32 | 34 | 0.125
FDA | RhAg | 09_25 | 0.32 | 35 | 0.127
FDA | RhAg | 09_25 | 0.32 | 36 | 0.128
FDA | RhAg | 09_25 | 0.32 | 37 | 0.129
FDA | RhAg | 09_25 | 0.32 | 38 | 0.13
FDA | RhAg | 09_25 | 0.32 | 39 | 0.131
FDA | RhAg | 09_25 | 0.32 | 40 | 0.132
FDA | RhAg | 09_25 | 0.32 | 41 | 0.133
FDA | RhAg | 09_25 | 0.32 | 42 | 0.133
FDA | RhAg | 09_25 | 0.32 | 43 | 0.134
FDA | RhAg | 09_25 | 0.34 | 20 | 0.096
FDA | RhAg | 09_25 | 0.34 | 21 | 0.1
FDA | RhAg | 09_25 | 0.34 | 22 | 0.103
FDA | RhAg | 09_25 | 0.34 | 23 | 0.107
FDA | RhAg | 09_25 | 0.34 | 24 | 0.111
FDA | RhAg | 09_25 | 0.34 | 25 | 0.115
FDA | RhAg | 09_25 | 0.34 | 26 | 0.119
FDA | RhAg | 09_25 | 0.34 | 27 | 0.121
FDA | RhAg | 09_25 | 0.34 | 28 | 0.124
FDA | RhAg | 09_25 | 0.34 | 29 | 0.126
FDA | RhAg | 09_25 | 0.34 | 30 | 0.127
FDA | RhAg | 09_25 | 0.34 | 31 | 0.129
FDA | RhAg | 09_25 | 0.34 | 32 | 0.13
FDA | RhAg | 09_25 | 0.34 | 33 | 0.131
FDA | RhAg | 09_25 | 0.34 | 34 | 0.132
FDA | RhAg | 09_25 | 0.34 | 35 | 0.134
FDA | RhAg | 09_25 | 0.34 | 36 | 0.135
FDA | RhAg | 09_25 | 0.34 | 37 | 0.136
FDA | RhAg | 09_25 | 0.34 | 38 | 0.137
FDA | RhAg | 09_25 | 0.34 | 39 | 0.138
FDA | RhAg | 09_25 | 0.34 | 40 | 0.139
FDA | RhAg | 09_25 | 0.34 | 41 | 0.139
FDA | RhAg | 09_25 | 0.34 | 42 | 0.14
FDA | RhAg | 09_25 | 0.34 | 43 | 0.141
FDA | RhAg | 09_25 | 0.34 | 44 | 0.142
FDA | RhAg | 09_25 | 0.34 | 45 | 0.142
FDA | RhAg | 09_25 | 0.34 | 46 | 0.143
FDA | RhAg | 09_25 | 0.36 | 20 | 0.103
FDA | RhAg | 09_25 | 0.36 | 21 | 0.107
FDA | RhAg | 09_25 | 0.36 | 22 | 0.11
FDA | RhAg | 09_25 | 0.36 | 23 | 0.114
FDA | RhAg | 09_25 | 0.36 | 24 | 0.119
FDA | RhAg | 09_25 | 0.36 | 25 | 0.122
FDA | RhAg | 09_25 | 0.36 | 26 | 0.126
FDA | RhAg | 09_25 | 0.36 | 27 | 0.129
FDA | RhAg | 09_25 | 0.36 | 28 | 0.131
FDA | RhAg | 09_25 | 0.36 | 29 | 0.133
FDA | RhAg | 09_25 | 0.36 | 30 | 0.134
FDA | RhAg | 09_25 | 0.36 | 31 | 0.136
FDA | RhAg | 09_25 | 0.36 | 32 | 0.137
FDA | RhAg | 09_25 | 0.36 | 33 | 0.138
FDA | RhAg | 09_25 | 0.36 | 34 | 0.139
FDA | RhAg | 09_25 | 0.36 | 35 | 0.14
FDA | RhAg | 09_25 | 0.36 | 36 | 0.141
FDA | RhAg | 09_25 | 0.36 | 37 | 0.142
FDA | RhAg | 09_25 | 0.36 | 38 | 0.143
FDA | RhAg | 09_25 | 0.36 | 39 | 0.144
FDA | RhAg | 09_25 | 0.36 | 40 | 0.145
FDA | RhAg | 09_25 | 0.36 | 41 | 0.146
FDA | RhAg | 09_25 | 0.36 | 42 | 0.147
FDA | RhAg | 09_25 | 0.36 | 43 | 0.148
FDA | RhAg | 09_25 | 0.36 | 44 | 0.148
FDA | RhAg | 09_25 | 0.36 | 45 | 0.149
FDA | RhAg | 09_25 | 0.36 | 46 | 0.15
FDA | RhAg | 09_25 | 0.36 | 47 | 0.15
FDA | RhAg | 09_25 | 0.36 | 48 | 0.151
FDA | RhAg | 09_25 | 0.36 | 49 | 0.151
FDA | RhAg | 09_25 | 0.38 | 20 | 0.11
FDA | RhAg | 09_25 | 0.38 | 21 | 0.115
FDA | RhAg | 09_25 | 0.38 | 22 | 0.118
FDA | RhAg | 09_25 | 0.38 | 23 | 0.122
FDA | RhAg | 09_25 | 0.38 | 24 | 0.126
FDA | RhAg | 09_25 | 0.38 | 25 | 0.13
FDA | RhAg | 09_25 | 0.38 | 26 | 0.133
FDA | RhAg | 09_25 | 0.38 | 27 | 0.136
FDA | RhAg | 09_25 | 0.38 | 28 | 0.138
FDA | RhAg | 09_25 | 0.38 | 29 | 0.14
FDA | RhAg | 09_25 | 0.38 | 30 | 0.141
FDA | RhAg | 09_25 | 0.38 | 31 | 0.143
FDA | RhAg | 09_25 | 0.38 | 32 | 0.144
FDA | RhAg | 09_25 | 0.38 | 33 | 0.145
FDA | RhAg | 09_25 | 0.38 | 34 | 0.146
FDA | RhAg | 09_25 | 0.38 | 35 | 0.147
FDA | RhAg | 09_25 | 0.38 | 36 | 0.148
FDA | RhAg | 09_25 | 0.38 | 37 | 0.149
FDA | RhAg | 09_25 | 0.38 | 38 | 0.15
FDA | RhAg | 09_25 | 0.38 | 39 | 0.151
FDA | RhAg | 09_25 | 0.38 | 40 | 0.152
FDA | RhAg | 09_25 | 0.38 | 41 | 0.152
FDA | RhAg | 09_25 | 0.38 | 42 | 0.153
FDA | RhAg | 09_25 | 0.38 | 43 | 0.154
FDA | RhAg | 09_25 | 0.38 | 44 | 0.155
FDA | RhAg | 09_25 | 0.38 | 45 | 0.155
FDA | RhAg | 09_25 | 0.38 | 46 | 0.156
FDA | RhAg | 09_25 | 0.38 | 47 | 0.156
FDA | RhAg | 09_25 | 0.38 | 48 | 0.157
FDA | RhAg | 09_25 | 0.38 | 49 | 0.158
FDA | RhAg | 09_25 | 0.4 | 20 | 0.117
FDA | RhAg | 09_25 | 0.4 | 21 | 0.122
FDA | RhAg | 09_25 | 0.4 | 22 | 0.125
FDA | RhAg | 09_25 | 0.4 | 23 | 0.129
FDA | RhAg | 09_25 | 0.4 | 24 | 0.133
FDA | RhAg | 09_25 | 0.4 | 25 | 0.137
FDA | RhAg | 09_25 | 0.4 | 26 | 0.14
FDA | RhAg | 09_25 | 0.4 | 27 | 0.143
FDA | RhAg | 09_25 | 0.4 | 28 | 0.145
FDA | RhAg | 09_25 | 0.4 | 29 | 0.147
FDA | RhAg | 09_25 | 0.4 | 30 | 0.148
FDA | RhAg | 09_25 | 0.4 | 31 | 0.149
FDA | RhAg | 09_25 | 0.4 | 32 | 0.15
FDA | RhAg | 09_25 | 0.4 | 33 | 0.151
FDA | RhAg | 09_25 | 0.4 | 34 | 0.152
FDA | RhAg | 09_25 | 0.4 | 35 | 0.153
FDA | RhAg | 09_25 | 0.4 | 36 | 0.154
FDA | RhAg | 09_25 | 0.4 | 37 | 0.155
FDA | RhAg | 09_25 | 0.4 | 38 | 0.156
FDA | RhAg | 09_25 | 0.4 | 39 | 0.157
FDA | RhAg | 09_25 | 0.4 | 40 | 0.158
FDA | RhAg | 09_25 | 0.4 | 41 | 0.159
FDA | RhAg | 09_25 | 0.4 | 42 | 0.159
FDA | RhAg | 09_25 | 0.4 | 43 | 0.16
FDA | RhAg | 09_25 | 0.4 | 44 | 0.161
FDA | RhAg | 09_25 | 0.4 | 45 | 0.161
FDA | RhAg | 09_25 | 0.4 | 46 | 0.162
FDA | RhAg | 09_25 | 0.4 | 47 | 0.163
FDA | RhAg | 09_25 | 0.4 | 48 | 0.163
FDA | RhAg | 09_25 | 0.4 | 49 | 0.164
FDA | RhAg | 09_25 | 0.42 | 21 | 0.129
FDA | RhAg | 09_25 | 0.42 | 22 | 0.133
FDA | RhAg | 09_25 | 0.42 | 23 | 0.136
FDA | RhAg | 09_25 | 0.42 | 24 | 0.141
FDA | RhAg | 09_25 | 0.42 | 25 | 0.144
FDA | RhAg | 09_25 | 0.42 | 26 | 0.148
FDA | RhAg | 09_25 | 0.42 | 27 | 0.15
FDA | RhAg | 09_25 | 0.42 | 28 | 0.152
FDA | RhAg | 09_25 | 0.42 | 29 | 0.154
FDA | RhAg | 09_25 | 0.42 | 30 | 0.155
FDA | RhAg | 09_25 | 0.42 | 31 | 0.156
FDA | RhAg | 09_25 | 0.42 | 32 | 0.157
FDA | RhAg | 09_25 | 0.42 | 33 | 0.158
FDA | RhAg | 09_25 | 0.42 | 34 | 0.159
FDA | RhAg | 09_25 | 0.42 | 35 | 0.16
FDA | RhAg | 09_25 | 0.42 | 36 | 0.161
FDA | RhAg | 09_25 | 0.42 | 37 | 0.161
FDA | RhAg | 09_25 | 0.42 | 38 | 0.162
FDA | RhAg | 09_25 | 0.42 | 39 | 0.163
FDA | RhAg | 09_25 | 0.42 | 40 | 0.164
FDA | RhAg | 09_25 | 0.42 | 41 | 0.165
FDA | RhAg | 09_25 | 0.42 | 42 | 0.165
FDA | RhAg | 09_25 | 0.42 | 43 | 0.166
FDA | RhAg | 09_25 | 0.42 | 44 | 0.167
FDA | RhAg | 09_25 | 0.42 | 45 | 0.167
FDA | RhAg | 09_25 | 0.42 | 46 | 0.168
FDA | RhAg | 09_25 | 0.42 | 47 | 0.169
FDA | RhAg | 09_25 | 0.42 | 48 | 0.169
FDA | RhAg | 09_25 | 0.42 | 49 | 0.17
FDA | RhAg | 09_25 | 0.44 | 21 | 0.137
FDA | RhAg | 09_25 | 0.44 | 22 | 0.14
FDA | RhAg | 09_25 | 0.44 | 23 | 0.144
FDA | RhAg | 09_25 | 0.44 | 24 | 0.148
FDA | RhAg | 09_25 | 0.44 | 25 | 0.151
FDA | RhAg | 09_25 | 0.44 | 26 | 0.155
FDA | RhAg | 09_25 | 0.44 | 27 | 0.157
FDA | RhAg | 09_25 | 0.44 | 28 | 0.159
FDA | RhAg | 09_25 | 0.44 | 29 | 0.16
FDA | RhAg | 09_25 | 0.44 | 30 | 0.161
FDA | RhAg | 09_25 | 0.44 | 31 | 0.163
FDA | RhAg | 09_25 | 0.44 | 32 | 0.163
FDA | RhAg | 09_25 | 0.44 | 33 | 0.164
FDA | RhAg | 09_25 | 0.44 | 34 | 0.165
FDA | RhAg | 09_25 | 0.44 | 35 | 0.166
FDA | RhAg | 09_25 | 0.44 | 36 | 0.167
FDA | RhAg | 09_25 | 0.44 | 37 | 0.168
FDA | RhAg | 09_25 | 0.44 | 38 | 0.168
FDA | RhAg | 09_25 | 0.44 | 39 | 0.169
FDA | RhAg | 09_25 | 0.44 | 40 | 0.17
FDA | RhAg | 09_25 | 0.44 | 41 | 0.171
FDA | RhAg | 09_25 | 0.44 | 42 | 0.171
FDA | RhAg | 09_25 | 0.44 | 43 | 0.172
FDA | RhAg | 09_25 | 0.44 | 44 | 0.173
FDA | RhAg | 09_25 | 0.44 | 45 | 0.173
FDA | RhAg | 09_25 | 0.44 | 46 | 0.174
FDA | RhAg | 09_25 | 0.44 | 47 | 0.174
FDA | RhAg | 09_25 | 0.44 | 48 | 0.175
FDA | RhAg | 09_25 | 0.44 | 49 | 0.175
FDA | RhAg | 09_25 | 0.46 | 22 | 0.147
FDA | RhAg | 09_25 | 0.46 | 23 | 0.151
FDA | RhAg | 09_25 | 0.46 | 24 | 0.155
FDA | RhAg | 09_25 | 0.46 | 25 | 0.158
FDA | RhAg | 09_25 | 0.46 | 26 | 0.161
FDA | RhAg | 09_25 | 0.46 | 27 | 0.164
FDA | RhAg | 09_25 | 0.46 | 28 | 0.166
FDA | RhAg | 09_25 | 0.46 | 29 | 0.167
FDA | RhAg | 09_25 | 0.46 | 30 | 0.168
FDA | RhAg | 09_25 | 0.46 | 31 | 0.169
FDA | RhAg | 09_25 | 0.46 | 32 | 0.17
FDA | RhAg | 09_25 | 0.46 | 33 | 0.171
FDA | RhAg | 09_25 | 0.46 | 34 | 0.171
FDA | RhAg | 09_25 | 0.46 | 35 | 0.172
FDA | RhAg | 09_25 | 0.46 | 36 | 0.173
FDA | RhAg | 09_25 | 0.46 | 37 | 0.173
FDA | RhAg | 09_25 | 0.46 | 38 | 0.174
FDA | RhAg | 09_25 | 0.46 | 39 | 0.175
FDA | RhAg | 09_25 | 0.46 | 40 | 0.176
FDA | RhAg | 09_25 | 0.46 | 41 | 0.176
FDA | RhAg | 09_25 | 0.46 | 42 | 0.177
FDA | RhAg | 09_25 | 0.46 | 43 | 0.178
FDA | RhAg | 09_25 | 0.46 | 44 | 0.178
FDA | RhAg | 09_25 | 0.46 | 45 | 0.179
FDA | RhAg | 09_25 | 0.46 | 46 | 0.18
FDA | RhAg | 09_25 | 0.46 | 47 | 0.18
FDA | RhAg | 09_25 | 0.46 | 48 | 0.181
FDA | RhAg | 09_25 | 0.46 | 49 | 0.181
FDA | RhAg | 09_25 | 0.48 | 22 | 0.155
FDA | RhAg | 09_25 | 0.48 | 23 | 0.158
FDA | RhAg | 09_25 | 0.48 | 24 | 0.162
FDA | RhAg | 09_25 | 0.48 | 25 | 0.165
FDA | RhAg | 09_25 | 0.48 | 26 | 0.168
FDA | RhAg | 09_25 | 0.48 | 27 | 0.171
FDA | RhAg | 09_25 | 0.48 | 28 | 0.172
FDA | RhAg | 09_25 | 0.48 | 29 | 0.174
FDA | RhAg | 09_25 | 0.48 | 30 | 0.175
FDA | RhAg | 09_25 | 0.48 | 31 | 0.175
FDA | RhAg | 09_25 | 0.48 | 32 | 0.176
FDA | RhAg | 09_25 | 0.48 | 33 | 0.177
FDA | RhAg | 09_25 | 0.48 | 34 | 0.177
FDA | RhAg | 09_25 | 0.48 | 35 | 0.178
FDA | RhAg | 09_25 | 0.48 | 36 | 0.179
FDA | RhAg | 09_25 | 0.48 | 37 | 0.18
FDA | RhAg | 09_25 | 0.48 | 38 | 0.18
FDA | RhAg | 09_25 | 0.48 | 39 | 0.181
FDA | RhAg | 09_25 | 0.48 | 40 | 0.181
FDA | RhAg | 09_25 | 0.48 | 41 | 0.182
FDA | RhAg | 09_25 | 0.48 | 42 | 0.183
FDA | RhAg | 09_25 | 0.48 | 43 | 0.183
FDA | RhAg | 09_25 | 0.48 | 44 | 0.184
FDA | RhAg | 09_25 | 0.48 | 45 | 0.185
FDA | RhAg | 09_25 | 0.48 | 46 | 0.185
FDA | RhAg | 09_25 | 0.48 | 47 | 0.186
FDA | RhAg | 09_25 | 0.48 | 48 | 0.186
FDA | RhAg | 09_25 | 0.48 | 49 | 0.187
FDA | RhAg | 09_25 | 0.5 | 23 | 0.166
FDA | RhAg | 09_25 | 0.5 | 24 | 0.169
FDA | RhAg | 09_25 | 0.5 | 25 | 0.172
FDA | RhAg | 09_25 | 0.5 | 26 | 0.175
FDA | RhAg | 09_25 | 0.5 | 27 | 0.177
FDA | RhAg | 09_25 | 0.5 | 28 | 0.179
FDA | RhAg | 09_25 | 0.5 | 29 | 0.18
FDA | RhAg | 09_25 | 0.5 | 30 | 0.181
FDA | RhAg | 09_25 | 0.5 | 31 | 0.182
FDA | RhAg | 09_25 | 0.5 | 32 | 0.182
FDA | RhAg | 09_25 | 0.5 | 33 | 0.183
FDA | RhAg | 09_25 | 0.5 | 34 | 0.183
FDA | RhAg | 09_25 | 0.5 | 35 | 0.184
FDA | RhAg | 09_25 | 0.5 | 36 | 0.185
FDA | RhAg | 09_25 | 0.5 | 37 | 0.185
FDA | RhAg | 09_25 | 0.5 | 38 | 0.186
FDA | RhAg | 09_25 | 0.5 | 39 | 0.187
FDA | RhAg | 09_25 | 0.5 | 40 | 0.187
FDA | RhAg | 09_25 | 0.5 | 41 | 0.188
FDA | RhAg | 09_25 | 0.5 | 42 | 0.189
FDA | RhAg | 09_25 | 0.5 | 43 | 0.189
FDA | RhAg | 09_25 | 0.5 | 44 | 0.19
FDA | RhAg | 09_25 | 0.5 | 45 | 0.19
FDA | RhAg | 09_25 | 0.5 | 46 | 0.191
FDA | RhAg | 09_25 | 0.5 | 47 | 0.191
FDA | RhAg | 09_25 | 0.5 | 48 | 0.192
FDA | RhAg | 09_25 | 0.5 | 49 | 0.192
FDA | RhAg | 09_25 | 0.52 | 23 | 0.173
FDA | RhAg | 09_25 | 0.52 | 24 | 0.176
FDA | RhAg | 09_25 | 0.52 | 25 | 0.179
FDA | RhAg | 09_25 | 0.52 | 26 | 0.182
FDA | RhAg | 09_25 | 0.52 | 27 | 0.184
FDA | RhAg | 09_25 | 0.52 | 28 | 0.186
FDA | RhAg | 09_25 | 0.52 | 29 | 0.187
FDA | RhAg | 09_25 | 0.52 | 30 | 0.187
FDA | RhAg | 09_25 | 0.52 | 31 | 0.188
FDA | RhAg | 09_25 | 0.52 | 32 | 0.188
FDA | RhAg | 09_25 | 0.52 | 33 | 0.189
FDA | RhAg | 09_25 | 0.52 | 34 | 0.19
FDA | RhAg | 09_25 | 0.52 | 35 | 0.19
FDA | RhAg | 09_25 | 0.52 | 36 | 0.191
FDA | RhAg | 09_25 | 0.52 | 37 | 0.191
FDA | RhAg | 09_25 | 0.52 | 38 | 0.192
FDA | RhAg | 09_25 | 0.52 | 39 | 0.192
FDA | RhAg | 09_25 | 0.52 | 40 | 0.193
FDA | RhAg | 09_25 | 0.52 | 41 | 0.193
FDA | RhAg | 09_25 | 0.52 | 42 | 0.194
FDA | RhAg | 09_25 | 0.52 | 43 | 0.195
FDA | RhAg | 09_25 | 0.52 | 44 | 0.195
FDA | RhAg | 09_25 | 0.52 | 45 | 0.196
FDA | RhAg | 09_25 | 0.52 | 46 | 0.196
FDA | RhAg | 09_25 | 0.52 | 47 | 0.197
FDA | RhAg | 09_25 | 0.52 | 48 | 0.197
FDA | RhAg | 09_25 | 0.52 | 49 | 0.198
FDA | RhAg | 09_25 | 0.54 | 24 | 0.183
FDA | RhAg | 09_25 | 0.54 | 25 | 0.186
FDA | RhAg | 09_25 | 0.54 | 26 | 0.189
FDA | RhAg | 09_25 | 0.54 | 27 | 0.191
FDA | RhAg | 09_25 | 0.54 | 28 | 0.192
FDA | RhAg | 09_25 | 0.54 | 29 | 0.193
FDA | RhAg | 09_25 | 0.54 | 30 | 0.194
FDA | RhAg | 09_25 | 0.54 | 31 | 0.194
FDA | RhAg | 09_25 | 0.54 | 32 | 0.195
FDA | RhAg | 09_25 | 0.54 | 33 | 0.195
FDA | RhAg | 09_25 | 0.54 | 34 | 0.196
FDA | RhAg | 09_25 | 0.54 | 35 | 0.196
FDA | RhAg | 09_25 | 0.54 | 36 | 0.196
FDA | RhAg | 09_25 | 0.54 | 37 | 0.197
FDA | RhAg | 09_25 | 0.54 | 38 | 0.197
FDA | RhAg | 09_25 | 0.54 | 39 | 0.198
FDA | RhAg | 09_25 | 0.54 | 40 | 0.199
FDA | RhAg | 09_25 | 0.54 | 41 | 0.199
FDA | RhAg | 09_25 | 0.54 | 42 | 0.2
FDA | RhAg | 09_25 | 0.54 | 43 | 0.2
FDA | RhAg | 09_25 | 0.54 | 44 | 0.201
FDA | RhAg | 09_25 | 0.54 | 45 | 0.201
FDA | RhAg | 09_25 | 0.54 | 46 | 0.202
FDA | RhAg | 09_25 | 0.54 | 47 | 0.202
FDA | RhAg | 09_25 | 0.54 | 48 | 0.203
FDA | RhAg | 09_25 | 0.54 | 49 | 0.203
FDA | RhAg | 09_25 | 0.56 | 26 | 0.195
FDA | RhAg | 09_25 | 0.56 | 27 | 0.197
FDA | RhAg | 09_25 | 0.56 | 28 | 0.198
FDA | RhAg | 09_25 | 0.56 | 29 | 0.199
FDA | RhAg | 09_25 | 0.56 | 30 | 0.2
FDA | RhAg | 09_25 | 0.56 | 31 | 0.2
FDA | RhAg | 09_25 | 0.56 | 32 | 0.201
FDA | RhAg | 09_25 | 0.56 | 33 | 0.201
FDA | RhAg | 09_25 | 0.56 | 34 | 0.201
FDA | RhAg | 09_25 | 0.56 | 35 | 0.202
FDA | RhAg | 09_25 | 0.56 | 36 | 0.202
FDA | RhAg | 09_25 | 0.56 | 37 | 0.203
FDA | RhAg | 09_25 | 0.56 | 38 | 0.203
FDA | RhAg | 09_25 | 0.56 | 39 | 0.204
FDA | RhAg | 09_25 | 0.56 | 40 | 0.204
FDA | RhAg | 09_25 | 0.56 | 41 | 0.205
FDA | RhAg | 09_25 | 0.56 | 42 | 0.205
FDA | RhAg | 09_25 | 0.56 | 43 | 0.206
FDA | RhAg | 09_25 | 0.56 | 44 | 0.206
FDA | RhAg | 09_25 | 0.56 | 45 | 0.207
FDA | RhAg | 09_25 | 0.56 | 46 | 0.207
FDA | RhAg | 09_25 | 0.56 | 47 | 0.208
FDA | RhAg | 09_25 | 0.56 | 48 | 0.208
FDA | RhAg | 09_25 | 0.56 | 49 | 0.209
FDA | RhAg | 09_25 | 0.58 | 28 | 0.205
FDA | RhAg | 09_25 | 0.58 | 29 | 0.206
FDA | RhAg | 09_25 | 0.58 | 30 | 0.206
FDA | RhAg | 09_25 | 0.58 | 31 | 0.206
FDA | RhAg | 09_25 | 0.58 | 32 | 0.207
FDA | RhAg | 09_25 | 0.58 | 33 | 0.207
FDA | RhAg | 09_25 | 0.58 | 34 | 0.207
FDA | RhAg | 09_25 | 0.58 | 35 | 0.208
FDA | RhAg | 09_25 | 0.58 | 36 | 0.208
FDA | RhAg | 09_25 | 0.58 | 37 | 0.208
FDA | RhAg | 09_25 | 0.58 | 38 | 0.209
FDA | RhAg | 09_25 | 0.58 | 39 | 0.209
FDA | RhAg | 09_25 | 0.58 | 40 | 0.21
FDA | RhAg | 09_25 | 0.58 | 41 | 0.21
FDA | RhAg | 09_25 | 0.58 | 42 | 0.211
FDA | RhAg | 09_25 | 0.58 | 43 | 0.211
FDA | RhAg | 09_25 | 0.58 | 44 | 0.212
FDA | RhAg | 09_25 | 0.58 | 45 | 0.212
FDA | RhAg | 09_25 | 0.58 | 46 | 0.213
FDA | RhAg | 09_25 | 0.58 | 47 | 0.213
FDA | RhAg | 09_25 | 0.58 | 48 | 0.214
FDA | RhAg | 09_25 | 0.58 | 49 | 0.214
FDA | RhAg | 09_25 | 0.6 | 30 | 0.212
FDA | RhAg | 09_25 | 0.6 | 31 | 0.213
FDA | RhAg | 09_25 | 0.6 | 32 | 0.213
FDA | RhAg | 09_25 | 0.6 | 33 | 0.213
FDA | RhAg | 09_25 | 0.6 | 34 | 0.213
FDA | RhAg | 09_25 | 0.6 | 35 | 0.214
FDA | RhAg | 09_25 | 0.6 | 36 | 0.214
FDA | RhAg | 09_25 | 0.6 | 37 | 0.214
FDA | RhAg | 09_25 | 0.6 | 38 | 0.215
FDA | RhAg | 09_25 | 0.6 | 39 | 0.215
FDA | RhAg | 09_25 | 0.6 | 40 | 0.215
FDA | RhAg | 09_25 | 0.6 | 41 | 0.216
FDA | RhAg | 09_25 | 0.6 | 42 | 0.216
FDA | RhAg | 09_25 | 0.6 | 43 | 0.217
FDA | RhAg | 09_25 | 0.6 | 44 | 0.217
FDA | RhAg | 09_25 | 0.6 | 45 | 0.218
FDA | RhAg | 09_25 | 0.6 | 46 | 0.218
FDA | RhAg | 09_25 | 0.6 | 47 | 0.218
FDA | RhAg | 09_25 | 0.6 | 48 | 0.219
FDA | RhAg | 09_25 | 0.6 | 49 | 0.219
FDA | RhAg | 09_25 | 0.62 | 33 | 0.219
FDA | RhAg | 09_25 | 0.62 | 34 | 0.219
FDA | RhAg | 09_25 | 0.62 | 35 | 0.219
FDA | RhAg | 09_25 | 0.62 | 36 | 0.22
FDA | RhAg | 09_25 | 0.62 | 37 | 0.22
FDA | RhAg | 09_25 | 0.62 | 38 | 0.22
FDA | RhAg | 09_25 | 0.62 | 39 | 0.221
FDA | RhAg | 09_25 | 0.62 | 40 | 0.221
FDA | RhAg | 09_25 | 0.62 | 41 | 0.221
FDA | RhAg | 09_25 | 0.62 | 42 | 0.222
FDA | RhAg | 09_25 | 0.62 | 43 | 0.222
FDA | RhAg | 09_25 | 0.62 | 44 | 0.222
FDA | RhAg | 09_25 | 0.62 | 45 | 0.223
FDA | RhAg | 09_25 | 0.62 | 46 | 0.223
FDA | RhAg | 09_25 | 0.62 | 47 | 0.224
FDA | RhAg | 09_25 | 0.62 | 48 | 0.224
FDA | RhAg | 09_25 | 0.62 | 49 | 0.225
FDA | RhAg | 09_25 | 0.64 | 35 | 0.225
FDA | RhAg | 09_25 | 0.64 | 36 | 0.225
FDA | RhAg | 09_25 | 0.64 | 37 | 0.226
FDA | RhAg | 09_25 | 0.64 | 38 | 0.226
FDA | RhAg | 09_25 | 0.64 | 39 | 0.226
FDA | RhAg | 09_25 | 0.64 | 40 | 0.227
FDA | RhAg | 09_25 | 0.64 | 41 | 0.227
FDA | RhAg | 09_25 | 0.64 | 42 | 0.227
FDA | RhAg | 09_25 | 0.64 | 43 | 0.228
FDA | RhAg | 09_25 | 0.64 | 44 | 0.228
FDA | RhAg | 09_25 | 0.64 | 45 | 0.228
FDA | RhAg | 09_25 | 0.64 | 46 | 0.229
FDA | RhAg | 09_25 | 0.64 | 47 | 0.229
FDA | RhAg | 09_25 | 0.64 | 48 | 0.23
FDA | RhAg | 09_25 | 0.64 | 49 | 0.23
FDA | RhAg | 09_25 | 0.66 | 38 | 0.231
FDA | RhAg | 09_25 | 0.66 | 39 | 0.232
FDA | RhAg | 09_25 | 0.66 | 40 | 0.232
FDA | RhAg | 09_25 | 0.66 | 41 | 0.232
FDA | RhAg | 09_25 | 0.66 | 42 | 0.233
FDA | RhAg | 09_25 | 0.66 | 43 | 0.233
FDA | RhAg | 09_25 | 0.66 | 44 | 0.233
FDA | RhAg | 09_25 | 0.66 | 45 | 0.234
FDA | RhAg | 09_25 | 0.66 | 46 | 0.234
FDA | RhAg | 09_25 | 0.66 | 47 | 0.235
FDA | RhAg | 09_25 | 0.66 | 48 | 0.235
FDA | RhAg | 09_25 | 0.66 | 49 | 0.235
FDA | RhAg | 09_25 | 0.68 | 44 | 0.239
FDA | RhAg | 09_25 | 0.68 | 45 | 0.239
FDA | RhAg | 09_25 | 0.68 | 46 | 0.24
FDA | RhAg | 09_25 | 0.68 | 47 | 0.24
FDA | RhAg | 09_25 | 0.68 | 48 | 0.24
FDA | RhAg | 09_25 | 0.68 | 49 | 0.241
FDA | RhAg | 09_25 | 0.7 | 47 | 0.246
FDA | RhAg | 09_25 | 0.7 | 48 | 0.246
FDA | RhAg | 09_25 | 0.7 | 49 | 0.246
FDA | RhRh | 09_25 | 0.2 | 20 | 0.046
FDA | RhRh | 09_25 | 0.2 | 21 | 0.048
FDA | RhRh | 09_25 | 0.2 | 22 | 0.05
FDA | RhRh | 09_25 | 0.2 | 23 | 0.053
FDA | RhRh | 09_25 | 0.2 | 24 | 0.056
FDA | RhRh | 09_25 | 0.2 | 25 | 0.058
FDA | RhRh | 09_25 | 0.22 | 20 | 0.053
FDA | RhRh | 09_25 | 0.22 | 21 | 0.056
FDA | RhRh | 09_25 | 0.22 | 22 | 0.058
FDA | RhRh | 09_25 | 0.22 | 23 | 0.061
FDA | RhRh | 09_25 | 0.22 | 24 | 0.064
FDA | RhRh | 09_25 | 0.22 | 25 | 0.067
FDA | RhRh | 09_25 | 0.22 | 26 | 0.07
FDA | RhRh | 09_25 | 0.22 | 27 | 0.072
FDA | RhRh | 09_25 | 0.24 | 20 | 0.06
FDA | RhRh | 09_25 | 0.24 | 21 | 0.063
FDA | RhRh | 09_25 | 0.24 | 22 | 0.066
FDA | RhRh | 09_25 | 0.24 | 23 | 0.069
FDA | RhRh | 09_25 | 0.24 | 24 | 0.072
FDA | RhRh | 09_25 | 0.24 | 25 | 0.076
FDA | RhRh | 09_25 | 0.24 | 26 | 0.078
FDA | RhRh | 09_25 | 0.24 | 27 | 0.081
FDA | RhRh | 09_25 | 0.24 | 28 | 0.083
FDA | RhRh | 09_25 | 0.24 | 29 | 0.085
FDA | RhRh | 09_25 | 0.26 | 20 | 0.067
FDA | RhRh | 09_25 | 0.26 | 21 | 0.071
FDA | RhRh | 09_25 | 0.26 | 22 | 0.073
FDA | RhRh | 09_25 | 0.26 | 23 | 0.076
FDA | RhRh | 09_25 | 0.26 | 24 | 0.081
FDA | RhRh | 09_25 | 0.26 | 25 | 0.084
FDA | RhRh | 09_25 | 0.26 | 26 | 0.087
FDA | RhRh | 09_25 | 0.26 | 27 | 0.089
FDA | RhRh | 09_25 | 0.26 | 28 | 0.091
FDA | RhRh | 09_25 | 0.26 | 29 | 0.093
FDA | RhRh | 09_25 | 0.26 | 30 | 0.095
FDA | RhRh | 09_25 | 0.26 | 31 | 0.097
FDA | RhRh | 09_25 | 0.26 | 32 | 0.098
FDA | RhRh | 09_25 | 0.28 | 20 | 0.074
FDA | RhRh | 09_25 | 0.28 | 21 | 0.078
FDA | RhRh | 09_25 | 0.28 | 22 | 0.081
FDA | RhRh | 09_25 | 0.28 | 23 | 0.084
FDA | RhRh | 09_25 | 0.28 | 24 | 0.088
FDA | RhRh | 09_25 | 0.28 | 25 | 0.092
FDA | RhRh | 09_25 | 0.28 | 26 | 0.095
FDA | RhRh | 09_25 | 0.28 | 27 | 0.097
FDA | RhRh | 09_25 | 0.28 | 28 | 0.099
FDA | RhRh | 09_25 | 0.28 | 29 | 0.101
FDA | RhRh | 09_25 | 0.28 | 30 | 0.103
FDA | RhRh | 09_25 | 0.28 | 31 | 0.104
FDA | RhRh | 09_25 | 0.28 | 32 | 0.106
FDA | RhRh | 09_25 | 0.28 | 33 | 0.108
FDA | RhRh | 09_25 | 0.28 | 34 | 0.109
FDA | RhRh | 09_25 | 0.3 | 20 | 0.081
FDA | RhRh | 09_25 | 0.3 | 21 | 0.085
FDA | RhRh | 09_25 | 0.3 | 22 | 0.088
FDA | RhRh | 09_25 | 0.3 | 23 | 0.092
FDA | RhRh | 09_25 | 0.3 | 24 | 0.096
FDA | RhRh | 09_25 | 0.3 | 25 | 0.1
FDA | RhRh | 09_25 | 0.3 | 26 | 0.102
FDA | RhRh | 09_25 | 0.3 | 27 | 0.105
FDA | RhRh | 09_25 | 0.3 | 28 | 0.107
FDA | RhRh | 09_25 | 0.3 | 29 | 0.109
FDA | RhRh | 09_25 | 0.3 | 30 | 0.11
FDA | RhRh | 09_25 | 0.3 | 31 | 0.112
FDA | RhRh | 09_25 | 0.3 | 32 | 0.113
FDA | RhRh | 09_25 | 0.3 | 33 | 0.115
FDA | RhRh | 09_25 | 0.3 | 34 | 0.116
FDA | RhRh | 09_25 | 0.3 | 35 | 0.118
FDA | RhRh | 09_25 | 0.3 | 36 | 0.119
FDA | RhRh | 09_25 | 0.3 | 37 | 0.12
FDA | RhRh | 09_25 | 0.32 | 20 | 0.088
FDA | RhRh | 09_25 | 0.32 | 21 | 0.093
FDA | RhRh | 09_25 | 0.32 | 22 | 0.096
FDA | RhRh | 09_25 | 0.32 | 23 | 0.099
FDA | RhRh | 09_25 | 0.32 | 24 | 0.104
FDA | RhRh | 09_25 | 0.32 | 25 | 0.107
FDA | RhRh | 09_25 | 0.32 | 26 | 0.11
FDA | RhRh | 09_25 | 0.32 | 27 | 0.112
FDA | RhRh | 09_25 | 0.32 | 28 | 0.114
FDA | RhRh | 09_25 | 0.32 | 29 | 0.116
FDA | RhRh | 09_25 | 0.32 | 30 | 0.118
FDA | RhRh | 09_25 | 0.32 | 31 | 0.119
FDA | RhRh | 09_25 | 0.32 | 32 | 0.121
FDA | RhRh | 09_25 | 0.32 | 33 | 0.122
FDA | RhRh | 09_25 | 0.32 | 34 | 0.123
FDA | RhRh | 09_25 | 0.32 | 35 | 0.125
FDA | RhRh | 09_25 | 0.32 | 36 | 0.126
FDA | RhRh | 09_25 | 0.32 | 37 | 0.127
FDA | RhRh | 09_25 | 0.32 | 38 | 0.128
FDA | RhRh | 09_25 | 0.32 | 39 | 0.129
FDA | RhRh | 09_25 | 0.34 | 20 | 0.096
FDA | RhRh | 09_25 | 0.34 | 21 | 0.1
FDA | RhRh | 09_25 | 0.34 | 22 | 0.103
FDA | RhRh | 09_25 | 0.34 | 23 | 0.107
FDA | RhRh | 09_25 | 0.34 | 24 | 0.111
FDA | RhRh | 09_25 | 0.34 | 25 | 0.115
FDA | RhRh | 09_25 | 0.34 | 26 | 0.117
FDA | RhRh | 09_25 | 0.34 | 27 | 0.12
FDA | RhRh | 09_25 | 0.34 | 28 | 0.121
FDA | RhRh | 09_25 | 0.34 | 29 | 0.123
FDA | RhRh | 09_25 | 0.34 | 30 | 0.125
FDA | RhRh | 09_25 | 0.34 | 31 | 0.126
FDA | RhRh | 09_25 | 0.34 | 32 | 0.128
FDA | RhRh | 09_25 | 0.34 | 33 | 0.129
FDA | RhRh | 09_25 | 0.34 | 34 | 0.13
FDA | RhRh | 09_25 | 0.34 | 35 | 0.131
FDA | RhRh | 09_25 | 0.34 | 36 | 0.133
FDA | RhRh | 09_25 | 0.34 | 37 | 0.134
FDA | RhRh | 09_25 | 0.34 | 38 | 0.135
FDA | RhRh | 09_25 | 0.34 | 39 | 0.136
FDA | RhRh | 09_25 | 0.34 | 40 | 0.137
FDA | RhRh | 09_25 | 0.34 | 41 | 0.138
FDA | RhRh | 09_25 | 0.34 | 42 | 0.139
FDA | RhRh | 09_25 | 0.36 | 20 | 0.103
FDA | RhRh | 09_25 | 0.36 | 21 | 0.107
FDA | RhRh | 09_25 | 0.36 | 22 | 0.111
FDA | RhRh | 09_25 | 0.36 | 23 | 0.114
FDA | RhRh | 09_25 | 0.36 | 24 | 0.119
FDA | RhRh | 09_25 | 0.36 | 25 | 0.122
FDA | RhRh | 09_25 | 0.36 | 26 | 0.125
FDA | RhRh | 09_25 | 0.36 | 27 | 0.127
FDA | RhRh | 09_25 | 0.36 | 28 | 0.129
FDA | RhRh | 09_25 | 0.36 | 29 | 0.13
FDA | RhRh | 09_25 | 0.36 | 30 | 0.132
FDA | RhRh | 09_25 | 0.36 | 31 | 0.133
FDA | RhRh | 09_25 | 0.36 | 32 | 0.134
FDA | RhRh | 09_25 | 0.36 | 33 | 0.136
FDA | RhRh | 09_25 | 0.36 | 34 | 0.137
FDA | RhRh | 09_25 | 0.36 | 35 | 0.138
FDA | RhRh | 09_25 | 0.36 | 36 | 0.139
FDA | RhRh | 09_25 | 0.36 | 37 | 0.14
FDA | RhRh | 09_25 | 0.36 | 38 | 0.141
FDA | RhRh | 09_25 | 0.36 | 39 | 0.142
FDA | RhRh | 09_25 | 0.36 | 40 | 0.143
FDA | RhRh | 09_25 | 0.36 | 41 | 0.144
FDA | RhRh | 09_25 | 0.36 | 42 | 0.145
FDA | RhRh | 09_25 | 0.36 | 43 | 0.146
FDA | RhRh | 09_25 | 0.36 | 44 | 0.147
FDA | RhRh | 09_25 | 0.36 | 45 | 0.147
FDA | RhRh | 09_25 | 0.36 | 46 | 0.148
FDA | RhRh | 09_25 | 0.38 | 20 | 0.11
FDA | RhRh | 09_25 | 0.38 | 21 | 0.115
FDA | RhRh | 09_25 | 0.38 | 22 | 0.118
FDA | RhRh | 09_25 | 0.38 | 23 | 0.121
FDA | RhRh | 09_25 | 0.38 | 24 | 0.126
FDA | RhRh | 09_25 | 0.38 | 25 | 0.129
FDA | RhRh | 09_25 | 0.38 | 26 | 0.132
FDA | RhRh | 09_25 | 0.38 | 27 | 0.134
FDA | RhRh | 09_25 | 0.38 | 28 | 0.136
FDA | RhRh | 09_25 | 0.38 | 29 | 0.137
FDA | RhRh | 09_25 | 0.38 | 30 | 0.138
FDA | RhRh | 09_25 | 0.38 | 31 | 0.14
FDA | RhRh | 09_25 | 0.38 | 32 | 0.141
FDA | RhRh | 09_25 | 0.38 | 33 | 0.142
FDA | RhRh | 09_25 | 0.38 | 34 | 0.143
FDA | RhRh | 09_25 | 0.38 | 35 | 0.144
FDA | RhRh | 09_25 | 0.38 | 36 | 0.145
FDA | RhRh | 09_25 | 0.38 | 37 | 0.147
FDA | RhRh | 09_25 | 0.38 | 38 | 0.148
FDA | RhRh | 09_25 | 0.38 | 39 | 0.149
FDA | RhRh | 09_25 | 0.38 | 40 | 0.149
FDA | RhRh | 09_25 | 0.38 | 41 | 0.15
FDA | RhRh | 09_25 | 0.38 | 42 | 0.151
FDA | RhRh | 09_25 | 0.38 | 43 | 0.152
FDA | RhRh | 09_25 | 0.38 | 44 | 0.153
FDA | RhRh | 09_25 | 0.38 | 45 | 0.154
FDA | RhRh | 09_25 | 0.38 | 46 | 0.154
FDA | RhRh | 09_25 | 0.38 | 47 | 0.155
FDA | RhRh | 09_25 | 0.38 | 48 | 0.156
FDA | RhRh | 09_25 | 0.38 | 49 | 0.156
FDA | RhRh | 09_25 | 0.4 | 20 | 0.117
FDA | RhRh | 09_25 | 0.4 | 21 | 0.122
FDA | RhRh | 09_25 | 0.4 | 22 | 0.125
FDA | RhRh | 09_25 | 0.4 | 23 | 0.129
FDA | RhRh | 09_25 | 0.4 | 24 | 0.133
FDA | RhRh | 09_25 | 0.4 | 25 | 0.137
FDA | RhRh | 09_25 | 0.4 | 26 | 0.139
FDA | RhRh | 09_25 | 0.4 | 27 | 0.141
FDA | RhRh | 09_25 | 0.4 | 28 | 0.142
FDA | RhRh | 09_25 | 0.4 | 29 | 0.144
FDA | RhRh | 09_25 | 0.4 | 30 | 0.145
FDA | RhRh | 09_25 | 0.4 | 31 | 0.146
FDA | RhRh | 09_25 | 0.4 | 32 | 0.147
FDA | RhRh | 09_25 | 0.4 | 33 | 0.148
FDA | RhRh | 09_25 | 0.4 | 34 | 0.15
FDA | RhRh | 09_25 | 0.4 | 35 | 0.151
FDA | RhRh | 09_25 | 0.4 | 36 | 0.152
FDA | RhRh | 09_25 | 0.4 | 37 | 0.153
FDA | RhRh | 09_25 | 0.4 | 38 | 0.154
FDA | RhRh | 09_25 | 0.4 | 39 | 0.155
FDA | RhRh | 09_25 | 0.4 | 40 | 0.156
FDA | RhRh | 09_25 | 0.4 | 41 | 0.156
FDA | RhRh | 09_25 | 0.4 | 42 | 0.157
FDA | RhRh | 09_25 | 0.4 | 43 | 0.158
FDA | RhRh | 09_25 | 0.4 | 44 | 0.159
FDA | RhRh | 09_25 | 0.4 | 45 | 0.16
FDA | RhRh | 09_25 | 0.4 | 46 | 0.16
FDA | RhRh | 09_25 | 0.4 | 47 | 0.161
FDA | RhRh | 09_25 | 0.4 | 48 | 0.162
FDA | RhRh | 09_25 | 0.4 | 49 | 0.162
FDA | RhRh | 09_25 | 0.42 | 21 | 0.129
FDA | RhRh | 09_25 | 0.42 | 22 | 0.133
FDA | RhRh | 09_25 | 0.42 | 23 | 0.136
FDA | RhRh | 09_25 | 0.42 | 24 | 0.141
FDA | RhRh | 09_25 | 0.42 | 25 | 0.144
FDA | RhRh | 09_25 | 0.42 | 26 | 0.146
FDA | RhRh | 09_25 | 0.42 | 27 | 0.148
FDA | RhRh | 09_25 | 0.42 | 28 | 0.149
FDA | RhRh | 09_25 | 0.42 | 29 | 0.15
FDA | RhRh | 09_25 | 0.42 | 30 | 0.152
FDA | RhRh | 09_25 | 0.42 | 31 | 0.153
FDA | RhRh | 09_25 | 0.42 | 32 | 0.154
FDA | RhRh | 09_25 | 0.42 | 33 | 0.155
FDA | RhRh | 09_25 | 0.42 | 34 | 0.156
FDA | RhRh | 09_25 | 0.42 | 35 | 0.157
FDA | RhRh | 09_25 | 0.42 | 36 | 0.158
FDA | RhRh | 09_25 | 0.42 | 37 | 0.159
FDA | RhRh | 09_25 | 0.42 | 38 | 0.16
FDA | RhRh | 09_25 | 0.42 | 39 | 0.161
FDA | RhRh | 09_25 | 0.42 | 40 | 0.161
FDA | RhRh | 09_25 | 0.42 | 41 | 0.162
FDA | RhRh | 09_25 | 0.42 | 42 | 0.163
FDA | RhRh | 09_25 | 0.42 | 43 | 0.164
FDA | RhRh | 09_25 | 0.42 | 44 | 0.165
FDA | RhRh | 09_25 | 0.42 | 45 | 0.165
FDA | RhRh | 09_25 | 0.42 | 46 | 0.166
FDA | RhRh | 09_25 | 0.42 | 47 | 0.167
FDA | RhRh | 09_25 | 0.42 | 48 | 0.167
FDA | RhRh | 09_25 | 0.42 | 49 | 0.168
FDA | RhRh | 09_25 | 0.44 | 21 | 0.137
FDA | RhRh | 09_25 | 0.44 | 22 | 0.14
FDA | RhRh | 09_25 | 0.44 | 23 | 0.144
FDA | RhRh | 09_25 | 0.44 | 24 | 0.148
FDA | RhRh | 09_25 | 0.44 | 25 | 0.151
FDA | RhRh | 09_25 | 0.44 | 26 | 0.153
FDA | RhRh | 09_25 | 0.44 | 27 | 0.155
FDA | RhRh | 09_25 | 0.44 | 28 | 0.156
FDA | RhRh | 09_25 | 0.44 | 29 | 0.157
FDA | RhRh | 09_25 | 0.44 | 30 | 0.158
FDA | RhRh | 09_25 | 0.44 | 31 | 0.159
FDA | RhRh | 09_25 | 0.44 | 32 | 0.16
FDA | RhRh | 09_25 | 0.44 | 33 | 0.161
FDA | RhRh | 09_25 | 0.44 | 34 | 0.162
FDA | RhRh | 09_25 | 0.44 | 35 | 0.163
FDA | RhRh | 09_25 | 0.44 | 36 | 0.164
FDA | RhRh | 09_25 | 0.44 | 37 | 0.165
FDA | RhRh | 09_25 | 0.44 | 38 | 0.166
FDA | RhRh | 09_25 | 0.44 | 39 | 0.166
FDA | RhRh | 09_25 | 0.44 | 40 | 0.167
FDA | RhRh | 09_25 | 0.44 | 41 | 0.168
FDA | RhRh | 09_25 | 0.44 | 42 | 0.169
FDA | RhRh | 09_25 | 0.44 | 43 | 0.17
FDA | RhRh | 09_25 | 0.44 | 44 | 0.17
FDA | RhRh | 09_25 | 0.44 | 45 | 0.171
FDA | RhRh | 09_25 | 0.44 | 46 | 0.172
FDA | RhRh | 09_25 | 0.44 | 47 | 0.172
FDA | RhRh | 09_25 | 0.44 | 48 | 0.173
FDA | RhRh | 09_25 | 0.44 | 49 | 0.174
FDA | RhRh | 09_25 | 0.46 | 22 | 0.147
FDA | RhRh | 09_25 | 0.46 | 23 | 0.151
FDA | RhRh | 09_25 | 0.46 | 24 | 0.155
FDA | RhRh | 09_25 | 0.46 | 25 | 0.158
FDA | RhRh | 09_25 | 0.46 | 26 | 0.16
FDA | RhRh | 09_25 | 0.46 | 27 | 0.161
FDA | RhRh | 09_25 | 0.46 | 28 | 0.163
FDA | RhRh | 09_25 | 0.46 | 29 | 0.164
FDA | RhRh | 09_25 | 0.46 | 30 | 0.164
FDA | RhRh | 09_25 | 0.46 | 31 | 0.165
FDA | RhRh | 09_25 | 0.46 | 32 | 0.166
FDA | RhRh | 09_25 | 0.46 | 33 | 0.167
FDA | RhRh | 09_25 | 0.46 | 34 | 0.168
FDA | RhRh | 09_25 | 0.46 | 35 | 0.169
FDA | RhRh | 09_25 | 0.46 | 36 | 0.17
FDA | RhRh | 09_25 | 0.46 | 37 | 0.17
FDA | RhRh | 09_25 | 0.46 | 38 | 0.171
FDA | RhRh | 09_25 | 0.46 | 39 | 0.172
FDA | RhRh | 09_25 | 0.46 | 40 | 0.173
FDA | RhRh | 09_25 | 0.46 | 41 | 0.174
FDA | RhRh | 09_25 | 0.46 | 42 | 0.175
FDA | RhRh | 09_25 | 0.46 | 43 | 0.175
FDA | RhRh | 09_25 | 0.46 | 44 | 0.176
FDA | RhRh | 09_25 | 0.46 | 45 | 0.177
FDA | RhRh | 09_25 | 0.46 | 46 | 0.177
FDA | RhRh | 09_25 | 0.46 | 47 | 0.178
FDA | RhRh | 09_25 | 0.46 | 48 | 0.179
FDA | RhRh | 09_25 | 0.46 | 49 | 0.179
FDA | RhRh | 09_25 | 0.48 | 22 | 0.155
FDA | RhRh | 09_25 | 0.48 | 23 | 0.158
FDA | RhRh | 09_25 | 0.48 | 24 | 0.162
FDA | RhRh | 09_25 | 0.48 | 25 | 0.165
FDA | RhRh | 09_25 | 0.48 | 26 | 0.167
FDA | RhRh | 09_25 | 0.48 | 27 | 0.168
FDA | RhRh | 09_25 | 0.48 | 28 | 0.169
FDA | RhRh | 09_25 | 0.48 | 29 | 0.17
FDA | RhRh | 09_25 | 0.48 | 30 | 0.171
FDA | RhRh | 09_25 | 0.48 | 31 | 0.172
FDA | RhRh | 09_25 | 0.48 | 32 | 0.172
FDA | RhRh | 09_25 | 0.48 | 33 | 0.173
FDA | RhRh | 09_25 | 0.48 | 34 | 0.174
FDA | RhRh | 09_25 | 0.48 | 35 | 0.175
FDA | RhRh | 09_25 | 0.48 | 36 | 0.175
FDA | RhRh | 09_25 | 0.48 | 37 | 0.176
FDA | RhRh | 09_25 | 0.48 | 38 | 0.177
FDA | RhRh | 09_25 | 0.48 | 39 | 0.178
FDA | RhRh | 09_25 | 0.48 | 40 | 0.179
FDA | RhRh | 09_25 | 0.48 | 41 | 0.179
FDA | RhRh | 09_25 | 0.48 | 42 | 0.18
FDA | RhRh | 09_25 | 0.48 | 43 | 0.181
FDA | RhRh | 09_25 | 0.48 | 44 | 0.182
FDA | RhRh | 09_25 | 0.48 | 45 | 0.182
FDA | RhRh | 09_25 | 0.48 | 46 | 0.183
FDA | RhRh | 09_25 | 0.48 | 47 | 0.183
FDA | RhRh | 09_25 | 0.48 | 48 | 0.184
FDA | RhRh | 09_25 | 0.48 | 49 | 0.185
FDA | RhRh | 09_25 | 0.5 | 23 | 0.165
FDA | RhRh | 09_25 | 0.5 | 24 | 0.169
FDA | RhRh | 09_25 | 0.5 | 25 | 0.172
FDA | RhRh | 09_25 | 0.5 | 26 | 0.174
FDA | RhRh | 09_25 | 0.5 | 27 | 0.175
FDA | RhRh | 09_25 | 0.5 | 28 | 0.176
FDA | RhRh | 09_25 | 0.5 | 29 | 0.176
FDA | RhRh | 09_25 | 0.5 | 30 | 0.177
FDA | RhRh | 09_25 | 0.5 | 31 | 0.178
FDA | RhRh | 09_25 | 0.5 | 32 | 0.178
FDA | RhRh | 09_25 | 0.5 | 33 | 0.179
FDA | RhRh | 09_25 | 0.5 | 34 | 0.18
FDA | RhRh | 09_25 | 0.5 | 35 | 0.18
FDA | RhRh | 09_25 | 0.5 | 36 | 0.181
FDA | RhRh | 09_25 | 0.5 | 37 | 0.182
FDA | RhRh | 09_25 | 0.5 | 38 | 0.183
FDA | RhRh | 09_25 | 0.5 | 39 | 0.183
FDA | RhRh | 09_25 | 0.5 | 40 | 0.184
FDA | RhRh | 09_25 | 0.5 | 41 | 0.185
FDA | RhRh | 09_25 | 0.5 | 42 | 0.186
FDA | RhRh | 09_25 | 0.5 | 43 | 0.186
FDA | RhRh | 09_25 | 0.5 | 44 | 0.187
FDA | RhRh | 09_25 | 0.5 | 45 | 0.188
FDA | RhRh | 09_25 | 0.5 | 46 | 0.188
FDA | RhRh | 09_25 | 0.5 | 47 | 0.189
FDA | RhRh | 09_25 | 0.5 | 48 | 0.19
FDA | RhRh | 09_25 | 0.5 | 49 | 0.19
FDA | RhRh | 09_25 | 0.52 | 23 | 0.173
FDA | RhRh | 09_25 | 0.52 | 24 | 0.176
FDA | RhRh | 09_25 | 0.52 | 25 | 0.179
FDA | RhRh | 09_25 | 0.52 | 26 | 0.18
FDA | RhRh | 09_25 | 0.52 | 27 | 0.181
FDA | RhRh | 09_25 | 0.52 | 28 | 0.182
FDA | RhRh | 09_25 | 0.52 | 29 | 0.183
FDA | RhRh | 09_25 | 0.52 | 30 | 0.183
FDA | RhRh | 09_25 | 0.52 | 31 | 0.184
FDA | RhRh | 09_25 | 0.52 | 32 | 0.184
FDA | RhRh | 09_25 | 0.52 | 33 | 0.185
FDA | RhRh | 09_25 | 0.52 | 34 | 0.186
FDA | RhRh | 09_25 | 0.52 | 35 | 0.186
FDA | RhRh | 09_25 | 0.52 | 36 | 0.187
FDA | RhRh | 09_25 | 0.52 | 37 | 0.188
FDA | RhRh | 09_25 | 0.52 | 38 | 0.188
FDA | RhRh | 09_25 | 0.52 | 39 | 0.189
FDA | RhRh | 09_25 | 0.52 | 40 | 0.19
FDA | RhRh | 09_25 | 0.52 | 41 | 0.19
FDA | RhRh | 09_25 | 0.52 | 42 | 0.191
FDA | RhRh | 09_25 | 0.52 | 43 | 0.192
FDA | RhRh | 09_25 | 0.52 | 44 | 0.192
FDA | RhRh | 09_25 | 0.52 | 45 | 0.193
FDA | RhRh | 09_25 | 0.52 | 46 | 0.194
FDA | RhRh | 09_25 | 0.52 | 47 | 0.194
FDA | RhRh | 09_25 | 0.52 | 48 | 0.195
FDA | RhRh | 09_25 | 0.52 | 49 | 0.195
FDA | RhRh | 09_25 | 0.54 | 24 | 0.183
FDA | RhRh | 09_25 | 0.54 | 25 | 0.186
FDA | RhRh | 09_25 | 0.54 | 26 | 0.187
FDA | RhRh | 09_25 | 0.54 | 27 | 0.188
FDA | RhRh | 09_25 | 0.54 | 28 | 0.189
FDA | RhRh | 09_25 | 0.54 | 29 | 0.189
FDA | RhRh | 09_25 | 0.54 | 30 | 0.19
FDA | RhRh | 09_25 | 0.54 | 31 | 0.19
FDA | RhRh | 09_25 | 0.54 | 32 | 0.191
FDA | RhRh | 09_25 | 0.54 | 33 | 0.191
FDA | RhRh | 09_25 | 0.54 | 34 | 0.192
FDA | RhRh | 09_25 | 0.54 | 35 | 0.192
FDA | RhRh | 09_25 | 0.54 | 36 | 0.193
FDA | RhRh | 09_25 | 0.54 | 37 | 0.193
FDA | RhRh | 09_25 | 0.54 | 38 | 0.194
FDA | RhRh | 09_25 | 0.54 | 39 | 0.195
FDA | RhRh | 09_25 | 0.54 | 40 | 0.195
FDA | RhRh | 09_25 | 0.54 | 41 | 0.196
FDA | RhRh | 09_25 | 0.54 | 42 | 0.197
FDA | RhRh | 09_25 | 0.54 | 43 | 0.197
FDA | RhRh | 09_25 | 0.54 | 44 | 0.198
FDA | RhRh | 09_25 | 0.54 | 45 | 0.198
FDA | RhRh | 09_25 | 0.54 | 46 | 0.199
FDA | RhRh | 09_25 | 0.54 | 47 | 0.2
FDA | RhRh | 09_25 | 0.54 | 48 | 0.2
FDA | RhRh | 09_25 | 0.54 | 49 | 0.201
FDA | RhRh | 09_25 | 0.56 | 26 | 0.194
FDA | RhRh | 09_25 | 0.56 | 27 | 0.194
FDA | RhRh | 09_25 | 0.56 | 28 | 0.195
FDA | RhRh | 09_25 | 0.56 | 29 | 0.195
FDA | RhRh | 09_25 | 0.56 | 30 | 0.196
FDA | RhRh | 09_25 | 0.56 | 31 | 0.196
FDA | RhRh | 09_25 | 0.56 | 32 | 0.197
FDA | RhRh | 09_25 | 0.56 | 33 | 0.197
FDA | RhRh | 09_25 | 0.56 | 34 | 0.197
FDA | RhRh | 09_25 | 0.56 | 35 | 0.198
FDA | RhRh | 09_25 | 0.56 | 36 | 0.198
FDA | RhRh | 09_25 | 0.56 | 37 | 0.199
FDA | RhRh | 09_25 | 0.56 | 38 | 0.2
FDA | RhRh | 09_25 | 0.56 | 39 | 0.2
FDA | RhRh | 09_25 | 0.56 | 40 | 0.201
FDA | RhRh | 09_25 | 0.56 | 41 | 0.201
FDA | RhRh | 09_25 | 0.56 | 42 | 0.202
FDA | RhRh | 09_25 | 0.56 | 43 | 0.203
FDA | RhRh | 09_25 | 0.56 | 44 | 0.203
FDA | RhRh | 09_25 | 0.56 | 45 | 0.204
FDA | RhRh | 09_25 | 0.56 | 46 | 0.204
FDA | RhRh | 09_25 | 0.56 | 47 | 0.205
FDA | RhRh | 09_25 | 0.56 | 48 | 0.205
FDA | RhRh | 09_25 | 0.56 | 49 | 0.206
FDA | RhRh | 09_25 | 0.58 | 28 | 0.201
FDA | RhRh | 09_25 | 0.58 | 29 | 0.202
FDA | RhRh | 09_25 | 0.58 | 30 | 0.202
FDA | RhRh | 09_25 | 0.58 | 31 | 0.202
FDA | RhRh | 09_25 | 0.58 | 32 | 0.203
FDA | RhRh | 09_25 | 0.58 | 33 | 0.203
FDA | RhRh | 09_25 | 0.58 | 34 | 0.203
FDA | RhRh | 09_25 | 0.58 | 35 | 0.204
FDA | RhRh | 09_25 | 0.58 | 36 | 0.204
FDA | RhRh | 09_25 | 0.58 | 37 | 0.205
FDA | RhRh | 09_25 | 0.58 | 38 | 0.205
FDA | RhRh | 09_25 | 0.58 | 39 | 0.206
FDA | RhRh | 09_25 | 0.58 | 40 | 0.206
FDA | RhRh | 09_25 | 0.58 | 41 | 0.207
FDA | RhRh | 09_25 | 0.58 | 42 | 0.207
FDA | RhRh | 09_25 | 0.58 | 43 | 0.208
FDA | RhRh | 09_25 | 0.58 | 44 | 0.209
FDA | RhRh | 09_25 | 0.58 | 45 | 0.209
FDA | RhRh | 09_25 | 0.58 | 46 | 0.21
FDA | RhRh | 09_25 | 0.58 | 47 | 0.21
FDA | RhRh | 09_25 | 0.58 | 48 | 0.211
FDA | RhRh | 09_25 | 0.58 | 49 | 0.211
FDA | RhRh | 09_25 | 0.6 | 30 | 0.208
FDA | RhRh | 09_25 | 0.6 | 31 | 0.208
FDA | RhRh | 09_25 | 0.6 | 32 | 0.209
FDA | RhRh | 09_25 | 0.6 | 33 | 0.209
FDA | RhRh | 09_25 | 0.6 | 34 | 0.209
FDA | RhRh | 09_25 | 0.6 | 35 | 0.209
FDA | RhRh | 09_25 | 0.6 | 36 | 0.21
FDA | RhRh | 09_25 | 0.6 | 37 | 0.21
FDA | RhRh | 09_25 | 0.6 | 38 | 0.211
FDA | RhRh | 09_25 | 0.6 | 39 | 0.211
FDA | RhRh | 09_25 | 0.6 | 40 | 0.212
FDA | RhRh | 09_25 | 0.6 | 41 | 0.212
FDA | RhRh | 09_25 | 0.6 | 42 | 0.213
FDA | RhRh | 09_25 | 0.6 | 43 | 0.213
FDA | RhRh | 09_25 | 0.6 | 44 | 0.214
FDA | RhRh | 09_25 | 0.6 | 45 | 0.214
FDA | RhRh | 09_25 | 0.6 | 46 | 0.215
FDA | RhRh | 09_25 | 0.6 | 47 | 0.215
FDA | RhRh | 09_25 | 0.6 | 48 | 0.216
FDA | RhRh | 09_25 | 0.6 | 49 | 0.217
FDA | RhRh | 09_25 | 0.62 | 33 | 0.215
FDA | RhRh | 09_25 | 0.62 | 34 | 0.215
FDA | RhRh | 09_25 | 0.62 | 35 | 0.215
FDA | RhRh | 09_25 | 0.62 | 36 | 0.216
FDA | RhRh | 09_25 | 0.62 | 37 | 0.216
FDA | RhRh | 09_25 | 0.62 | 38 | 0.216
FDA | RhRh | 09_25 | 0.62 | 39 | 0.217
FDA | RhRh | 09_25 | 0.62 | 40 | 0.217
FDA | RhRh | 09_25 | 0.62 | 41 | 0.218
FDA | RhRh | 09_25 | 0.62 | 42 | 0.218
FDA | RhRh | 09_25 | 0.62 | 43 | 0.219
FDA | RhRh | 09_25 | 0.62 | 44 | 0.219
FDA | RhRh | 09_25 | 0.62 | 45 | 0.22
FDA | RhRh | 09_25 | 0.62 | 46 | 0.22
FDA | RhRh | 09_25 | 0.62 | 47 | 0.221
FDA | RhRh | 09_25 | 0.62 | 48 | 0.221
FDA | RhRh | 09_25 | 0.62 | 49 | 0.222
FDA | RhRh | 09_25 | 0.64 | 35 | 0.221
FDA | RhRh | 09_25 | 0.64 | 36 | 0.221
FDA | RhRh | 09_25 | 0.64 | 37 | 0.222
FDA | RhRh | 09_25 | 0.64 | 38 | 0.222
FDA | RhRh | 09_25 | 0.64 | 39 | 0.222
FDA | RhRh | 09_25 | 0.64 | 40 | 0.223
FDA | RhRh | 09_25 | 0.64 | 41 | 0.223
FDA | RhRh | 09_25 | 0.64 | 42 | 0.224
FDA | RhRh | 09_25 | 0.64 | 43 | 0.224
FDA | RhRh | 09_25 | 0.64 | 44 | 0.225
FDA | RhRh | 09_25 | 0.64 | 45 | 0.225
FDA | RhRh | 09_25 | 0.64 | 46 | 0.226
FDA | RhRh | 09_25 | 0.64 | 47 | 0.226
FDA | RhRh | 09_25 | 0.64 | 48 | 0.227
FDA | RhRh | 09_25 | 0.64 | 49 | 0.227
FDA | RhRh | 09_25 | 0.66 | 38 | 0.228
FDA | RhRh | 09_25 | 0.66 | 39 | 0.228
FDA | RhRh | 09_25 | 0.66 | 40 | 0.228
FDA | RhRh | 09_25 | 0.66 | 41 | 0.229
FDA | RhRh | 09_25 | 0.66 | 42 | 0.229
FDA | RhRh | 09_25 | 0.66 | 43 | 0.23
FDA | RhRh | 09_25 | 0.66 | 44 | 0.23
FDA | RhRh | 09_25 | 0.66 | 45 | 0.23
FDA | RhRh | 09_25 | 0.66 | 46 | 0.231
FDA | RhRh | 09_25 | 0.66 | 47 | 0.231
FDA | RhRh | 09_25 | 0.66 | 48 | 0.232
FDA | RhRh | 09_25 | 0.66 | 49 | 0.232
FDA | RhRh | 09_25 | 0.68 | 40 | 0.234
FDA | RhRh | 09_25 | 0.68 | 41 | 0.234
FDA | RhRh | 09_25 | 0.68 | 42 | 0.235
FDA | RhRh | 09_25 | 0.68 | 43 | 0.235
FDA | RhRh | 09_25 | 0.68 | 44 | 0.235
FDA | RhRh | 09_25 | 0.68 | 45 | 0.236
FDA | RhRh | 09_25 | 0.68 | 46 | 0.236
FDA | RhRh | 09_25 | 0.68 | 47 | 0.237
FDA | RhRh | 09_25 | 0.68 | 48 | 0.237
FDA | RhRh | 09_25 | 0.68 | 49 | 0.238
FDA | RhRh | 09_25 | 0.7 | 43 | 0.241
FDA | RhRh | 09_25 | 0.7 | 44 | 0.241
FDA | RhRh | 09_25 | 0.7 | 45 | 0.241
FDA | RhRh | 09_25 | 0.7 | 46 | 0.242
FDA | RhRh | 09_25 | 0.7 | 47 | 0.242
FDA | RhRh | 09_25 | 0.7 | 48 | 0.243
FDA | RhRh | 09_25 | 0.7 | 49 | 0.243
FDA | RhRh | 09_25 | 0.72 | 47 | 0.248
FDA | RhRh | 09_25 | 0.72 | 48 | 0.248
FDA | RhRh | 09_25 | 0.72 | 49 | 0.249
FDA | WAl | 15_15 | 0.2 | 20 | 0.048
FDA | WAl | 15_15 | 0.2 | 21 | 0.051
FDA | WAl | 15_15 | 0.25 | 20 | 0.065
FDA | WAl | 15_15 | 0.25 | 21 | 0.069
FDA | WAl | 15_15 | 0.25 | 22 | 0.072
FDA | WAl | 15_15 | 0.25 | 23 | 0.076
FDA | WAl | 15_15 | 0.25 | 24 | 0.079
FDA | WAl | 15_15 | 0.25 | 25 | 0.083
FDA | WAl | 15_15 | 0.3 | 20 | 0.082
FDA | WAl | 15_15 | 0.3 | 21 | 0.086
FDA | WAl | 15_15 | 0.3 | 22 | 0.09
FDA | WAl | 15_15 | 0.3 | 23 | 0.094
FDA | WAl | 15_15 | 0.3 | 24 | 0.098
FDA | WAl | 15_15 | 0.3 | 25 | 0.101
FDA | WAl | 15_15 | 0.3 | 26 | 0.105
FDA | WAl | 15_15 | 0.3 | 27 | 0.108
FDA | WAl | 15_15 | 0.3 | 28 | 0.111
FDA | WAl | 15_15 | 0.3 | 29 | 0.114
FDA | WAl | 15_15 | 0.3 | 30 | 0.116
FDA | WAl | 15_15 | 0.35 | 20 | 0.1
FDA | WAl | 15_15 | 0.35 | 21 | 0.104
FDA | WAl | 15_15 | 0.35 | 22 | 0.108
FDA | WAl | 15_15 | 0.35 | 23 | 0.112
FDA | WAl | 15_15 | 0.35 | 24 | 0.116
FDA | WAl | 15_15 | 0.35 | 25 | 0.119
FDA | WAl | 15_15 | 0.35 | 26 | 0.123
FDA | WAl | 15_15 | 0.35 | 27 | 0.126
FDA | WAl | 15_15 | 0.35 | 28 | 0.129
FDA | WAl | 15_15 | 0.35 | 29 | 0.132
FDA | WAl | 15_15 | 0.35 | 30 | 0.135
FDA | WAl | 15_15 | 0.35 | 31 | 0.138
FDA | WAl | 15_15 | 0.35 | 32 | 0.14
FDA | WAl | 15_15 | 0.35 | 33 | 0.143
FDA | WAl | 15_15 | 0.35 | 34 | 0.145
FDA | WAl | 15_15 | 0.35 | 35 | 0.147
FDA | WAl | 15_15 | 0.4 | 20 | 0.118
FDA | WAl | 15_15 | 0.4 | 21 | 0.122
FDA | WAl | 15_15 | 0.4 | 22 | 0.126
FDA | WAl | 15_15 | 0.4 | 23 | 0.13
FDA | WAl | 15_15 | 0.4 | 24 | 0.134
FDA | WAl | 15_15 | 0.4 | 25 | 0.137
FDA | WAl | 15_15 | 0.4 | 26 | 0.14
FDA | WAl | 15_15 | 0.4 | 27 | 0.144
FDA | WAl | 15_15 | 0.4 | 28 | 0.147
FDA | WAl | 15_15 | 0.4 | 29 | 0.15
FDA | WAl | 15_15 | 0.4 | 30 | 0.153
FDA | WAl | 15_15 | 0.4 | 31 | 0.155
FDA | WAl | 15_15 | 0.4 | 32 | 0.158
FDA | WAl | 15_15 | 0.4 | 33 | 0.16
FDA | WAl | 15_15 | 0.4 | 34 | 0.163
FDA | WAl | 15_15 | 0.4 | 35 | 0.165
FDA | WAl | 15_15 | 0.4 | 36 | 0.167
FDA | WAl | 15_15 | 0.4 | 37 | 0.169
FDA | WAl | 15_15 | 0.4 | 38 | 0.171
FDA | WAl | 15_15 | 0.4 | 39 | 0.173
FDA | WAl | 15_15 | 0.4 | 40 | 0.175
FDA | WAl | 15_15 | 0.45 | 20 | 0.137
FDA | WAl | 15_15 | 0.45 | 21 | 0.141
FDA | WAl | 15_15 | 0.45 | 22 | 0.145
FDA | WAl | 15_15 | 0.45 | 23 | 0.148
FDA | WAl | 15_15 | 0.45 | 24 | 0.152
FDA | WAl | 15_15 | 0.45 | 25 | 0.155
FDA | WAl | 15_15 | 0.45 | 26 | 0.158
FDA | WAl | 15_15 | 0.45 | 27 | 0.161
FDA | WAl | 15_15 | 0.45 | 28 | 0.164
FDA | WAl | 15_15 | 0.45 | 29 | 0.167
FDA | WAl | 15_15 | 0.45 | 30 | 0.17
FDA | WAl | 15_15 | 0.45 | 31 | 0.173
FDA | WAl | 15_15 | 0.45 | 32 | 0.175
FDA | WAl | 15_15 | 0.45 | 33 | 0.177
FDA | WAl | 15_15 | 0.45 | 34 | 0.18
FDA | WAl | 15_15 | 0.45 | 35 | 0.182
FDA | WAl | 15_15 | 0.45 | 36 | 0.184
FDA | WAl | 15_15 | 0.45 | 37 | 0.186
FDA | WAl | 15_15 | 0.45 | 38 | 0.188
FDA | WAl | 15_15 | 0.45 | 39 | 0.19
FDA | WAl | 15_15 | 0.45 | 40 | 0.192
FDA | WAl | 15_15 | 0.45 | 41 | 0.193
FDA | WAl | 15_15 | 0.45 | 42 | 0.195
FDA | WAl | 15_15 | 0.45 | 43 | 0.197
FDA | WAl | 15_15 | 0.45 | 44 | 0.198
FDA | WAl | 15_15 | 0.5 | 21 | 0.159
FDA | WAl | 15_15 | 0.5 | 22 | 0.163
FDA | WAl | 15_15 | 0.5 | 23 | 0.166
FDA | WAl | 15_15 | 0.5 | 24 | 0.17
FDA | WAl | 15_15 | 0.5 | 25 | 0.173
FDA | WAl | 15_15 | 0.5 | 26 | 0.176
FDA | WAl | 15_15 | 0.5 | 27 | 0.179
FDA | WAl | 15_15 | 0.5 | 28 | 0.182
FDA | WAl | 15_15 | 0.5 | 29 | 0.184
FDA | WAl | 15_15 | 0.5 | 30 | 0.187
FDA | WAl | 15_15 | 0.5 | 31 | 0.189
FDA | WAl | 15_15 | 0.5 | 32 | 0.192
FDA | WAl | 15_15 | 0.5 | 33 | 0.194
FDA | WAl | 15_15 | 0.5 | 34 | 0.196
FDA | WAl | 15_15 | 0.5 | 35 | 0.198
FDA | WAl | 15_15 | 0.5 | 36 | 0.2
FDA | WAl | 15_15 | 0.5 | 37 | 0.202
FDA | WAl | 15_15 | 0.5 | 38 | 0.204
FDA | WAl | 15_15 | 0.5 | 39 | 0.206
FDA | WAl | 15_15 | 0.5 | 40 | 0.207
FDA | WAl | 15_15 | 0.5 | 41 | 0.209
FDA | WAl | 15_15 | 0.5 | 42 | 0.211
FDA | WAl | 15_15 | 0.5 | 43 | 0.212
FDA | WAl | 15_15 | 0.5 | 44 | 0.214
FDA | WAl | 15_15 | 0.5 | 45 | 0.215
FDA | WAl | 15_15 | 0.5 | 46 | 0.217
FDA | WAl | 15_15 | 0.5 | 47 | 0.218
FDA | WAl | 15_15 | 0.5 | 48 | 0.219
FDA | WAl | 15_15 | 0.5 | 49 | 0.221
FDA | WAl | 15_15 | 0.55 | 22 | 0.181
FDA | WAl | 15_15 | 0.55 | 23 | 0.184
FDA | WAl | 15_15 | 0.55 | 24 | 0.187
FDA | WAl | 15_15 | 0.55 | 25 | 0.19
FDA | WAl | 15_15 | 0.55 | 26 | 0.193
FDA | WAl | 15_15 | 0.55 | 27 | 0.196
FDA | WAl | 15_15 | 0.55 | 28 | 0.199
FDA | WAl | 15_15 | 0.55 | 29 | 0.201
FDA | WAl | 15_15 | 0.55 | 30 | 0.204
FDA | WAl | 15_15 | 0.55 | 31 | 0.206
FDA | WAl | 15_15 | 0.55 | 32 | 0.208
FDA | WAl | 15_15 | 0.55 | 33 | 0.21
FDA | WAl | 15_15 | 0.55 | 34 | 0.212
FDA | WAl | 15_15 | 0.55 | 35 | 0.214
FDA | WAl | 15_15 | 0.55 | 36 | 0.216
FDA | WAl | 15_15 | 0.55 | 37 | 0.218
FDA | WAl | 15_15 | 0.55 | 38 | 0.22
FDA | WAl | 15_15 | 0.55 | 39 | 0.221
FDA | WAl | 15_15 | 0.55 | 40 | 0.223
FDA | WAl | 15_15 | 0.55 | 41 | 0.225
FDA | WAl | 15_15 | 0.55 | 42 | 0.226
FDA | WAl | 15_15 | 0.55 | 43 | 0.228
FDA | WAl | 15_15 | 0.55 | 44 | 0.229
FDA | WAl | 15_15 | 0.55 | 45 | 0.23
FDA | WAl | 15_15 | 0.55 | 46 | 0.232
FDA | WAl | 15_15 | 0.55 | 47 | 0.233
FDA | WAl | 15_15 | 0.55 | 48 | 0.234
FDA | WAl | 15_15 | 0.55 | 49 | 0.236
FDA | WAl | 15_15 | 0.6 | 23 | 0.202
FDA | WAl | 15_15 | 0.6 | 24 | 0.205
FDA | WAl | 15_15 | 0.6 | 25 | 0.207
FDA | WAl | 15_15 | 0.6 | 26 | 0.21
FDA | WAl | 15_15 | 0.6 | 27 | 0.212
FDA | WAl | 15_15 | 0.6 | 28 | 0.215
FDA | WAl | 15_15 | 0.6 | 29 | 0.218
FDA | WAl | 15_15 | 0.6 | 30 | 0.22
FDA | WAl | 15_15 | 0.6 | 31 | 0.222
FDA | WAl | 15_15 | 0.6 | 32 | 0.224
FDA | WAl | 15_15 | 0.6 | 33 | 0.226
FDA | WAl | 15_15 | 0.6 | 34 | 0.228
FDA | WAl | 15_15 | 0.6 | 35 | 0.23
FDA | WAl | 15_15 | 0.6 | 36 | 0.232
FDA | WAl | 15_15 | 0.6 | 37 | 0.233
FDA | WAl | 15_15 | 0.6 | 38 | 0.235
FDA | WAl | 15_15 | 0.6 | 39 | 0.237
FDA | WAl | 15_15 | 0.6 | 40 | 0.238
FDA | WAl | 15_15 | 0.6 | 41 | 0.24
FDA | WAl | 15_15 | 0.6 | 42 | 0.241
FDA | WAl | 15_15 | 0.6 | 43 | 0.243
FDA | WAl | 15_15 | 0.6 | 44 | 0.244
FDA | WAl | 15_15 | 0.6 | 45 | 0.245
FDA | WAl | 15_15 | 0.6 | 46 | 0.246
FDA | WAl | 15_15 | 0.6 | 47 | 0.247
FDA | WAl | 15_15 | 0.6 | 48 | 0.249
FDA | WAl | 15_15 | 0.6 | 49 | 0.25
FDA | WAl | 15_15 | 0.65 | 24 | 0.221
FDA | WAl | 15_15 | 0.65 | 25 | 0.224
FDA | WAl | 15_15 | 0.65 | 26 | 0.226
FDA | WAl | 15_15 | 0.65 | 27 | 0.229
FDA | WAl | 15_15 | 0.65 | 28 | 0.231
FDA | WAl | 15_15 | 0.65 | 29 | 0.234
FDA | WAl | 15_15 | 0.65 | 30 | 0.236
FDA | WAl | 15_15 | 0.65 | 31 | 0.238
FDA | WAl | 15_15 | 0.65 | 32 | 0.24
FDA | WAl | 15_15 | 0.65 | 33 | 0.242
FDA | WAl | 15_15 | 0.65 | 34 | 0.244
FDA | WAl | 15_15 | 0.65 | 35 | 0.245
FDA | WAl | 15_15 | 0.65 | 36 | 0.247
FDA | WAl | 15_15 | 0.65 | 37 | 0.248
FDA | WAl | 15_15 | 0.65 | 38 | 0.25
FDA | WAl | 15_15 | 0.65 | 39 | 0.252
FDA | WAl | 15_15 | 0.65 | 40 | 0.253
FDA | WAl | 15_15 | 0.65 | 41 | 0.254
FDA | WAl | 15_15 | 0.65 | 42 | 0.256
FDA | WAl | 15_15 | 0.65 | 43 | 0.257
FDA | WAl | 15_15 | 0.65 | 44 | 0.258
FDA | WAl | 15_15 | 0.65 | 45 | 0.259
FDA | WAl | 15_15 | 0.65 | 46 | 0.26
FDA | WAl | 15_15 | 0.65 | 47 | 0.262
FDA | WAl | 15_15 | 0.65 | 48 | 0.263
FDA | WAl | 15_15 | 0.65 | 49 | 0.264
FDA | WAl | 15_15 | 0.7 | 25 | 0.24
FDA | WAl | 15_15 | 0.7 | 26 | 0.242
FDA | WAl | 15_15 | 0.7 | 27 | 0.244
FDA | WAl | 15_15 | 0.7 | 28 | 0.247
FDA | WAl | 15_15 | 0.7 | 29 | 0.249
FDA | WAl | 15_15 | 0.7 | 30 | 0.251
FDA | WAl | 15_15 | 0.7 | 31 | 0.253
FDA | WAl | 15_15 | 0.7 | 32 | 0.255
FDA | WAl | 15_15 | 0.7 | 33 | 0.257
FDA | WAl | 15_15 | 0.7 | 34 | 0.258
FDA | WAl | 15_15 | 0.7 | 35 | 0.26
FDA | WAl | 15_15 | 0.7 | 36 | 0.262
FDA | WAl | 15_15 | 0.7 | 37 | 0.263
FDA | WAl | 15_15 | 0.7 | 38 | 0.265
FDA | WAl | 15_15 | 0.7 | 39 | 0.266
FDA | WAl | 15_15 | 0.7 | 40 | 0.267
FDA | WAl | 15_15 | 0.7 | 41 | 0.268
FDA | WAl | 15_15 | 0.7 | 42 | 0.27
FDA | WAl | 15_15 | 0.7 | 43 | 0.271
FDA | WAl | 15_15 | 0.7 | 44 | 0.272
FDA | WAl | 15_15 | 0.7 | 45 | 0.273
FDA | WAl | 15_15 | 0.7 | 46 | 0.274
FDA | WAl | 15_15 | 0.7 | 47 | 0.275
FDA | WAl | 15_15 | 0.7 | 48 | 0.276
FDA | WAl | 15_15 | 0.7 | 49 | 0.277
FDA | WAl | 15_15 | 0.75 | 26 | 0.258
FDA | WAl | 15_15 | 0.75 | 27 | 0.26
FDA | WAl | 15_15 | 0.75 | 28 | 0.262
FDA | WAl | 15_15 | 0.75 | 29 | 0.264
FDA | WAl | 15_15 | 0.75 | 30 | 0.266
FDA | WAl | 15_15 | 0.75 | 31 | 0.268
FDA | WAl | 15_15 | 0.75 | 32 | 0.27
FDA | WAl | 15_15 | 0.75 | 33 | 0.271
FDA | WAl | 15_15 | 0.75 | 34 | 0.273
FDA | WAl | 15_15 | 0.75 | 35 | 0.274
FDA | WAl | 15_15 | 0.75 | 36 | 0.276
FDA | WAl | 15_15 | 0.75 | 37 | 0.277
FDA | WAl | 15_15 | 0.75 | 38 | 0.279
FDA | WAl | 15_15 | 0.75 | 39 | 0.28
FDA | WAl | 15_15 | 0.75 | 40 | 0.281
FDA | WAl | 15_15 | 0.75 | 41 | 0.282
FDA | WAl | 15_15 | 0.75 | 42 | 0.283
FDA | WAl | 15_15 | 0.75 | 43 | 0.285
FDA | WAl | 15_15 | 0.75 | 44 | 0.286
FDA | WAl | 15_15 | 0.75 | 45 | 0.287
FDA | WAl | 15_15 | 0.75 | 46 | 0.288
FDA | WAl | 15_15 | 0.75 | 47 | 0.289
FDA | WAl | 15_15 | 0.75 | 48 | 0.29
FDA | WAl | 15_15 | 0.75 | 49 | 0.29
FDA | WAl | 15_15 | 0.8 | 27 | 0.275
FDA | WAl | 15_15 | 0.8 | 28 | 0.277
FDA | WAl | 15_15 | 0.8 | 29 | 0.279
FDA | WAl | 15_15 | 0.8 | 30 | 0.28
FDA | WAl | 15_15 | 0.8 | 31 | 0.282
FDA | WAl | 15_15 | 0.8 | 32 | 0.284
FDA | WAl | 15_15 | 0.8 | 33 | 0.285
FDA | WAl | 15_15 | 0.8 | 34 | 0.287
FDA | WAl | 15_15 | 0.8 | 35 | 0.288
FDA | WAl | 15_15 | 0.8 | 36 | 0.29
FDA | WAl | 15_15 | 0.8 | 37 | 0.291
FDA | WAl | 15_15 | 0.8 | 38 | 0.292
FDA | WAl | 15_15 | 0.8 | 39 | 0.294
FDA | WAl | 15_15 | 0.8 | 40 | 0.295
FDA | WAl | 15_15 | 0.8 | 41 | 0.296
FDA | WAl | 15_15 | 0.8 | 42 | 0.297
FDA | WAl | 15_15 | 0.8 | 43 | 0.298
FDA | WAl | 15_15 | 0.8 | 44 | 0.299
FDA | WAl | 15_15 | 0.8 | 45 | 0.3
FDA | WAl | 15_15 | 0.8 | 46 | 0.301
FDA | WAl | 15_15 | 0.8 | 47 | 0.301
FDA | WAl | 15_15 | 0.8 | 48 | 0.302
FDA | WAl | 15_15 | 0.8 | 49 | 0.303
FDA | WAl | 15_15 | 0.85 | 28 | 0.291
FDA | WAl | 15_15 | 0.85 | 29 | 0.293
FDA | WAl | 15_15 | 0.85 | 30 | 0.294
FDA | WAl | 15_15 | 0.85 | 31 | 0.296
FDA | WAl | 15_15 | 0.85 | 32 | 0.297
FDA | WAl | 15_15 | 0.85 | 33 | 0.299
FDA | WAl | 15_15 | 0.85 | 34 | 0.3
FDA | WAl | 15_15 | 0.85 | 35 | 0.302
FDA | WAl | 15_15 | 0.85 | 36 | 0.303
FDA | WAl | 15_15 | 0.85 | 37 | 0.304
FDA | WAl | 15_15 | 0.85 | 38 | 0.305
FDA | WAl | 15_15 | 0.85 | 39 | 0.307
FDA | WAl | 15_15 | 0.85 | 40 | 0.308
FDA | WAl | 15_15 | 0.85 | 41 | 0.309
FDA | WAl | 15_15 | 0.85 | 42 | 0.31
FDA | WAl | 15_15 | 0.85 | 43 | 0.311
FDA | WAl | 15_15 | 0.85 | 44 | 0.311
FDA | WAl | 15_15 | 0.85 | 45 | 0.312
FDA | WAl | 15_15 | 0.85 | 46 | 0.313
FDA | WAl | 15_15 | 0.85 | 47 | 0.314
FDA | WAl | 15_15 | 0.85 | 48 | 0.315
FDA | WAl | 15_15 | 0.85 | 49 | 0.316
FDA | WAl | 15_15 | 0.9 | 28 | 0.305
FDA | WAl | 15_15 | 0.9 | 29 | 0.306
FDA | WAl | 15_15 | 0.9 | 30 | 0.308
FDA | WAl | 15_15 | 0.9 | 31 | 0.309
FDA | WAl | 15_15 | 0.9 | 32 | 0.311
FDA | WAl | 15_15 | 0.9 | 33 | 0.312
FDA | WAl | 15_15 | 0.9 | 34 | 0.313
FDA | WAl | 15_15 | 0.9 | 35 | 0.315
FDA | WAl | 15_15 | 0.9 | 36 | 0.316
FDA | WAl | 15_15 | 0.9 | 37 | 0.317
FDA | WAl | 15_15 | 0.9 | 38 | 0.318
FDA | WAl | 15_15 | 0.9 | 39 | 0.319
FDA | WAl | 15_15 | 0.9 | 40 | 0.32
FDA | WAl | 15_15 | 0.9 | 41 | 0.321
FDA | WAl | 15_15 | 0.9 | 42 | 0.322
FDA | WAl | 15_15 | 0.9 | 43 | 0.323
FDA | WAl | 15_15 | 0.9 | 44 | 0.324
FDA | WAl | 15_15 | 0.9 | 45 | 0.325
FDA | WAl | 15_15 | 0.9 | 46 | 0.325
FDA | WAl | 15_15 | 0.9 | 47 | 0.326
FDA | WAl | 15_15 | 0.9 | 48 | 0.327
FDA | WAl | 15_15 | 0.9 | 49 | 0.328
FDA | WAl | 15_15 | 0.95 | 29 | 0.319
FDA | WAl | 15_15 | 0.95 | 30 | 0.321
FDA | WAl | 15_15 | 0.95 | 31 | 0.322
FDA | WAl | 15_15 | 0.95 | 32 | 0.323
FDA | WAl | 15_15 | 0.95 | 33 | 0.325
FDA | WAl | 15_15 | 0.95 | 34 | 0.326
FDA | WAl | 15_15 | 0.95 | 35 | 0.327
FDA | WAl | 15_15 | 0.95 | 36 | 0.328
FDA | WAl | 15_15 | 0.95 | 37 | 0.329
FDA | WAl | 15_15 | 0.95 | 38 | 0.33
FDA | WAl | 15_15 | 0.95 | 39 | 0.331
FDA | WAl | 15_15 | 0.95 | 40 | 0.332
FDA | WAl | 15_15 | 0.95 | 41 | 0.333
FDA | WAl | 15_15 | 0.95 | 42 | 0.334
FDA | WAl | 15_15 | 0.95 | 43 | 0.335
FDA | WAl | 15_15 | 0.95 | 44 | 0.336
FDA | WAl | 15_15 | 0.95 | 45 | 0.336
FDA | WAl | 15_15 | 0.95 | 46 | 0.337
FDA | WAl | 15_15 | 0.95 | 47 | 0.338
FDA | WAl | 15_15 | 0.95 | 48 | 0.339
FDA | WAl | 15_15 | 0.95 | 49 | 0.339
FDA | WAl | 15_15 | 1 | 30 | 0.333
FDA | WAl | 15_15 | 1 | 31 | 0.334
FDA | WAl | 15_15 | 1 | 32 | 0.336
FDA | WAl | 15_15 | 1 | 33 | 0.337
FDA | WAl | 15_15 | 1 | 34 | 0.338
FDA | WAl | 15_15 | 1 | 35 | 0.339
FDA | WAl | 15_15 | 1 | 36 | 0.34
FDA | WAl | 15_15 | 1 | 37 | 0.341
FDA | WAl | 15_15 | 1 | 38 | 0.342
FDA | WAl | 15_15 | 1 | 39 | 0.343
FDA | WAl | 15_15 | 1 | 40 | 0.344
FDA | WAl | 15_15 | 1 | 41 | 0.345
FDA | WAl | 15_15 | 1 | 42 | 0.346
FDA | WAl | 15_15 | 1 | 43 | 0.346
FDA | WAl | 15_15 | 1 | 44 | 0.347
FDA | WAl | 15_15 | 1 | 45 | 0.348
FDA | WAl | 15_15 | 1 | 46 | 0.348
FDA | WAl | 15_15 | 1 | 47 | 0.349
FDA | WAl | 15_15 | 1 | 48 | 0.35
FDA | WAl | 15_15 | 1 | 49 | 0.35
FDA | WAl | 15_15 | 1.05 | 31 | 0.346
FDA | WAl | 15_15 | 1.05 | 32 | 0.348
FDA | WAl | 15_15 | 1.05 | 33 | 0.349
FDA | WAl | 15_15 | 1.05 | 34 | 0.35
FDA | WAl | 15_15 | 1.05 | 35 | 0.351
FDA | WAl | 15_15 | 1.05 | 36 | 0.352
FDA | WAl | 15_15 | 1.05 | 37 | 0.353
FDA | WAl | 15_15 | 1.05 | 38 | 0.354
FDA | WAl | 15_15 | 1.05 | 39 | 0.354
FDA | WAl | 15_15 | 1.05 | 40 | 0.355
FDA | WAl | 15_15 | 1.05 | 41 | 0.356
FDA | WAl | 15_15 | 1.05 | 42 | 0.357
FDA | WAl | 15_15 | 1.05 | 43 | 0.358
FDA | WAl | 15_15 | 1.05 | 44 | 0.358
FDA | WAl | 15_15 | 1.05 | 45 | 0.359
FDA | WAl | 15_15 | 1.05 | 46 | 0.359
FDA | WAl | 15_15 | 1.05 | 47 | 0.36
FDA | WAl | 15_15 | 1.05 | 48 | 0.361
FDA | WAl | 15_15 | 1.05 | 49 | 0.361
FDA | WAl | 15_15 | 1.1 | 32 | 0.359
FDA | WAl | 15_15 | 1.1 | 33 | 0.36
FDA | WAl | 15_15 | 1.1 | 34 | 0.361
FDA | WAl | 15_15 | 1.1 | 35 | 0.362
FDA | WAl | 15_15 | 1.1 | 36 | 0.363
FDA | WAl | 15_15 | 1.1 | 37 | 0.364
FDA | WAl | 15_15 | 1.1 | 38 | 0.365
FDA | WAl | 15_15 | 1.1 | 39 | 0.365
FDA | WAl | 15_15 | 1.1 | 40 | 0.366
FDA | WAl | 15_15 | 1.1 | 41 | 0.367
FDA | WAl | 15_15 | 1.1 | 42 | 0.368
FDA | WAl | 15_15 | 1.1 | 43 | 0.368
FDA | WAl | 15_15 | 1.1 | 44 | 0.369
FDA | WAl | 15_15 | 1.1 | 45 | 0.369
FDA | WAl | 15_15 | 1.1 | 46 | 0.37
FDA | WAl | 15_15 | 1.1 | 47 | 0.371
FDA | WAl | 15_15 | 1.1 | 48 | 0.371
FDA | WAl | 15_15 | 1.1 | 49 | 0.372
FDA | WAl | 15_15 | 1.15 | 33 | 0.371
FDA | WAl | 15_15 | 1.15 | 34 | 0.372
FDA | WAl | 15_15 | 1.15 | 35 | 0.373
FDA | WAl | 15_15 | 1.15 | 36 | 0.374
FDA | WAl | 15_15 | 1.15 | 37 | 0.374
FDA | WAl | 15_15 | 1.15 | 38 | 0.375
FDA | WAl | 15_15 | 1.15 | 39 | 0.376
FDA | WAl | 15_15 | 1.15 | 40 | 0.377
FDA | WAl | 15_15 | 1.15 | 41 | 0.377
FDA | WAl | 15_15 | 1.15 | 42 | 0.378
FDA | WAl | 15_15 | 1.15 | 43 | 0.379
FDA | WAl | 15_15 | 1.15 | 44 | 0.379
FDA | WAl | 15_15 | 1.15 | 45 | 0.38
FDA | WAl | 15_15 | 1.15 | 46 | 0.38
FDA | WAl | 15_15 | 1.15 | 47 | 0.381
FDA | WAl | 15_15 | 1.15 | 48 | 0.381
FDA | WAl | 15_15 | 1.15 | 49 | 0.382
FDA | WAl | 15_15 | 1.2 | 34 | 0.383
FDA | WAl | 15_15 | 1.2 | 35 | 0.383
FDA | WAl | 15_15 | 1.2 | 36 | 0.384
FDA | WAl | 15_15 | 1.2 | 37 | 0.385
FDA | WAl | 15_15 | 1.2 | 38 | 0.386
FDA | WAl | 15_15 | 1.2 | 39 | 0.386
FDA | WAl | 15_15 | 1.2 | 40 | 0.387
FDA | WAl | 15_15 | 1.2 | 41 | 0.387
FDA | WAl | 15_15 | 1.2 | 42 | 0.388
FDA | WAl | 15_15 | 1.2 | 43 | 0.389
FDA | WAl | 15_15 | 1.2 | 44 | 0.389
FDA | WAl | 15_15 | 1.2 | 45 | 0.39
FDA | WAl | 15_15 | 1.2 | 46 | 0.39
FDA | WAl | 15_15 | 1.2 | 47 | 0.391
FDA | WAl | 15_15 | 1.2 | 48 | 0.391
FDA | WAl | 15_15 | 1.2 | 49 | 0.392
FDA | WAl | 15_15 | 1.25 | 36 | 0.394
FDA | WAl | 15_15 | 1.25 | 37 | 0.395
FDA | WAl | 15_15 | 1.25 | 38 | 0.395
FDA | WAl | 15_15 | 1.25 | 39 | 0.396
FDA | WAl | 15_15 | 1.25 | 40 | 0.397
FDA | WAl | 15_15 | 1.25 | 41 | 0.397
FDA | WAl | 15_15 | 1.25 | 42 | 0.398
FDA | WAl | 15_15 | 1.25 | 43 | 0.398
FDA | WAl | 15_15 | 1.25 | 44 | 0.399
FDA | WAl | 15_15 | 1.25 | 45 | 0.399
FDA | WAl | 15_15 | 1.25 | 46 | 0.4
FDA | WAl | 15_15 | 1.25 | 47 | 0.4
FDA | WAl | 15_15 | 1.25 | 48 | 0.401
FDA | WAl | 15_15 | 1.25 | 49 | 0.401
FDA | WAl | 15_15 | 1.3 | 41 | 0.407
FDA | WAl | 15_15 | 1.3 | 42 | 0.407
FDA | WAl | 15_15 | 1.3 | 43 | 0.408
FDA | WAl | 15_15 | 1.3 | 44 | 0.408
FDA | WAl | 15_15 | 1.3 | 45 | 0.409
FDA | WAl | 15_15 | 1.3 | 46 | 0.409
FDA | WAl | 15_15 | 1.3 | 47 | 0.409
FDA | WAl | 15_15 | 1.3 | 48 | 0.41
FDA | WAl | 15_15 | 1.3 | 49 | 0.41
FDA | WAl | 15_15 | 1.35 | 46 | 0.418
FDA | WAl | 15_15 | 1.35 | 47 | 0.418
FDA | WAl | 15_15 | 1.35 | 48 | 0.419
FDA | WAl | 15_15 | 1.35 | 49 | 0.419
FDA | WAl | 30_15 | 0.2 | 20 | 0.048
FDA | WAl | 30_15 | 0.2 | 21 | 0.051
FDA | WAl | 30_15 | 0.25 | 20 | 0.065
FDA | WAl | 30_15 | 0.25 | 21 | 0.069
FDA | WAl | 30_15 | 0.25 | 22 | 0.072
FDA | WAl | 30_15 | 0.25 | 23 | 0.076
FDA | WAl | 30_15 | 0.25 | 24 | 0.079
FDA | WAl | 30_15 | 0.25 | 25 | 0.083
FDA | WAl | 30_15 | 0.3 | 20 | 0.082
FDA | WAl | 30_15 | 0.3 | 21 | 0.086
FDA | WAl | 30_15 | 0.3 | 22 | 0.09
FDA | WAl | 30_15 | 0.3 | 23 | 0.094
FDA | WAl | 30_15 | 0.3 | 24 | 0.098
FDA | WAl | 30_15 | 0.3 | 25 | 0.101
FDA | WAl | 30_15 | 0.3 | 26 | 0.105
FDA | WAl | 30_15 | 0.3 | 27 | 0.108
FDA | WAl | 30_15 | 0.3 | 28 | 0.111
FDA | WAl | 30_15 | 0.3 | 29 | 0.114
FDA | WAl | 30_15 | 0.3 | 30 | 0.117
FDA | WAl | 30_15 | 0.35 | 20 | 0.1
FDA | WAl | 30_15 | 0.35 | 21 | 0.104
FDA | WAl | 30_15 | 0.35 | 22 | 0.108
FDA | WAl | 30_15 | 0.35 | 23 | 0.112
FDA | WAl | 30_15 | 0.35 | 24 | 0.116
FDA | WAl | 30_15 | 0.35 | 25 | 0.119
FDA | WAl | 30_15 | 0.35 | 26 | 0.123
FDA | WAl | 30_15 | 0.35 | 27 | 0.126
FDA | WAl | 30_15 | 0.35 | 28 | 0.129
FDA | WAl | 30_15 | 0.35 | 29 | 0.132
FDA | WAl | 30_15 | 0.35 | 30 | 0.135
FDA | WAl | 30_15 | 0.35 | 31 | 0.138
FDA | WAl | 30_15 | 0.35 | 32 | 0.14
FDA | WAl | 30_15 | 0.35 | 33 | 0.143
FDA | WAl | 30_15 | 0.35 | 34 | 0.145
FDA | WAl | 30_15 | 0.35 | 35 | 0.147
FDA | WAl | 30_15 | 0.4 | 20 | 0.118
FDA | WAl | 30_15 | 0.4 | 21 | 0.122
FDA | WAl | 30_15 | 0.4 | 22 | 0.126
FDA | WAl | 30_15 | 0.4 | 23 | 0.13
FDA | WAl | 30_15 | 0.4 | 24 | 0.134
FDA | WAl | 30_15 | 0.4 | 25 | 0.137
FDA | WAl | 30_15 | 0.4 | 26 | 0.14
FDA | WAl | 30_15 | 0.4 | 27 | 0.144
FDA | WAl | 30_15 | 0.4 | 28 | 0.147
FDA | WAl | 30_15 | 0.4 | 29 | 0.15
FDA | WAl | 30_15 | 0.4 | 30 | 0.153
FDA | WAl | 30_15 | 0.4 | 31 | 0.155
FDA | WAl | 30_15 | 0.4 | 32 | 0.158
FDA | WAl | 30_15 | 0.4 | 33 | 0.16
FDA | WAl | 30_15 | 0.4 | 34 | 0.163
FDA | WAl | 30_15 | 0.4 | 35 | 0.165
FDA | WAl | 30_15 | 0.4 | 36 | 0.167
FDA | WAl | 30_15 | 0.4 | 37 | 0.169
FDA | WAl | 30_15 | 0.4 | 38 | 0.171
FDA | WAl | 30_15 | 0.4 | 39 | 0.173
FDA | WAl | 30_15 | 0.4 | 40 | 0.175
FDA | WAl | 30_15 | 0.45 | 20 | 0.137
FDA | WAl | 30_15 | 0.45 | 21 | 0.141
FDA | WAl | 30_15 | 0.45 | 22 | 0.145
FDA | WAl | 30_15 | 0.45 | 23 | 0.148
FDA | WAl | 30_15 | 0.45 | 24 | 0.152
FDA | WAl | 30_15 | 0.45 | 25 | 0.155
FDA | WAl | 30_15 | 0.45 | 26 | 0.158
FDA | WAl | 30_15 | 0.45 | 27 | 0.161
FDA | WAl | 30_15 | 0.45 | 28 | 0.164
FDA | WAl | 30_15 | 0.45 | 29 | 0.167
FDA | WAl | 30_15 | 0.45 | 30 | 0.17
FDA | WAl | 30_15 | 0.45 | 31 | 0.173
FDA | WAl | 30_15 | 0.45 | 32 | 0.175
FDA | WAl | 30_15 | 0.45 | 33 | 0.177
FDA | WAl | 30_15 | 0.45 | 34 | 0.18
FDA | WAl | 30_15 | 0.45 | 35 | 0.182
FDA | WAl | 30_15 | 0.45 | 36 | 0.184
FDA | WAl | 30_15 | 0.45 | 37 | 0.186
FDA | WAl | 30_15 | 0.45 | 38 | 0.188
FDA | WAl | 30_15 | 0.45 | 39 | 0.19
FDA | WAl | 30_15 | 0.45 | 40 | 0.192
FDA | WAl | 30_15 | 0.45 | 41 | 0.193
FDA | WAl | 30_15 | 0.45 | 42 | 0.195
FDA | WAl | 30_15 | 0.45 | 43 | 0.197
FDA | WAl | 30_15 | 0.45 | 44 | 0.198
FDA | WAl | 30_15 | 0.5 | 21 | 0.159
FDA | WAl | 30_15 | 0.5 | 22 | 0.163
FDA | WAl | 30_15 | 0.5 | 23 | 0.166
FDA | WAl | 30_15 | 0.5 | 24 | 0.17
FDA | WAl | 30_15 | 0.5 | 25 | 0.173
FDA | WAl | 30_15 | 0.5 | 26 | 0.176
FDA | WAl | 30_15 | 0.5 | 27 | 0.179
FDA | WAl | 30_15 | 0.5 | 28 | 0.182
FDA | WAl | 30_15 | 0.5 | 29 | 0.184
FDA | WAl | 30_15 | 0.5 | 30 | 0.187
FDA | WAl | 30_15 | 0.5 | 31 | 0.189
FDA | WAl | 30_15 | 0.5 | 32 | 0.192
FDA | WAl | 30_15 | 0.5 | 33 | 0.194
FDA | WAl | 30_15 | 0.5 | 34 | 0.196
FDA | WAl | 30_15 | 0.5 | 35 | 0.198
FDA | WAl | 30_15 | 0.5 | 36 | 0.2
FDA | WAl | 30_15 | 0.5 | 37 | 0.202
FDA | WAl | 30_15 | 0.5 | 38 | 0.204
FDA | WAl | 30_15 | 0.5 | 39 | 0.206
FDA | WAl | 30_15 | 0.5 | 40 | 0.207
FDA | WAl | 30_15 | 0.5 | 41 | 0.209
FDA | WAl | 30_15 | 0.5 | 42 | 0.211
FDA | WAl | 30_15 | 0.5 | 43 | 0.212
FDA | WAl | 30_15 | 0.5 | 44 | 0.214
FDA | WAl | 30_15 | 0.5 | 45 | 0.215
FDA | WAl | 30_15 | 0.5 | 46 | 0.217
FDA | WAl | 30_15 | 0.5 | 47 | 0.218
FDA | WAl | 30_15 | 0.5 | 48 | 0.219
FDA | WAl | 30_15 | 0.5 | 49 | 0.221
FDA | WAl | 30_15 | 0.55 | 22 | 0.181
FDA | WAl | 30_15 | 0.55 | 23 | 0.184
FDA | WAl | 30_15 | 0.55 | 24 | 0.187
FDA | WAl | 30_15 | 0.55 | 25 | 0.19
FDA | WAl | 30_15 | 0.55 | 26 | 0.193
FDA | WAl | 30_15 | 0.55 | 27 | 0.196
FDA | WAl | 30_15 | 0.55 | 28 | 0.199
FDA | WAl | 30_15 | 0.55 | 29 | 0.201
FDA | WAl | 30_15 | 0.55 | 30 | 0.204
FDA | WAl | 30_15 | 0.55 | 31 | 0.206
FDA | WAl | 30_15 | 0.55 | 32 | 0.208
FDA | WAl | 30_15 | 0.55 | 33 | 0.21
FDA | WAl | 30_15 | 0.55 | 34 | 0.212
FDA | WAl | 30_15 | 0.55 | 35 | 0.214
FDA | WAl | 30_15 | 0.55 | 36 | 0.216
FDA | WAl | 30_15 | 0.55 | 37 | 0.218
FDA | WAl | 30_15 | 0.55 | 38 | 0.22
FDA | WAl | 30_15 | 0.55 | 39 | 0.221
FDA | WAl | 30_15 | 0.55 | 40 | 0.223
FDA | WAl | 30_15 | 0.55 | 41 | 0.225
FDA | WAl | 30_15 | 0.55 | 42 | 0.226
FDA | WAl | 30_15 | 0.55 | 43 | 0.228
FDA | WAl | 30_15 | 0.55 | 44 | 0.229
FDA | WAl | 30_15 | 0.55 | 45 | 0.23
FDA | WAl | 30_15 | 0.55 | 46 | 0.232
FDA | WAl | 30_15 | 0.55 | 47 | 0.233
FDA | WAl | 30_15 | 0.55 | 48 | 0.234
FDA | WAl | 30_15 | 0.55 | 49 | 0.236
FDA | WAl | 30_15 | 0.6 | 23 | 0.202
FDA | WAl | 30_15 | 0.6 | 24 | 0.205
FDA | WAl | 30_15 | 0.6 | 25 | 0.207
FDA | WAl | 30_15 | 0.6 | 26 | 0.21
FDA | WAl | 30_15 | 0.6 | 27 | 0.212
FDA | WAl | 30_15 | 0.6 | 28 | 0.215
FDA | WAl | 30_15 | 0.6 | 29 | 0.218
FDA | WAl | 30_15 | 0.6 | 30 | 0.22
FDA | WAl | 30_15 | 0.6 | 31 | 0.222
FDA | WAl | 30_15 | 0.6 | 32 | 0.224
FDA | WAl | 30_15 | 0.6 | 33 | 0.226
FDA | WAl | 30_15 | 0.6 | 34 | 0.228
FDA | WAl | 30_15 | 0.6 | 35 | 0.23
FDA | WAl | 30_15 | 0.6 | 36 | 0.232
FDA | WAl | 30_15 | 0.6 | 37 | 0.233
FDA | WAl | 30_15 | 0.6 | 38 | 0.235
FDA | WAl | 30_15 | 0.6 | 39 | 0.237
FDA | WAl | 30_15 | 0.6 | 40 | 0.238
FDA | WAl | 30_15 | 0.6 | 41 | 0.24
FDA | WAl | 30_15 | 0.6 | 42 | 0.241
FDA | WAl | 30_15 | 0.6 | 43 | 0.243
FDA | WAl | 30_15 | 0.6 | 44 | 0.244
FDA | WAl | 30_15 | 0.6 | 45 | 0.245
FDA | WAl | 30_15 | 0.6 | 46 | 0.246
FDA | WAl | 30_15 | 0.6 | 47 | 0.247
FDA | WAl | 30_15 | 0.6 | 48 | 0.249
FDA | WAl | 30_15 | 0.6 | 49 | 0.25
FDA | WAl | 30_15 | 0.65 | 24 | 0.221
FDA | WAl | 30_15 | 0.65 | 25 | 0.224
FDA | WAl | 30_15 | 0.65 | 26 | 0.226
FDA | WAl | 30_15 | 0.65 | 27 | 0.229
FDA | WAl | 30_15 | 0.65 | 28 | 0.231
FDA | WAl | 30_15 | 0.65 | 29 | 0.234
FDA | WAl | 30_15 | 0.65 | 30 | 0.236
FDA | WAl | 30_15 | 0.65 | 31 | 0.238
FDA | WAl | 30_15 | 0.65 | 32 | 0.24
FDA | WAl | 30_15 | 0.65 | 33 | 0.242
FDA | WAl | 30_15 | 0.65 | 34 | 0.244
FDA | WAl | 30_15 | 0.65 | 35 | 0.245
FDA | WAl | 30_15 | 0.65 | 36 | 0.247
FDA | WAl | 30_15 | 0.65 | 37 | 0.248
FDA | WAl | 30_15 | 0.65 | 38 | 0.25
FDA | WAl | 30_15 | 0.65 | 39 | 0.252
FDA | WAl | 30_15 | 0.65 | 40 | 0.253
FDA | WAl | 30_15 | 0.65 | 41 | 0.254
FDA | WAl | 30_15 | 0.65 | 42 | 0.256
FDA | WAl | 30_15 | 0.65 | 43 | 0.257
FDA | WAl | 30_15 | 0.65 | 44 | 0.258
FDA | WAl | 30_15 | 0.65 | 45 | 0.259
FDA | WAl | 30_15 | 0.65 | 46 | 0.261
FDA | WAl | 30_15 | 0.65 | 47 | 0.262
FDA | WAl | 30_15 | 0.65 | 48 | 0.263
FDA | WAl | 30_15 | 0.65 | 49 | 0.264
FDA | WAl | 30_15 | 0.7 | 25 | 0.24
FDA | WAl | 30_15 | 0.7 | 26 | 0.242
FDA | WAl | 30_15 | 0.7 | 27 | 0.244
FDA | WAl | 30_15 | 0.7 | 28 | 0.247
FDA | WAl | 30_15 | 0.7 | 29 | 0.249
FDA | WAl | 30_15 | 0.7 | 30 | 0.251
FDA | WAl | 30_15 | 0.7 | 31 | 0.253
FDA | WAl | 30_15 | 0.7 | 32 | 0.255
FDA | WAl | 30_15 | 0.7 | 33 | 0.257
FDA | WAl | 30_15 | 0.7 | 34 | 0.259
FDA | WAl | 30_15 | 0.7 | 35 | 0.26
FDA | WAl | 30_15 | 0.7 | 36 | 0.262
FDA | WAl | 30_15 | 0.7 | 37 | 0.263
FDA | WAl | 30_15 | 0.7 | 38 | 0.265
FDA | WAl | 30_15 | 0.7 | 39 | 0.266
FDA | WAl | 30_15 | 0.7 | 40 | 0.267
FDA | WAl | 30_15 | 0.7 | 41 | 0.269
FDA | WAl | 30_15 | 0.7 | 42 | 0.27
FDA | WAl | 30_15 | 0.7 | 43 | 0.271
FDA | WAl | 30_15 | 0.7 | 44 | 0.272
FDA | WAl | 30_15 | 0.7 | 45 | 0.273
FDA | WAl | 30_15 | 0.7 | 46 | 0.274
FDA | WAl | 30_15 | 0.7 | 47 | 0.275
FDA | WAl | 30_15 | 0.7 | 48 | 0.276
FDA | WAl | 30_15 | 0.7 | 49 | 0.277
FDA | WAl | 30_15 | 0.75 | 26 | 0.258
FDA | WAl | 30_15 | 0.75 | 27 | 0.26
FDA | WAl | 30_15 | 0.75 | 28 | 0.262
FDA | WAl | 30_15 | 0.75 | 29 | 0.264
FDA | WAl | 30_15 | 0.75 | 30 | 0.266
FDA | WAl | 30_15 | 0.75 | 31 | 0.268
FDA | WAl | 30_15 | 0.75 | 32 | 0.27
FDA | WAl | 30_15 | 0.75 | 33 | 0.271
FDA | WAl | 30_15 | 0.75 | 34 | 0.273
FDA | WAl | 30_15 | 0.75 | 35 | 0.274
FDA | WAl | 30_15 | 0.75 | 36 | 0.276
FDA | WAl | 30_15 | 0.75 | 37 | 0.277
FDA | WAl | 30_15 | 0.75 | 38 | 0.279
FDA | WAl | 30_15 | 0.75 | 39 | 0.28
FDA | WAl | 30_15 | 0.75 | 40 | 0.281
FDA | WAl | 30_15 | 0.75 | 41 | 0.282
FDA | WAl | 30_15 | 0.75 | 42 | 0.283
FDA | WAl | 30_15 | 0.75 | 43 | 0.285
FDA | WAl | 30_15 | 0.75 | 44 | 0.286
FDA | WAl | 30_15 | 0.75 | 45 | 0.287
FDA | WAl | 30_15 | 0.75 | 46 | 0.288
FDA | WAl | 30_15 | 0.75 | 47 | 0.289
FDA | WAl | 30_15 | 0.75 | 48 | 0.29
FDA | WAl | 30_15 | 0.75 | 49 | 0.29
FDA | WAl | 30_15 | 0.8 | 27 | 0.275
FDA | WAl | 30_15 | 0.8 | 28 | 0.277
FDA | WAl | 30_15 | 0.8 | 29 | 0.279
FDA | WAl | 30_15 | 0.8 | 30 | 0.28
FDA | WAl | 30_15 | 0.8 | 31 | 0.282
FDA | WAl | 30_15 | 0.8 | 32 | 0.284
FDA | WAl | 30_15 | 0.8 | 33 | 0.285
FDA | WAl | 30_15 | 0.8 | 34 | 0.287
FDA | WAl | 30_15 | 0.8 | 35 | 0.288
FDA | WAl | 30_15 | 0.8 | 36 | 0.29
FDA | WAl | 30_15 | 0.8 | 37 | 0.291
FDA | WAl | 30_15 | 0.8 | 38 | 0.292
FDA | WAl | 30_15 | 0.8 | 39 | 0.294
FDA | WAl | 30_15 | 0.8 | 40 | 0.295
FDA | WAl | 30_15 | 0.8 | 41 | 0.296
FDA | WAl | 30_15 | 0.8 | 42 | 0.297
FDA | WAl | 30_15 | 0.8 | 43 | 0.298
FDA | WAl | 30_15 | 0.8 | 44 | 0.299
FDA | WAl | 30_15 | 0.8 | 45 | 0.3
FDA | WAl | 30_15 | 0.8 | 46 | 0.301
FDA | WAl | 30_15 | 0.8 | 47 | 0.301
FDA | WAl | 30_15 | 0.8 | 48 | 0.302
FDA | WAl | 30_15 | 0.8 | 49 | 0.303
FDA | WAl | 30_15 | 0.85 | 28 | 0.291
FDA | WAl | 30_15 | 0.85 | 29 | 0.293
FDA | WAl | 30_15 | 0.85 | 30 | 0.294
FDA | WAl | 30_15 | 0.85 | 31 | 0.296
FDA | WAl | 30_15 | 0.85 | 32 | 0.297
FDA | WAl | 30_15 | 0.85 | 33 | 0.299
FDA | WAl | 30_15 | 0.85 | 34 | 0.3
FDA | WAl | 30_15 | 0.85 | 35 | 0.302
FDA | WAl | 30_15 | 0.85 | 36 | 0.303
FDA | WAl | 30_15 | 0.85 | 37 | 0.304
FDA | WAl | 30_15 | 0.85 | 38 | 0.305
FDA | WAl | 30_15 | 0.85 | 39 | 0.307
FDA | WAl | 30_15 | 0.85 | 40 | 0.308
FDA | WAl | 30_15 | 0.85 | 41 | 0.309
FDA | WAl | 30_15 | 0.85 | 42 | 0.31
FDA | WAl | 30_15 | 0.85 | 43 | 0.311
FDA | WAl | 30_15 | 0.85 | 44 | 0.311
FDA | WAl | 30_15 | 0.85 | 45 | 0.312
FDA | WAl | 30_15 | 0.85 | 46 | 0.313
FDA | WAl | 30_15 | 0.85 | 47 | 0.314
FDA | WAl | 30_15 | 0.85 | 48 | 0.315
FDA | WAl | 30_15 | 0.85 | 49 | 0.316
FDA | WAl | 30_15 | 0.9 | 28 | 0.305
FDA | WAl | 30_15 | 0.9 | 29 | 0.306
FDA | WAl | 30_15 | 0.9 | 30 | 0.308
FDA | WAl | 30_15 | 0.9 | 31 | 0.309
FDA | WAl | 30_15 | 0.9 | 32 | 0.311
FDA | WAl | 30_15 | 0.9 | 33 | 0.312
FDA | WAl | 30_15 | 0.9 | 34 | 0.313
FDA | WAl | 30_15 | 0.9 | 35 | 0.315
FDA | WAl | 30_15 | 0.9 | 36 | 0.316
FDA | WAl | 30_15 | 0.9 | 37 | 0.317
FDA | WAl | 30_15 | 0.9 | 38 | 0.318
FDA | WAl | 30_15 | 0.9 | 39 | 0.319
FDA | WAl | 30_15 | 0.9 | 40 | 0.32
FDA | WAl | 30_15 | 0.9 | 41 | 0.321
FDA | WAl | 30_15 | 0.9 | 42 | 0.322
FDA | WAl | 30_15 | 0.9 | 43 | 0.323
FDA | WAl | 30_15 | 0.9 | 44 | 0.324
FDA | WAl | 30_15 | 0.9 | 45 | 0.325
FDA | WAl | 30_15 | 0.9 | 46 | 0.325
FDA | WAl | 30_15 | 0.9 | 47 | 0.326
FDA | WAl | 30_15 | 0.9 | 48 | 0.327
FDA | WAl | 30_15 | 0.9 | 49 | 0.328
FDA | WAl | 30_15 | 0.95 | 29 | 0.319
FDA | WAl | 30_15 | 0.95 | 30 | 0.321
FDA | WAl | 30_15 | 0.95 | 31 | 0.322
FDA | WAl | 30_15 | 0.95 | 32 | 0.323
FDA | WAl | 30_15 | 0.95 | 33 | 0.325
FDA | WAl | 30_15 | 0.95 | 34 | 0.326
FDA | WAl | 30_15 | 0.95 | 35 | 0.327
FDA | WAl | 30_15 | 0.95 | 36 | 0.328
FDA | WAl | 30_15 | 0.95 | 37 | 0.329
FDA | WAl | 30_15 | 0.95 | 38 | 0.33
FDA | WAl | 30_15 | 0.95 | 39 | 0.331
FDA | WAl | 30_15 | 0.95 | 40 | 0.332
FDA | WAl | 30_15 | 0.95 | 41 | 0.333
FDA | WAl | 30_15 | 0.95 | 42 | 0.334
FDA | WAl | 30_15 | 0.95 | 43 | 0.335
FDA | WAl | 30_15 | 0.95 | 44 | 0.336
FDA | WAl | 30_15 | 0.95 | 45 | 0.336
FDA | WAl | 30_15 | 0.95 | 46 | 0.337
FDA | WAl | 30_15 | 0.95 | 47 | 0.338
FDA | WAl | 30_15 | 0.95 | 48 | 0.339
FDA | WAl | 30_15 | 0.95 | 49 | 0.339
FDA | WAl | 30_15 | 1 | 30 | 0.333
FDA | WAl | 30_15 | 1 | 31 | 0.334
FDA | WAl | 30_15 | 1 | 32 | 0.336
FDA | WAl | 30_15 | 1 | 33 | 0.337
FDA | WAl | 30_15 | 1 | 34 | 0.338
FDA | WAl | 30_15 | 1 | 35 | 0.339
FDA | WAl | 30_15 | 1 | 36 | 0.34
FDA | WAl | 30_15 | 1 | 37 | 0.341
FDA | WAl | 30_15 | 1 | 38 | 0.342
FDA | WAl | 30_15 | 1 | 39 | 0.343
FDA | WAl | 30_15 | 1 | 40 | 0.344
FDA | WAl | 30_15 | 1 | 41 | 0.345
FDA | WAl | 30_15 | 1 | 42 | 0.346
FDA | WAl | 30_15 | 1 | 43 | 0.346
FDA | WAl | 30_15 | 1 | 44 | 0.347
FDA | WAl | 30_15 | 1 | 45 | 0.348
FDA | WAl | 30_15 | 1 | 46 | 0.348
FDA | WAl | 30_15 | 1 | 47 | 0.349
FDA | WAl | 30_15 | 1 | 48 | 0.35
FDA | WAl | 30_15 | 1 | 49 | 0.35
FDA | WAl | 30_15 | 1.05 | 31 | 0.346
FDA | WAl | 30_15 | 1.05 | 32 | 0.348
FDA | WAl | 30_15 | 1.05 | 33 | 0.349
FDA | WAl | 30_15 | 1.05 | 34 | 0.35
FDA | WAl | 30_15 | 1.05 | 35 | 0.351
FDA | WAl | 30_15 | 1.05 | 36 | 0.352
FDA | WAl | 30_15 | 1.05 | 37 | 0.353
FDA | WAl | 30_15 | 1.05 | 38 | 0.354
FDA | WAl | 30_15 | 1.05 | 39 | 0.354
FDA | WAl | 30_15 | 1.05 | 40 | 0.355
FDA | WAl | 30_15 | 1.05 | 41 | 0.356
FDA | WAl | 30_15 | 1.05 | 42 | 0.357
FDA | WAl | 30_15 | 1.05 | 43 | 0.358
FDA | WAl | 30_15 | 1.05 | 44 | 0.358
FDA | WAl | 30_15 | 1.05 | 45 | 0.359
FDA | WAl | 30_15 | 1.05 | 46 | 0.359
FDA | WAl | 30_15 | 1.05 | 47 | 0.36
FDA | WAl | 30_15 | 1.05 | 48 | 0.361
FDA | WAl | 30_15 | 1.05 | 49 | 0.361
FDA | WAl | 30_15 | 1.1 | 32 | 0.359
FDA | WAl | 30_15 | 1.1 | 33 | 0.36
FDA | WAl | 30_15 | 1.1 | 34 | 0.361
FDA | WAl | 30_15 | 1.1 | 35 | 0.362
FDA | WAl | 30_15 | 1.1 | 36 | 0.363
FDA | WAl | 30_15 | 1.1 | 37 | 0.364
FDA | WAl | 30_15 | 1.1 | 38 | 0.365
FDA | WAl | 30_15 | 1.1 | 39 | 0.365
FDA | WAl | 30_15 | 1.1 | 40 | 0.366
FDA | WAl | 30_15 | 1.1 | 41 | 0.367
FDA | WAl | 30_15 | 1.1 | 42 | 0.368
FDA | WAl | 30_15 | 1.1 | 43 | 0.368
FDA | WAl | 30_15 | 1.1 | 44 | 0.369
FDA | WAl | 30_15 | 1.1 | 45 | 0.369
FDA | WAl | 30_15 | 1.1 | 46 | 0.37
FDA | WAl | 30_15 | 1.1 | 47 | 0.371
FDA | WAl | 30_15 | 1.1 | 48 | 0.371
FDA | WAl | 30_15 | 1.1 | 49 | 0.372
FDA | WAl | 30_15 | 1.15 | 33 | 0.371
FDA | WAl | 30_15 | 1.15 | 34 | 0.372
FDA | WAl | 30_15 | 1.15 | 35 | 0.373
FDA | WAl | 30_15 | 1.15 | 36 | 0.374
FDA | WAl | 30_15 | 1.15 | 37 | 0.374
FDA | WAl | 30_15 | 1.15 | 38 | 0.375
FDA | WAl | 30_15 | 1.15 | 39 | 0.376
FDA | WAl | 30_15 | 1.15 | 40 | 0.377
FDA | WAl | 30_15 | 1.15 | 41 | 0.377
FDA | WAl | 30_15 | 1.15 | 42 | 0.378
FDA | WAl | 30_15 | 1.15 | 43 | 0.379
FDA | WAl | 30_15 | 1.15 | 44 | 0.379
FDA | WAl | 30_15 | 1.15 | 45 | 0.38
FDA | WAl | 30_15 | 1.15 | 46 | 0.38
FDA | WAl | 30_15 | 1.15 | 47 | 0.381
FDA | WAl | 30_15 | 1.15 | 48 | 0.381
FDA | WAl | 30_15 | 1.15 | 49 | 0.382
FDA | WAl | 30_15 | 1.2 | 34 | 0.383
FDA | WAl | 30_15 | 1.2 | 35 | 0.383
FDA | WAl | 30_15 | 1.2 | 36 | 0.384
FDA | WAl | 30_15 | 1.2 | 37 | 0.385
FDA | WAl | 30_15 | 1.2 | 38 | 0.386
FDA | WAl | 30_15 | 1.2 | 39 | 0.386
FDA | WAl | 30_15 | 1.2 | 40 | 0.387
FDA | WAl | 30_15 | 1.2 | 41 | 0.387
FDA | WAl | 30_15 | 1.2 | 42 | 0.388
FDA | WAl | 30_15 | 1.2 | 43 | 0.389
FDA | WAl | 30_15 | 1.2 | 44 | 0.389
FDA | WAl | 30_15 | 1.2 | 45 | 0.39
FDA | WAl | 30_15 | 1.2 | 46 | 0.39
FDA | WAl | 30_15 | 1.2 | 47 | 0.391
FDA | WAl | 30_15 | 1.2 | 48 | 0.391
FDA | WAl | 30_15 | 1.2 | 49 | 0.392
FDA | WAl | 30_15 | 1.25 | 36 | 0.394
FDA | WAl | 30_15 | 1.25 | 37 | 0.395
FDA | WAl | 30_15 | 1.25 | 38 | 0.395
FDA | WAl | 30_15 | 1.25 | 39 | 0.396
FDA | WAl | 30_15 | 1.25 | 40 | 0.397
FDA | WAl | 30_15 | 1.25 | 41 | 0.397
FDA | WAl | 30_15 | 1.25 | 42 | 0.398
FDA | WAl | 30_15 | 1.25 | 43 | 0.398
FDA | WAl | 30_15 | 1.25 | 44 | 0.399
FDA | WAl | 30_15 | 1.25 | 45 | 0.399
FDA | WAl | 30_15 | 1.25 | 46 | 0.4
FDA | WAl | 30_15 | 1.25 | 47 | 0.4
FDA | WAl | 30_15 | 1.25 | 48 | 0.401
FDA | WAl | 30_15 | 1.25 | 49 | 0.401
FDA | WAl | 30_15 | 1.3 | 41 | 0.407
FDA | WAl | 30_15 | 1.3 | 42 | 0.407
FDA | WAl | 30_15 | 1.3 | 43 | 0.408
FDA | WAl | 30_15 | 1.3 | 44 | 0.408
FDA | WAl | 30_15 | 1.3 | 45 | 0.409
FDA | WAl | 30_15 | 1.3 | 46 | 0.409
FDA | WAl | 30_15 | 1.3 | 47 | 0.409
FDA | WAl | 30_15 | 1.3 | 48 | 0.41
FDA | WAl | 30_15 | 1.3 | 49 | 0.41
FDA | WAl | 30_15 | 1.35 | 46 | 0.418
FDA | WAl | 30_15 | 1.35 | 47 | 0.418
FDA | WAl | 30_15 | 1.35 | 48 | 0.419
FDA | WAl | 30_15 | 1.35 | 49 | 0.419
FDA | WAl | 25_50 | 0.2 | 20 | 0.046
FDA | WAl | 25_50 | 0.2 | 21 | 0.049
FDA | WAl | 25_50 | 0.25 | 20 | 0.063
FDA | WAl | 25_50 | 0.25 | 21 | 0.066
FDA | WAl | 25_50 | 0.25 | 22 | 0.07
FDA | WAl | 25_50 | 0.25 | 23 | 0.074
FDA | WAl | 25_50 | 0.25 | 24 | 0.077
FDA | WAl | 25_50 | 0.25 | 25 | 0.08
FDA | WAl | 25_50 | 0.3 | 20 | 0.079
FDA | WAl | 25_50 | 0.3 | 21 | 0.084
FDA | WAl | 25_50 | 0.3 | 22 | 0.088
FDA | WAl | 25_50 | 0.3 | 23 | 0.092
FDA | WAl | 25_50 | 0.3 | 24 | 0.095
FDA | WAl | 25_50 | 0.3 | 25 | 0.099
FDA | WAl | 25_50 | 0.3 | 26 | 0.102
FDA | WAl | 25_50 | 0.3 | 27 | 0.105
FDA | WAl | 25_50 | 0.3 | 28 | 0.108
FDA | WAl | 25_50 | 0.3 | 29 | 0.111
FDA | WAl | 25_50 | 0.3 | 30 | 0.114
FDA | WAl | 25_50 | 0.35 | 20 | 0.097
FDA | WAl | 25_50 | 0.35 | 21 | 0.101
FDA | WAl | 25_50 | 0.35 | 22 | 0.105
FDA | WAl | 25_50 | 0.35 | 23 | 0.109
FDA | WAl | 25_50 | 0.35 | 24 | 0.113
FDA | WAl | 25_50 | 0.35 | 25 | 0.117
FDA | WAl | 25_50 | 0.35 | 26 | 0.12
FDA | WAl | 25_50 | 0.35 | 27 | 0.123
FDA | WAl | 25_50 | 0.35 | 28 | 0.126
FDA | WAl | 25_50 | 0.35 | 29 | 0.13
FDA | WAl | 25_50 | 0.35 | 30 | 0.132
FDA | WAl | 25_50 | 0.35 | 31 | 0.135
FDA | WAl | 25_50 | 0.35 | 32 | 0.138
FDA | WAl | 25_50 | 0.35 | 33 | 0.14
FDA | WAl | 25_50 | 0.35 | 34 | 0.143
FDA | WAl | 25_50 | 0.35 | 35 | 0.145
FDA | WAl | 25_50 | 0.4 | 20 | 0.115
FDA | WAl | 25_50 | 0.4 | 21 | 0.119
FDA | WAl | 25_50 | 0.4 | 22 | 0.123
FDA | WAl | 25_50 | 0.4 | 23 | 0.127
FDA | WAl | 25_50 | 0.4 | 24 | 0.131
FDA | WAl | 25_50 | 0.4 | 25 | 0.134
FDA | WAl | 25_50 | 0.4 | 26 | 0.138
FDA | WAl | 25_50 | 0.4 | 27 | 0.141
FDA | WAl | 25_50 | 0.4 | 28 | 0.144
FDA | WAl | 25_50 | 0.4 | 29 | 0.147
FDA | WAl | 25_50 | 0.4 | 30 | 0.15
FDA | WAl | 25_50 | 0.4 | 31 | 0.153
FDA | WAl | 25_50 | 0.4 | 32 | 0.155
FDA | WAl | 25_50 | 0.4 | 33 | 0.158
FDA | WAl | 25_50 | 0.4 | 34 | 0.16
FDA | WAl | 25_50 | 0.4 | 35 | 0.163
FDA | WAl | 25_50 | 0.4 | 36 | 0.165
FDA | WAl | 25_50 | 0.4 | 37 | 0.167
FDA | WAl | 25_50 | 0.4 | 38 | 0.169
FDA | WAl | 25_50 | 0.4 | 39 | 0.171
FDA | WAl | 25_50 | 0.4 | 40 | 0.173
FDA | WAl | 25_50 | 0.45 | 20 | 0.133
FDA | WAl | 25_50 | 0.45 | 21 | 0.137
FDA | WAl | 25_50 | 0.45 | 22 | 0.141
FDA | WAl | 25_50 | 0.45 | 23 | 0.145
FDA | WAl | 25_50 | 0.45 | 24 | 0.149
FDA | WAl | 25_50 | 0.45 | 25 | 0.152
FDA | WAl | 25_50 | 0.45 | 26 | 0.155
FDA | WAl | 25_50 | 0.45 | 27 | 0.158
FDA | WAl | 25_50 | 0.45 | 28 | 0.161
FDA | WAl | 25_50 | 0.45 | 29 | 0.164
FDA | WAl | 25_50 | 0.45 | 30 | 0.167
FDA | WAl | 25_50 | 0.45 | 31 | 0.17
FDA | WAl | 25_50 | 0.45 | 32 | 0.172
FDA | WAl | 25_50 | 0.45 | 33 | 0.175
FDA | WAl | 25_50 | 0.45 | 34 | 0.177
FDA | WAl | 25_50 | 0.45 | 35 | 0.179
FDA | WAl | 25_50 | 0.45 | 36 | 0.181
FDA | WAl | 25_50 | 0.45 | 37 | 0.184
FDA | WAl | 25_50 | 0.45 | 38 | 0.186
FDA | WAl | 25_50 | 0.45 | 39 | 0.188
FDA | WAl | 25_50 | 0.45 | 40 | 0.189
FDA | WAl | 25_50 | 0.45 | 41 | 0.191
FDA | WAl | 25_50 | 0.45 | 42 | 0.193
FDA | WAl | 25_50 | 0.45 | 43 | 0.195
FDA | WAl | 25_50 | 0.45 | 44 | 0.196
FDA | WAl | 25_50 | 0.5 | 21 | 0.156
FDA | WAl | 25_50 | 0.5 | 22 | 0.159
FDA | WAl | 25_50 | 0.5 | 23 | 0.163
FDA | WAl | 25_50 | 0.5 | 24 | 0.166
FDA | WAl | 25_50 | 0.5 | 25 | 0.169
FDA | WAl | 25_50 | 0.5 | 26 | 0.172
FDA | WAl | 25_50 | 0.5 | 27 | 0.175
FDA | WAl | 25_50 | 0.5 | 28 | 0.179
FDA | WAl | 25_50 | 0.5 | 29 | 0.181
FDA | WAl | 25_50 | 0.5 | 30 | 0.184
FDA | WAl | 25_50 | 0.5 | 31 | 0.187
FDA | WAl | 25_50 | 0.5 | 32 | 0.189
FDA | WAl | 25_50 | 0.5 | 33 | 0.191
FDA | WAl | 25_50 | 0.5 | 34 | 0.194
FDA | WAl | 25_50 | 0.5 | 35 | 0.196
FDA | WAl | 25_50 | 0.5 | 36 | 0.198
FDA | WAl | 25_50 | 0.5 | 37 | 0.2
FDA | WAl | 25_50 | 0.5 | 38 | 0.202
FDA | WAl | 25_50 | 0.5 | 39 | 0.203
FDA | WAl | 25_50 | 0.5 | 40 | 0.205
FDA | WAl | 25_50 | 0.5 | 41 | 0.207
FDA | WAl | 25_50 | 0.5 | 42 | 0.209
FDA | WAl | 25_50 | 0.5 | 43 | 0.21
FDA | WAl | 25_50 | 0.5 | 44 | 0.212
FDA | WAl | 25_50 | 0.5 | 45 | 0.213
FDA | WAl | 25_50 | 0.5 | 46 | 0.215
FDA | WAl | 25_50 | 0.5 | 47 | 0.216
FDA | WAl | 25_50 | 0.5 | 48 | 0.218
FDA | WAl | 25_50 | 0.5 | 49 | 0.219
FDA | WAl | 25_50 | 0.55 | 22 | 0.177
FDA | WAl | 25_50 | 0.55 | 23 | 0.18
FDA | WAl | 25_50 | 0.55 | 24 | 0.184
FDA | WAl | 25_50 | 0.55 | 25 | 0.187
FDA | WAl | 25_50 | 0.55 | 26 | 0.19
FDA | WAl | 25_50 | 0.55 | 27 | 0.192
FDA | WAl | 25_50 | 0.55 | 28 | 0.195
FDA | WAl | 25_50 | 0.55 | 29 | 0.198
FDA | WAl | 25_50 | 0.55 | 30 | 0.201
FDA | WAl | 25_50 | 0.55 | 31 | 0.203
FDA | WAl | 25_50 | 0.55 | 32 | 0.205
FDA | WAl | 25_50 | 0.55 | 33 | 0.208
FDA | WAl | 25_50 | 0.55 | 34 | 0.21
FDA | WAl | 25_50 | 0.55 | 35 | 0.212
FDA | WAl | 25_50 | 0.55 | 36 | 0.214
FDA | WAl | 25_50 | 0.55 | 37 | 0.215
FDA | WAl | 25_50 | 0.55 | 38 | 0.217
FDA | WAl | 25_50 | 0.55 | 39 | 0.219
FDA | WAl | 25_50 | 0.55 | 40 | 0.221
FDA | WAl | 25_50 | 0.55 | 41 | 0.222
FDA | WAl | 25_50 | 0.55 | 42 | 0.224
FDA | WAl | 25_50 | 0.55 | 43 | 0.225
FDA | WAl | 25_50 | 0.55 | 44 | 0.227
FDA | WAl | 25_50 | 0.55 | 45 | 0.228
FDA | WAl | 25_50 | 0.55 | 46 | 0.23
FDA | WAl | 25_50 | 0.55 | 47 | 0.231
FDA | WAl | 25_50 | 0.55 | 48 | 0.232
FDA | WAl | 25_50 | 0.55 | 49 | 0.234
FDA | WAl | 25_50 | 0.6 | 23 | 0.198
FDA | WAl | 25_50 | 0.6 | 24 | 0.201
FDA | WAl | 25_50 | 0.6 | 25 | 0.204
FDA | WAl | 25_50 | 0.6 | 26 | 0.206
FDA | WAl | 25_50 | 0.6 | 27 | 0.209
FDA | WAl | 25_50 | 0.6 | 28 | 0.212
FDA | WAl | 25_50 | 0.6 | 29 | 0.215
FDA | WAl | 25_50 | 0.6 | 30 | 0.217
FDA | WAl | 25_50 | 0.6 | 31 | 0.219
FDA | WAl | 25_50 | 0.6 | 32 | 0.221
FDA | WAl | 25_50 | 0.6 | 33 | 0.223
FDA | WAl | 25_50 | 0.6 | 34 | 0.225
FDA | WAl | 25_50 | 0.6 | 35 | 0.227
FDA | WAl | 25_50 | 0.6 | 36 | 0.229
FDA | WAl | 25_50 | 0.6 | 37 | 0.231
FDA | WAl | 25_50 | 0.6 | 38 | 0.233
FDA | WAl | 25_50 | 0.6 | 39 | 0.234
FDA | WAl | 25_50 | 0.6 | 40 | 0.236
FDA | WAl | 25_50 | 0.6 | 41 | 0.237
FDA | WAl | 25_50 | 0.6 | 42 | 0.239
FDA | WAl | 25_50 | 0.6 | 43 | 0.24
FDA | WAl | 25_50 | 0.6 | 44 | 0.242
FDA | WAl | 25_50 | 0.6 | 45 | 0.243
FDA | WAl | 25_50 | 0.6 | 46 | 0.244
FDA | WAl | 25_50 | 0.6 | 47 | 0.245
FDA | WAl | 25_50 | 0.6 | 48 | 0.247
FDA | WAl | 25_50 | 0.6 | 49 | 0.248
FDA | WAl | 25_50 | 0.65 | 24 | 0.218
FDA | WAl | 25_50 | 0.65 | 25 | 0.22
FDA | WAl | 25_50 | 0.65 | 26 | 0.223
FDA | WAl | 25_50 | 0.65 | 27 | 0.225
FDA | WAl | 25_50 | 0.65 | 28 | 0.228
FDA | WAl | 25_50 | 0.65 | 29 | 0.231
FDA | WAl | 25_50 | 0.65 | 30 | 0.233
FDA | WAl | 25_50 | 0.65 | 31 | 0.235
FDA | WAl | 25_50 | 0.65 | 32 | 0.237
FDA | WAl | 25_50 | 0.65 | 33 | 0.239
FDA | WAl | 25_50 | 0.65 | 34 | 0.241
FDA | WAl | 25_50 | 0.65 | 35 | 0.243
FDA | WAl | 25_50 | 0.65 | 36 | 0.244
FDA | WAl | 25_50 | 0.65 | 37 | 0.246
FDA | WAl | 25_50 | 0.65 | 38 | 0.248
FDA | WAl | 25_50 | 0.65 | 39 | 0.249
FDA | WAl | 25_50 | 0.65 | 40 | 0.251
FDA | WAl | 25_50 | 0.65 | 41 | 0.252
FDA | WAl | 25_50 | 0.65 | 42 | 0.253
FDA | WAl | 25_50 | 0.65 | 43 | 0.255
FDA | WAl | 25_50 | 0.65 | 44 | 0.256
FDA | WAl | 25_50 | 0.65 | 45 | 0.257
FDA | WAl | 25_50 | 0.65 | 46 | 0.258
FDA | WAl | 25_50 | 0.65 | 47 | 0.26
FDA | WAl | 25_50 | 0.65 | 48 | 0.261
FDA | WAl | 25_50 | 0.65 | 49 | 0.262
FDA | WAl | 25_50 | 0.7 | 25 | 0.236
FDA | WAl | 25_50 | 0.7 | 26 | 0.239
FDA | WAl | 25_50 | 0.7 | 27 | 0.241
FDA | WAl | 25_50 | 0.7 | 28 | 0.244
FDA | WAl | 25_50 | 0.7 | 29 | 0.246
FDA | WAl | 25_50 | 0.7 | 30 | 0.248
FDA | WAl | 25_50 | 0.7 | 31 | 0.25
FDA | WAl | 25_50 | 0.7 | 32 | 0.252
FDA | WAl | 25_50 | 0.7 | 33 | 0.254
FDA | WAl | 25_50 | 0.7 | 34 | 0.256
FDA | WAl | 25_50 | 0.7 | 35 | 0.258
FDA | WAl | 25_50 | 0.7 | 36 | 0.259
FDA | WAl | 25_50 | 0.7 | 37 | 0.261
FDA | WAl | 25_50 | 0.7 | 38 | 0.262
FDA | WAl | 25_50 | 0.7 | 39 | 0.264
FDA | WAl | 25_50 | 0.7 | 40 | 0.265
FDA | WAl | 25_50 | 0.7 | 41 | 0.266
FDA | WAl | 25_50 | 0.7 | 42 | 0.268
FDA | WAl | 25_50 | 0.7 | 43 | 0.269
FDA | WAl | 25_50 | 0.7 | 44 | 0.27
FDA | WAl | 25_50 | 0.7 | 45 | 0.271
FDA | WAl | 25_50 | 0.7 | 46 | 0.272
FDA | WAl | 25_50 | 0.7 | 47 | 0.273
FDA | WAl | 25_50 | 0.7 | 48 | 0.274
FDA | WAl | 25_50 | 0.7 | 49 | 0.275
FDA | WAl | 25_50 | 0.75 | 26 | 0.254
FDA | WAl | 25_50 | 0.75 | 27 | 0.257
FDA | WAl | 25_50 | 0.75 | 28 | 0.259
FDA | WAl | 25_50 | 0.75 | 29 | 0.261
FDA | WAl | 25_50 | 0.75 | 30 | 0.263
FDA | WAl | 25_50 | 0.75 | 31 | 0.265
FDA | WAl | 25_50 | 0.75 | 32 | 0.267
FDA | WAl | 25_50 | 0.75 | 33 | 0.269
FDA | WAl | 25_50 | 0.75 | 34 | 0.27
FDA | WAl | 25_50 | 0.75 | 35 | 0.272
FDA | WAl | 25_50 | 0.75 | 36 | 0.273
FDA | WAl | 25_50 | 0.75 | 37 | 0.275
FDA | WAl | 25_50 | 0.75 | 38 | 0.276
FDA | WAl | 25_50 | 0.75 | 39 | 0.278
FDA | WAl | 25_50 | 0.75 | 40 | 0.279
FDA | WAl | 25_50 | 0.75 | 41 | 0.28
FDA | WAl | 25_50 | 0.75 | 42 | 0.281
FDA | WAl | 25_50 | 0.75 | 43 | 0.283
FDA | WAl | 25_50 | 0.75 | 44 | 0.284
FDA | WAl | 25_50 | 0.75 | 45 | 0.285
FDA | WAl | 25_50 | 0.75 | 46 | 0.286
FDA | WAl | 25_50 | 0.75 | 47 | 0.287
FDA | WAl | 25_50 | 0.75 | 48 | 0.288
FDA | WAl | 25_50 | 0.75 | 49 | 0.289
FDA | WAl | 25_50 | 0.8 | 27 | 0.271
FDA | WAl | 25_50 | 0.8 | 28 | 0.274
FDA | WAl | 25_50 | 0.8 | 29 | 0.276
FDA | WAl | 25_50 | 0.8 | 30 | 0.277
FDA | WAl | 25_50 | 0.8 | 31 | 0.279
FDA | WAl | 25_50 | 0.8 | 32 | 0.281
FDA | WAl | 25_50 | 0.8 | 33 | 0.283
FDA | WAl | 25_50 | 0.8 | 34 | 0.284
FDA | WAl | 25_50 | 0.8 | 35 | 0.286
FDA | WAl | 25_50 | 0.8 | 36 | 0.287
FDA | WAl | 25_50 | 0.8 | 37 | 0.289
FDA | WAl | 25_50 | 0.8 | 38 | 0.29
FDA | WAl | 25_50 | 0.8 | 39 | 0.291
FDA | WAl | 25_50 | 0.8 | 40 | 0.292
FDA | WAl | 25_50 | 0.8 | 41 | 0.294
FDA | WAl | 25_50 | 0.8 | 42 | 0.295
FDA | WAl | 25_50 | 0.8 | 43 | 0.296
FDA | WAl | 25_50 | 0.8 | 44 | 0.297
FDA | WAl | 25_50 | 0.8 | 45 | 0.298
FDA | WAl | 25_50 | 0.8 | 46 | 0.299
FDA | WAl | 25_50 | 0.8 | 47 | 0.3
FDA | WAl | 25_50 | 0.8 | 48 | 0.301
FDA | WAl | 25_50 | 0.8 | 49 | 0.302
FDA | WAl | 25_50 | 0.85 | 28 | 0.288
FDA | WAl | 25_50 | 0.85 | 29 | 0.29
FDA | WAl | 25_50 | 0.85 | 30 | 0.291
FDA | WAl | 25_50 | 0.85 | 31 | 0.293
FDA | WAl | 25_50 | 0.85 | 32 | 0.295
FDA | WAl | 25_50 | 0.85 | 33 | 0.296
FDA | WAl | 25_50 | 0.85 | 34 | 0.298
FDA | WAl | 25_50 | 0.85 | 35 | 0.299
FDA | WAl | 25_50 | 0.85 | 36 | 0.301
FDA | WAl | 25_50 | 0.85 | 37 | 0.302
FDA | WAl | 25_50 | 0.85 | 38 | 0.303
FDA | WAl | 25_50 | 0.85 | 39 | 0.304
FDA | WAl | 25_50 | 0.85 | 40 | 0.306
FDA | WAl | 25_50 | 0.85 | 41 | 0.307
FDA | WAl | 25_50 | 0.85 | 42 | 0.308
FDA | WAl | 25_50 | 0.85 | 43 | 0.309
FDA | WAl | 25_50 | 0.85 | 44 | 0.31
FDA | WAl | 25_50 | 0.85 | 45 | 0.311
FDA | WAl | 25_50 | 0.85 | 46 | 0.311
FDA | WAl | 25_50 | 0.85 | 47 | 0.312
FDA | WAl | 25_50 | 0.85 | 48 | 0.313
FDA | WAl | 25_50 | 0.85 | 49 | 0.314
FDA | WAl | 25_50 | 0.9 | 28 | 0.302
FDA | WAl | 25_50 | 0.9 | 29 | 0.303
FDA | WAl | 25_50 | 0.9 | 30 | 0.305
FDA | WAl | 25_50 | 0.9 | 31 | 0.306
FDA | WAl | 25_50 | 0.9 | 32 | 0.308
FDA | WAl | 25_50 | 0.9 | 33 | 0.31
FDA | WAl | 25_50 | 0.9 | 34 | 0.311
FDA | WAl | 25_50 | 0.9 | 35 | 0.312
FDA | WAl | 25_50 | 0.9 | 36 | 0.314
FDA | WAl | 25_50 | 0.9 | 37 | 0.315
FDA | WAl | 25_50 | 0.9 | 38 | 0.316
FDA | WAl | 25_50 | 0.9 | 39 | 0.317
FDA | WAl | 25_50 | 0.9 | 40 | 0.318
FDA | WAl | 25_50 | 0.9 | 41 | 0.319
FDA | WAl | 25_50 | 0.9 | 42 | 0.32
FDA | WAl | 25_50 | 0.9 | 43 | 0.321
FDA | WAl | 25_50 | 0.9 | 44 | 0.322
FDA | WAl | 25_50 | 0.9 | 45 | 0.323
FDA | WAl | 25_50 | 0.9 | 46 | 0.324
FDA | WAl | 25_50 | 0.9 | 47 | 0.325
FDA | WAl | 25_50 | 0.9 | 48 | 0.325
FDA | WAl | 25_50 | 0.9 | 49 | 0.326
FDA | WAl | 25_50 | 0.95 | 29 | 0.317
FDA | WAl | 25_50 | 0.95 | 30 | 0.318
FDA | WAl | 25_50 | 0.95 | 31 | 0.32
FDA | WAl | 25_50 | 0.95 | 32 | 0.321
FDA | WAl | 25_50 | 0.95 | 33 | 0.322
FDA | WAl | 25_50 | 0.95 | 34 | 0.324
FDA | WAl | 25_50 | 0.95 | 35 | 0.325
FDA | WAl | 25_50 | 0.95 | 36 | 0.326
FDA | WAl | 25_50 | 0.95 | 37 | 0.327
FDA | WAl | 25_50 | 0.95 | 38 | 0.328
FDA | WAl | 25_50 | 0.95 | 39 | 0.329
FDA | WAl | 25_50 | 0.95 | 40 | 0.33
FDA | WAl | 25_50 | 0.95 | 41 | 0.331
FDA | WAl | 25_50 | 0.95 | 42 | 0.332
FDA | WAl | 25_50 | 0.95 | 43 | 0.333
FDA | WAl | 25_50 | 0.95 | 44 | 0.334
FDA | WAl | 25_50 | 0.95 | 45 | 0.335
FDA | WAl | 25_50 | 0.95 | 46 | 0.335
FDA | WAl | 25_50 | 0.95 | 47 | 0.336
FDA | WAl | 25_50 | 0.95 | 48 | 0.337
FDA | WAl | 25_50 | 0.95 | 49 | 0.338
FDA | WAl | 25_50 | 1 | 30 | 0.331
FDA | WAl | 25_50 | 1 | 31 | 0.332
FDA | WAl | 25_50 | 1 | 32 | 0.333
FDA | WAl | 25_50 | 1 | 33 | 0.335
FDA | WAl | 25_50 | 1 | 34 | 0.336
FDA | WAl | 25_50 | 1 | 35 | 0.337
FDA | WAl | 25_50 | 1 | 36 | 0.338
FDA | WAl | 25_50 | 1 | 37 | 0.339
FDA | WAl | 25_50 | 1 | 38 | 0.34
FDA | WAl | 25_50 | 1 | 39 | 0.341
FDA | WAl | 25_50 | 1 | 40 | 0.342
FDA | WAl | 25_50 | 1 | 41 | 0.343
FDA | WAl | 25_50 | 1 | 42 | 0.344
FDA | WAl | 25_50 | 1 | 43 | 0.345
FDA | WAl | 25_50 | 1 | 44 | 0.346
FDA | WAl | 25_50 | 1 | 45 | 0.346
FDA | WAl | 25_50 | 1 | 46 | 0.347
FDA | WAl | 25_50 | 1 | 47 | 0.348
FDA | WAl | 25_50 | 1 | 48 | 0.348
FDA | WAl | 25_50 | 1 | 49 | 0.349
FDA | WAl | 25_50 | 1.05 | 31 | 0.344
FDA | WAl | 25_50 | 1.05 | 32 | 0.345
FDA | WAl | 25_50 | 1.05 | 33 | 0.347
FDA | WAl | 25_50 | 1.05 | 34 | 0.348
FDA | WAl | 25_50 | 1.05 | 35 | 0.349
FDA | WAl | 25_50 | 1.05 | 36 | 0.35
FDA | WAl | 25_50 | 1.05 | 37 | 0.351
FDA | WAl | 25_50 | 1.05 | 38 | 0.352
FDA | WAl | 25_50 | 1.05 | 39 | 0.353
FDA | WAl | 25_50 | 1.05 | 40 | 0.354
FDA | WAl | 25_50 | 1.05 | 41 | 0.354
FDA | WAl | 25_50 | 1.05 | 42 | 0.355
FDA | WAl | 25_50 | 1.05 | 43 | 0.356
FDA | WAl | 25_50 | 1.05 | 44 | 0.357
FDA | WAl | 25_50 | 1.05 | 45 | 0.357
FDA | WAl | 25_50 | 1.05 | 46 | 0.358
FDA | WAl | 25_50 | 1.05 | 47 | 0.359
FDA | WAl | 25_50 | 1.05 | 48 | 0.359
FDA | WAl | 25_50 | 1.05 | 49 | 0.36
FDA | WAl | 25_50 | 1.1 | 32 | 0.357
FDA | WAl | 25_50 | 1.1 | 33 | 0.358
FDA | WAl | 25_50 | 1.1 | 34 | 0.359
FDA | WAl | 25_50 | 1.1 | 35 | 0.36
FDA | WAl | 25_50 | 1.1 | 36 | 0.361
FDA | WAl | 25_50 | 1.1 | 37 | 0.362
FDA | WAl | 25_50 | 1.1 | 38 | 0.363
FDA | WAl | 25_50 | 1.1 | 39 | 0.364
FDA | WAl | 25_50 | 1.1 | 40 | 0.365
FDA | WAl | 25_50 | 1.1 | 41 | 0.365
FDA | WAl | 25_50 | 1.1 | 42 | 0.366
FDA | WAl | 25_50 | 1.1 | 43 | 0.367
FDA | WAl | 25_50 | 1.1 | 44 | 0.368
FDA | WAl | 25_50 | 1.1 | 45 | 0.368
FDA | WAl | 25_50 | 1.1 | 46 | 0.369
FDA | WAl | 25_50 | 1.1 | 47 | 0.369
FDA | WAl | 25_50 | 1.1 | 48 | 0.37
FDA | WAl | 25_50 | 1.1 | 49 | 0.371
FDA | WAl | 25_50 | 1.15 | 33 | 0.369
FDA | WAl | 25_50 | 1.15 | 34 | 0.37
FDA | WAl | 25_50 | 1.15 | 35 | 0.371
FDA | WAl | 25_50 | 1.15 | 36 | 0.372
FDA | WAl | 25_50 | 1.15 | 37 | 0.373
FDA | WAl | 25_50 | 1.15 | 38 | 0.374
FDA | WAl | 25_50 | 1.15 | 39 | 0.375
FDA | WAl | 25_50 | 1.15 | 40 | 0.375
FDA | WAl | 25_50 | 1.15 | 41 | 0.376
FDA | WAl | 25_50 | 1.15 | 42 | 0.377
FDA | WAl | 25_50 | 1.15 | 43 | 0.377
FDA | WAl | 25_50 | 1.15 | 44 | 0.378
FDA | WAl | 25_50 | 1.15 | 45 | 0.379
FDA | WAl | 25_50 | 1.15 | 46 | 0.379
FDA | WAl | 25_50 | 1.15 | 47 | 0.38
FDA | WAl | 25_50 | 1.15 | 48 | 0.38
FDA | WAl | 25_50 | 1.15 | 49 | 0.381
FDA | WAl | 25_50 | 1.2 | 34 | 0.381
FDA | WAl | 25_50 | 1.2 | 35 | 0.382
FDA | WAl | 25_50 | 1.2 | 36 | 0.382
FDA | WAl | 25_50 | 1.2 | 37 | 0.383
FDA | WAl | 25_50 | 1.2 | 38 | 0.384
FDA | WAl | 25_50 | 1.2 | 39 | 0.385
FDA | WAl | 25_50 | 1.2 | 40 | 0.386
FDA | WAl | 25_50 | 1.2 | 41 | 0.386
FDA | WAl | 25_50 | 1.2 | 42 | 0.387
FDA | WAl | 25_50 | 1.2 | 43 | 0.387
FDA | WAl | 25_50 | 1.2 | 44 | 0.388
FDA | WAl | 25_50 | 1.2 | 45 | 0.389
FDA | WAl | 25_50 | 1.2 | 46 | 0.389
FDA | WAl | 25_50 | 1.2 | 47 | 0.39
FDA | WAl | 25_50 | 1.2 | 48 | 0.39
FDA | WAl | 25_50 | 1.2 | 49 | 0.391
FDA | WAl | 25_50 | 1.25 | 36 | 0.393
FDA | WAl | 25_50 | 1.25 | 37 | 0.393
FDA | WAl | 25_50 | 1.25 | 38 | 0.394
FDA | WAl | 25_50 | 1.25 | 39 | 0.395
FDA | WAl | 25_50 | 1.25 | 40 | 0.395
FDA | WAl | 25_50 | 1.25 | 41 | 0.396
FDA | WAl | 25_50 | 1.25 | 42 | 0.397
FDA | WAl | 25_50 | 1.25 | 43 | 0.397
FDA | WAl | 25_50 | 1.25 | 44 | 0.398
FDA | WAl | 25_50 | 1.25 | 45 | 0.398
FDA | WAl | 25_50 | 1.25 | 46 | 0.399
FDA | WAl | 25_50 | 1.25 | 47 | 0.399
FDA | WAl | 25_50 | 1.25 | 48 | 0.4
FDA | WAl | 25_50 | 1.25 | 49 | 0.4
FDA | WAl | 25_50 | 1.3 | 41 | 0.406
FDA | WAl | 25_50 | 1.3 | 42 | 0.406
FDA | WAl | 25_50 | 1.3 | 43 | 0.407
FDA | WAl | 25_50 | 1.3 | 44 | 0.407
FDA | WAl | 25_50 | 1.3 | 45 | 0.408
FDA | WAl | 25_50 | 1.3 | 46 | 0.408
FDA | WAl | 25_50 | 1.3 | 47 | 0.409
FDA | WAl | 25_50 | 1.3 | 48 | 0.409
FDA | WAl | 25_50 | 1.3 | 49 | 0.41
FDA | WAl | 25_50 | 1.35 | 46 | 0.417
FDA | WAl | 25_50 | 1.35 | 47 | 0.418
FDA | WAl | 25_50 | 1.35 | 48 | 0.418
FDA | WAl | 25_50 | 1.35 | 49 | 0.418
FDA | WRh | 25_50 | 0.2 | 20 | 0.046
FDA | WRh | 25_50 | 0.2 | 21 | 0.049
FDA | WRh | 25_50 | 0.225 | 20 | 0.055
FDA | WRh | 25_50 | 0.225 | 21 | 0.058
FDA | WRh | 25_50 | 0.225 | 22 | 0.062
FDA | WRh | 25_50 | 0.225 | 23 | 0.065
FDA | WRh | 25_50 | 0.225 | 24 | 0.068
FDA | WRh | 25_50 | 0.25 | 20 | 0.063
FDA | WRh | 25_50 | 0.25 | 21 | 0.067
FDA | WRh | 25_50 | 0.25 | 22 | 0.071
FDA | WRh | 25_50 | 0.25 | 23 | 0.074
FDA | WRh | 25_50 | 0.25 | 24 | 0.078
FDA | WRh | 25_50 | 0.25 | 25 | 0.081
FDA | WRh | 25_50 | 0.25 | 26 | 0.083
FDA | WRh | 25_50 | 0.275 | 20 | 0.071
FDA | WRh | 25_50 | 0.275 | 21 | 0.076
FDA | WRh | 25_50 | 0.275 | 22 | 0.079
FDA | WRh | 25_50 | 0.275 | 23 | 0.083
FDA | WRh | 25_50 | 0.275 | 24 | 0.087
FDA | WRh | 25_50 | 0.275 | 25 | 0.09
FDA | WRh | 25_50 | 0.275 | 26 | 0.093
FDA | WRh | 25_50 | 0.275 | 27 | 0.095
FDA | WRh | 25_50 | 0.275 | 28 | 0.097
FDA | WRh | 25_50 | 0.275 | 29 | 0.099
FDA | WRh | 25_50 | 0.275 | 30 | 0.1
FDA | WRh | 25_50 | 0.275 | 31 | 0.102
FDA | WRh | 25_50 | 0.275 | 32 | 0.104
FDA | WRh | 25_50 | 0.3 | 20 | 0.08
FDA | WRh | 25_50 | 0.3 | 21 | 0.084
FDA | WRh | 25_50 | 0.3 | 22 | 0.088
FDA | WRh | 25_50 | 0.3 | 23 | 0.092
FDA | WRh | 25_50 | 0.3 | 24 | 0.096
FDA | WRh | 25_50 | 0.3 | 25 | 0.099
FDA | WRh | 25_50 | 0.3 | 26 | 0.102
FDA | WRh | 25_50 | 0.3 | 27 | 0.104
FDA | WRh | 25_50 | 0.3 | 28 | 0.106
FDA | WRh | 25_50 | 0.3 | 29 | 0.107
FDA | WRh | 25_50 | 0.3 | 30 | 0.109
FDA | WRh | 25_50 | 0.3 | 31 | 0.111
FDA | WRh | 25_50 | 0.3 | 32 | 0.112
FDA | WRh | 25_50 | 0.3 | 33 | 0.114
FDA | WRh | 25_50 | 0.3 | 34 | 0.116
FDA | WRh | 25_50 | 0.3 | 35 | 0.118
FDA | WRh | 25_50 | 0.3 | 36 | 0.12
FDA | WRh | 25_50 | 0.325 | 20 | 0.088
FDA | WRh | 25_50 | 0.325 | 21 | 0.093
FDA | WRh | 25_50 | 0.325 | 22 | 0.097
FDA | WRh | 25_50 | 0.325 | 23 | 0.101
FDA | WRh | 25_50 | 0.325 | 24 | 0.105
FDA | WRh | 25_50 | 0.325 | 25 | 0.108
FDA | WRh | 25_50 | 0.325 | 26 | 0.111
FDA | WRh | 25_50 | 0.325 | 27 | 0.113
FDA | WRh | 25_50 | 0.325 | 28 | 0.114
FDA | WRh | 25_50 | 0.325 | 29 | 0.116
FDA | WRh | 25_50 | 0.325 | 30 | 0.117
FDA | WRh | 25_50 | 0.325 | 31 | 0.119
FDA | WRh | 25_50 | 0.325 | 32 | 0.121
FDA | WRh | 25_50 | 0.325 | 33 | 0.123
FDA | WRh | 25_50 | 0.325 | 34 | 0.124
FDA | WRh | 25_50 | 0.325 | 35 | 0.126
FDA | WRh | 25_50 | 0.325 | 36 | 0.128
FDA | WRh | 25_50 | 0.325 | 37 | 0.129
FDA | WRh | 25_50 | 0.325 | 38 | 0.131
FDA | WRh | 25_50 | 0.325 | 39 | 0.133
FDA | WRh | 25_50 | 0.325 | 40 | 0.135
FDA | WRh | 25_50 | 0.325 | 41 | 0.136
FDA | WRh | 25_50 | 0.325 | 42 | 0.138
FDA | WRh | 25_50 | 0.35 | 20 | 0.097
FDA | WRh | 25_50 | 0.35 | 21 | 0.102
FDA | WRh | 25_50 | 0.35 | 22 | 0.106
FDA | WRh | 25_50 | 0.35 | 23 | 0.11
FDA | WRh | 25_50 | 0.35 | 24 | 0.114
FDA | WRh | 25_50 | 0.35 | 25 | 0.117
FDA | WRh | 25_50 | 0.35 | 26 | 0.119
FDA | WRh | 25_50 | 0.35 | 27 | 0.121
FDA | WRh | 25_50 | 0.35 | 28 | 0.123
FDA | WRh | 25_50 | 0.35 | 29 | 0.124
FDA | WRh | 25_50 | 0.35 | 30 | 0.126
FDA | WRh | 25_50 | 0.35 | 31 | 0.127
FDA | WRh | 25_50 | 0.35 | 32 | 0.129
FDA | WRh | 25_50 | 0.35 | 33 | 0.131
FDA | WRh | 25_50 | 0.35 | 34 | 0.132
FDA | WRh | 25_50 | 0.35 | 35 | 0.134
FDA | WRh | 25_50 | 0.35 | 36 | 0.136
FDA | WRh | 25_50 | 0.35 | 37 | 0.137
FDA | WRh | 25_50 | 0.35 | 38 | 0.139
FDA | WRh | 25_50 | 0.35 | 39 | 0.141
FDA | WRh | 25_50 | 0.35 | 40 | 0.143
FDA | WRh | 25_50 | 0.35 | 41 | 0.144
FDA | WRh | 25_50 | 0.35 | 42 | 0.146
FDA | WRh | 25_50 | 0.35 | 43 | 0.148
FDA | WRh | 25_50 | 0.35 | 44 | 0.149
FDA | WRh | 25_50 | 0.35 | 45 | 0.151
FDA | WRh | 25_50 | 0.35 | 46 | 0.153
FDA | WRh | 25_50 | 0.35 | 47 | 0.154
FDA | WRh | 25_50 | 0.375 | 20 | 0.106
FDA | WRh | 25_50 | 0.375 | 21 | 0.11
FDA | WRh | 25_50 | 0.375 | 22 | 0.115
FDA | WRh | 25_50 | 0.375 | 23 | 0.119
FDA | WRh | 25_50 | 0.375 | 24 | 0.123
FDA | WRh | 25_50 | 0.375 | 25 | 0.126
FDA | WRh | 25_50 | 0.375 | 26 | 0.128
FDA | WRh | 25_50 | 0.375 | 27 | 0.13
FDA | WRh | 25_50 | 0.375 | 28 | 0.131
FDA | WRh | 25_50 | 0.375 | 29 | 0.133
FDA | WRh | 25_50 | 0.375 | 30 | 0.134
FDA | WRh | 25_50 | 0.375 | 31 | 0.135
FDA | WRh | 25_50 | 0.375 | 32 | 0.137
FDA | WRh | 25_50 | 0.375 | 33 | 0.139
FDA | WRh | 25_50 | 0.375 | 34 | 0.14
FDA | WRh | 25_50 | 0.375 | 35 | 0.142
FDA | WRh | 25_50 | 0.375 | 36 | 0.143
FDA | WRh | 25_50 | 0.375 | 37 | 0.145
FDA | WRh | 25_50 | 0.375 | 38 | 0.147
FDA | WRh | 25_50 | 0.375 | 39 | 0.148
FDA | WRh | 25_50 | 0.375 | 40 | 0.15
FDA | WRh | 25_50 | 0.375 | 41 | 0.152
FDA | WRh | 25_50 | 0.375 | 42 | 0.154
FDA | WRh | 25_50 | 0.375 | 43 | 0.155
FDA | WRh | 25_50 | 0.375 | 44 | 0.157
FDA | WRh | 25_50 | 0.375 | 45 | 0.159
FDA | WRh | 25_50 | 0.375 | 46 | 0.16
FDA | WRh | 25_50 | 0.375 | 47 | 0.162
FDA | WRh | 25_50 | 0.375 | 48 | 0.164
FDA | WRh | 25_50 | 0.375 | 49 | 0.165
FDA | WRh | 25_50 | 0.4 | 20 | 0.115
FDA | WRh | 25_50 | 0.4 | 21 | 0.119
FDA | WRh | 25_50 | 0.4 | 22 | 0.124
FDA | WRh | 25_50 | 0.4 | 23 | 0.128
FDA | WRh | 25_50 | 0.4 | 24 | 0.132
FDA | WRh | 25_50 | 0.4 | 25 | 0.135
FDA | WRh | 25_50 | 0.4 | 26 | 0.137
FDA | WRh | 25_50 | 0.4 | 27 | 0.138
FDA | WRh | 25_50 | 0.4 | 28 | 0.14
FDA | WRh | 25_50 | 0.4 | 29 | 0.141
FDA | WRh | 25_50 | 0.4 | 30 | 0.142
FDA | WRh | 25_50 | 0.4 | 31 | 0.144
FDA | WRh | 25_50 | 0.4 | 32 | 0.145
FDA | WRh | 25_50 | 0.4 | 33 | 0.146
FDA | WRh | 25_50 | 0.4 | 34 | 0.148
FDA | WRh | 25_50 | 0.4 | 35 | 0.149
FDA | WRh | 25_50 | 0.4 | 36 | 0.151
FDA | WRh | 25_50 | 0.4 | 37 | 0.152
FDA | WRh | 25_50 | 0.4 | 38 | 0.154
FDA | WRh | 25_50 | 0.4 | 39 | 0.156
FDA | WRh | 25_50 | 0.4 | 40 | 0.158
FDA | WRh | 25_50 | 0.4 | 41 | 0.159
FDA | WRh | 25_50 | 0.4 | 42 | 0.161
FDA | WRh | 25_50 | 0.4 | 43 | 0.163
FDA | WRh | 25_50 | 0.4 | 44 | 0.164
FDA | WRh | 25_50 | 0.4 | 45 | 0.166
FDA | WRh | 25_50 | 0.4 | 46 | 0.168
FDA | WRh | 25_50 | 0.4 | 47 | 0.169
FDA | WRh | 25_50 | 0.4 | 48 | 0.171
FDA | WRh | 25_50 | 0.4 | 49 | 0.173
FDA | WRh | 25_50 | 0.425 | 20 | 0.124
FDA | WRh | 25_50 | 0.425 | 21 | 0.128
FDA | WRh | 25_50 | 0.425 | 22 | 0.133
FDA | WRh | 25_50 | 0.425 | 23 | 0.137
FDA | WRh | 25_50 | 0.425 | 24 | 0.14
FDA | WRh | 25_50 | 0.425 | 25 | 0.143
FDA | WRh | 25_50 | 0.425 | 26 | 0.145
FDA | WRh | 25_50 | 0.425 | 27 | 0.147
FDA | WRh | 25_50 | 0.425 | 28 | 0.148
FDA | WRh | 25_50 | 0.425 | 29 | 0.149
FDA | WRh | 25_50 | 0.425 | 30 | 0.15
FDA | WRh | 25_50 | 0.425 | 31 | 0.151
FDA | WRh | 25_50 | 0.425 | 32 | 0.153
FDA | WRh | 25_50 | 0.425 | 33 | 0.154
FDA | WRh | 25_50 | 0.425 | 34 | 0.155
FDA | WRh | 25_50 | 0.425 | 35 | 0.157
FDA | WRh | 25_50 | 0.425 | 36 | 0.158
FDA | WRh | 25_50 | 0.425 | 37 | 0.16
FDA | WRh | 25_50 | 0.425 | 38 | 0.162
FDA | WRh | 25_50 | 0.425 | 39 | 0.163
FDA | WRh | 25_50 | 0.425 | 40 | 0.165
FDA | WRh | 25_50 | 0.425 | 41 | 0.167
FDA | WRh | 25_50 | 0.425 | 42 | 0.168
FDA | WRh | 25_50 | 0.425 | 43 | 0.17
FDA | WRh | 25_50 | 0.425 | 44 | 0.172
FDA | WRh | 25_50 | 0.425 | 45 | 0.173
FDA | WRh | 25_50 | 0.425 | 46 | 0.175
FDA | WRh | 25_50 | 0.425 | 47 | 0.176
FDA | WRh | 25_50 | 0.425 | 48 | 0.178
FDA | WRh | 25_50 | 0.425 | 49 | 0.18
FDA | WRh | 25_50 | 0.45 | 20 | 0.133
FDA | WRh | 25_50 | 0.45 | 21 | 0.138
FDA | WRh | 25_50 | 0.45 | 22 | 0.142
FDA | WRh | 25_50 | 0.45 | 23 | 0.146
FDA | WRh | 25_50 | 0.45 | 24 | 0.149
FDA | WRh | 25_50 | 0.45 | 25 | 0.152
FDA | WRh | 25_50 | 0.45 | 26 | 0.154
FDA | WRh | 25_50 | 0.45 | 27 | 0.155
FDA | WRh | 25_50 | 0.45 | 28 | 0.157
FDA | WRh | 25_50 | 0.45 | 29 | 0.158
FDA | WRh | 25_50 | 0.45 | 30 | 0.158
FDA | WRh | 25_50 | 0.45 | 31 | 0.16
FDA | WRh | 25_50 | 0.45 | 32 | 0.161
FDA | WRh | 25_50 | 0.45 | 33 | 0.162
FDA | WRh | 25_50 | 0.45 | 34 | 0.163
FDA | WRh | 25_50 | 0.45 | 35 | 0.164
FDA | WRh | 25_50 | 0.45 | 36 | 0.166
FDA | WRh | 25_50 | 0.45 | 37 | 0.167
FDA | WRh | 25_50 | 0.45 | 38 | 0.169
FDA | WRh | 25_50 | 0.45 | 39 | 0.17
FDA | WRh | 25_50 | 0.45 | 40 | 0.172
FDA | WRh | 25_50 | 0.45 | 41 | 0.174
FDA | WRh | 25_50 | 0.45 | 42 | 0.175
FDA | WRh | 25_50 | 0.45 | 43 | 0.177
FDA | WRh | 25_50 | 0.45 | 44 | 0.179
FDA | WRh | 25_50 | 0.45 | 45 | 0.18
FDA | WRh | 25_50 | 0.45 | 46 | 0.182
FDA | WRh | 25_50 | 0.45 | 47 | 0.184
FDA | WRh | 25_50 | 0.45 | 48 | 0.185
FDA | WRh | 25_50 | 0.45 | 49 | 0.187
FDA | WRh | 25_50 | 0.475 | 20 | 0.143
FDA | WRh | 25_50 | 0.475 | 21 | 0.147
FDA | WRh | 25_50 | 0.475 | 22 | 0.151
FDA | WRh | 25_50 | 0.475 | 23 | 0.154
FDA | WRh | 25_50 | 0.475 | 24 | 0.158
FDA | WRh | 25_50 | 0.475 | 25 | 0.161
FDA | WRh | 25_50 | 0.475 | 26 | 0.163
FDA | WRh | 25_50 | 0.475 | 27 | 0.164
FDA | WRh | 25_50 | 0.475 | 28 | 0.165
FDA | WRh | 25_50 | 0.475 | 29 | 0.166
FDA | WRh | 25_50 | 0.475 | 30 | 0.167
FDA | WRh | 25_50 | 0.475 | 31 | 0.167
FDA | WRh | 25_50 | 0.475 | 32 | 0.168
FDA | WRh | 25_50 | 0.475 | 33 | 0.17
FDA | WRh | 25_50 | 0.475 | 34 | 0.171
FDA | WRh | 25_50 | 0.475 | 35 | 0.172
FDA | WRh | 25_50 | 0.475 | 36 | 0.173
FDA | WRh | 25_50 | 0.475 | 37 | 0.175
FDA | WRh | 25_50 | 0.475 | 38 | 0.176
FDA | WRh | 25_50 | 0.475 | 39 | 0.178
FDA | WRh | 25_50 | 0.475 | 40 | 0.179
FDA | WRh | 25_50 | 0.475 | 41 | 0.181
FDA | WRh | 25_50 | 0.475 | 42 | 0.182
FDA | WRh | 25_50 | 0.475 | 43 | 0.184
FDA | WRh | 25_50 | 0.475 | 44 | 0.186
FDA | WRh | 25_50 | 0.475 | 45 | 0.187
FDA | WRh | 25_50 | 0.475 | 46 | 0.189
FDA | WRh | 25_50 | 0.475 | 47 | 0.191
FDA | WRh | 25_50 | 0.475 | 48 | 0.192
FDA | WRh | 25_50 | 0.475 | 49 | 0.194
FDA | WRh | 25_50 | 0.5 | 20 | 0.152
FDA | WRh | 25_50 | 0.5 | 21 | 0.156
FDA | WRh | 25_50 | 0.5 | 22 | 0.16
FDA | WRh | 25_50 | 0.5 | 23 | 0.163
FDA | WRh | 25_50 | 0.5 | 24 | 0.167
FDA | WRh | 25_50 | 0.5 | 25 | 0.17
FDA | WRh | 25_50 | 0.5 | 26 | 0.171
FDA | WRh | 25_50 | 0.5 | 27 | 0.172
FDA | WRh | 25_50 | 0.5 | 28 | 0.173
FDA | WRh | 25_50 | 0.5 | 29 | 0.174
FDA | WRh | 25_50 | 0.5 | 30 | 0.175
FDA | WRh | 25_50 | 0.5 | 31 | 0.175
FDA | WRh | 25_50 | 0.5 | 32 | 0.176
FDA | WRh | 25_50 | 0.5 | 33 | 0.177
FDA | WRh | 25_50 | 0.5 | 34 | 0.178
FDA | WRh | 25_50 | 0.5 | 35 | 0.179
FDA | WRh | 25_50 | 0.5 | 36 | 0.181
FDA | WRh | 25_50 | 0.5 | 37 | 0.182
FDA | WRh | 25_50 | 0.5 | 38 | 0.183
FDA | WRh | 25_50 | 0.5 | 39 | 0.185
FDA | WRh | 25_50 | 0.5 | 40 | 0.186
FDA | WRh | 25_50 | 0.5 | 41 | 0.188
FDA | WRh | 25_50 | 0.5 | 42 | 0.189
FDA | WRh | 25_50 | 0.5 | 43 | 0.191
FDA | WRh | 25_50 | 0.5 | 44 | 0.193
FDA | WRh | 25_50 | 0.5 | 45 | 0.194
FDA | WRh | 25_50 | 0.5 | 46 | 0.196
FDA | WRh | 25_50 | 0.5 | 47 | 0.198
FDA | WRh | 25_50 | 0.5 | 48 | 0.199
FDA | WRh | 25_50 | 0.5 | 49 | 0.201
FDA | WRh | 25_50 | 0.525 | 21 | 0.165
FDA | WRh | 25_50 | 0.525 | 22 | 0.169
FDA | WRh | 25_50 | 0.525 | 23 | 0.172
FDA | WRh | 25_50 | 0.525 | 24 | 0.176
FDA | WRh | 25_50 | 0.525 | 25 | 0.178
FDA | WRh | 25_50 | 0.525 | 26 | 0.18
FDA | WRh | 25_50 | 0.525 | 27 | 0.181
FDA | WRh | 25_50 | 0.525 | 28 | 0.182
FDA | WRh | 25_50 | 0.525 | 29 | 0.182
FDA | WRh | 25_50 | 0.525 | 30 | 0.183
FDA | WRh | 25_50 | 0.525 | 31 | 0.183
FDA | WRh | 25_50 | 0.525 | 32 | 0.184
FDA | WRh | 25_50 | 0.525 | 33 | 0.185
FDA | WRh | 25_50 | 0.525 | 34 | 0.186
FDA | WRh | 25_50 | 0.525 | 35 | 0.187
FDA | WRh | 25_50 | 0.525 | 36 | 0.188
FDA | WRh | 25_50 | 0.525 | 37 | 0.189
FDA | WRh | 25_50 | 0.525 | 38 | 0.191
FDA | WRh | 25_50 | 0.525 | 39 | 0.192
FDA | WRh | 25_50 | 0.525 | 40 | 0.193
FDA | WRh | 25_50 | 0.525 | 41 | 0.195
FDA | WRh | 25_50 | 0.525 | 42 | 0.196
FDA | WRh | 25_50 | 0.525 | 43 | 0.198
FDA | WRh | 25_50 | 0.525 | 44 | 0.2
FDA | WRh | 25_50 | 0.525 | 45 | 0.201
FDA | WRh | 25_50 | 0.525 | 46 | 0.203
FDA | WRh | 25_50 | 0.525 | 47 | 0.204
FDA | WRh | 25_50 | 0.525 | 48 | 0.206
FDA | WRh | 25_50 | 0.525 | 49 | 0.208
FDA | WRh | 25_50 | 0.55 | 21 | 0.174
FDA | WRh | 25_50 | 0.55 | 22 | 0.177
FDA | WRh | 25_50 | 0.55 | 23 | 0.181
FDA | WRh | 25_50 | 0.55 | 24 | 0.184
FDA | WRh | 25_50 | 0.55 | 25 | 0.187
FDA | WRh | 25_50 | 0.55 | 26 | 0.188
FDA | WRh | 25_50 | 0.55 | 27 | 0.189
FDA | WRh | 25_50 | 0.55 | 28 | 0.19
FDA | WRh | 25_50 | 0.55 | 29 | 0.19
FDA | WRh | 25_50 | 0.55 | 30 | 0.191
FDA | WRh | 25_50 | 0.55 | 31 | 0.191
FDA | WRh | 25_50 | 0.55 | 32 | 0.192
FDA | WRh | 25_50 | 0.55 | 33 | 0.193
FDA | WRh | 25_50 | 0.55 | 34 | 0.194
FDA | WRh | 25_50 | 0.55 | 35 | 0.194
FDA | WRh | 25_50 | 0.55 | 36 | 0.195
FDA | WRh | 25_50 | 0.55 | 37 | 0.197
FDA | WRh | 25_50 | 0.55 | 38 | 0.198
FDA | WRh | 25_50 | 0.55 | 39 | 0.199
FDA | WRh | 25_50 | 0.55 | 40 | 0.2
FDA | WRh | 25_50 | 0.55 | 41 | 0.202
FDA | WRh | 25_50 | 0.55 | 42 | 0.203
FDA | WRh | 25_50 | 0.55 | 43 | 0.205
FDA | WRh | 25_50 | 0.55 | 44 | 0.207
FDA | WRh | 25_50 | 0.55 | 45 | 0.208
FDA | WRh | 25_50 | 0.55 | 46 | 0.21
FDA | WRh | 25_50 | 0.55 | 47 | 0.211
FDA | WRh | 25_50 | 0.55 | 48 | 0.213
FDA | WRh | 25_50 | 0.55 | 49 | 0.214
FDA | WRh | 25_50 | 0.575 | 22 | 0.186
FDA | WRh | 25_50 | 0.575 | 23 | 0.19
FDA | WRh | 25_50 | 0.575 | 24 | 0.193
FDA | WRh | 25_50 | 0.575 | 25 | 0.195
FDA | WRh | 25_50 | 0.575 | 26 | 0.196
FDA | WRh | 25_50 | 0.575 | 27 | 0.197
FDA | WRh | 25_50 | 0.575 | 28 | 0.198
FDA | WRh | 25_50 | 0.575 | 29 | 0.198
FDA | WRh | 25_50 | 0.575 | 30 | 0.199
FDA | WRh | 25_50 | 0.575 | 31 | 0.199
FDA | WRh | 25_50 | 0.575 | 32 | 0.2
FDA | WRh | 25_50 | 0.575 | 33 | 0.201
FDA | WRh | 25_50 | 0.575 | 34 | 0.201
FDA | WRh | 25_50 | 0.575 | 35 | 0.202
FDA | WRh | 25_50 | 0.575 | 36 | 0.203
FDA | WRh | 25_50 | 0.575 | 37 | 0.204
FDA | WRh | 25_50 | 0.575 | 38 | 0.205
FDA | WRh | 25_50 | 0.575 | 39 | 0.206
FDA | WRh | 25_50 | 0.575 | 40 | 0.208
FDA | WRh | 25_50 | 0.575 | 41 | 0.209
FDA | WRh | 25_50 | 0.575 | 42 | 0.21
FDA | WRh | 25_50 | 0.575 | 43 | 0.212
FDA | WRh | 25_50 | 0.575 | 44 | 0.213
FDA | WRh | 25_50 | 0.575 | 45 | 0.215
FDA | WRh | 25_50 | 0.575 | 46 | 0.216
FDA | WRh | 25_50 | 0.575 | 47 | 0.218
FDA | WRh | 25_50 | 0.575 | 48 | 0.22
FDA | WRh | 25_50 | 0.575 | 49 | 0.221
FDA | WRh | 25_50 | 0.6 | 22 | 0.195
FDA | WRh | 25_50 | 0.6 | 23 | 0.198
FDA | WRh | 25_50 | 0.6 | 24 | 0.201
FDA | WRh | 25_50 | 0.6 | 25 | 0.204
FDA | WRh | 25_50 | 0.6 | 26 | 0.205
FDA | WRh | 25_50 | 0.6 | 27 | 0.206
FDA | WRh | 25_50 | 0.6 | 28 | 0.206
FDA | WRh | 25_50 | 0.6 | 29 | 0.207
FDA | WRh | 25_50 | 0.6 | 30 | 0.207
FDA | WRh | 25_50 | 0.6 | 31 | 0.207
FDA | WRh | 25_50 | 0.6 | 32 | 0.208
FDA | WRh | 25_50 | 0.6 | 33 | 0.208
FDA | WRh | 25_50 | 0.6 | 34 | 0.209
FDA | WRh | 25_50 | 0.6 | 35 | 0.21
FDA | WRh | 25_50 | 0.6 | 36 | 0.21
FDA | WRh | 25_50 | 0.6 | 37 | 0.211
FDA | WRh | 25_50 | 0.6 | 38 | 0.212
FDA | WRh | 25_50 | 0.6 | 39 | 0.213
FDA | WRh | 25_50 | 0.6 | 40 | 0.215
FDA | WRh | 25_50 | 0.6 | 41 | 0.216
FDA | WRh | 25_50 | 0.6 | 42 | 0.217
FDA | WRh | 25_50 | 0.6 | 43 | 0.219
FDA | WRh | 25_50 | 0.6 | 44 | 0.22
FDA | WRh | 25_50 | 0.6 | 45 | 0.222
FDA | WRh | 25_50 | 0.6 | 46 | 0.223
FDA | WRh | 25_50 | 0.6 | 47 | 0.225
FDA | WRh | 25_50 | 0.6 | 48 | 0.226
FDA | WRh | 25_50 | 0.6 | 49 | 0.228
FDA | WRh | 25_50 | 0.625 | 22 | 0.204
FDA | WRh | 25_50 | 0.625 | 23 | 0.207
FDA | WRh | 25_50 | 0.625 | 24 | 0.21
FDA | WRh | 25_50 | 0.625 | 25 | 0.212
FDA | WRh | 25_50 | 0.625 | 26 | 0.213
FDA | WRh | 25_50 | 0.625 | 27 | 0.214
FDA | WRh | 25_50 | 0.625 | 28 | 0.214
FDA | WRh | 25_50 | 0.625 | 29 | 0.215
FDA | WRh | 25_50 | 0.625 | 30 | 0.215
FDA | WRh | 25_50 | 0.625 | 31 | 0.215
FDA | WRh | 25_50 | 0.625 | 32 | 0.216
FDA | WRh | 25_50 | 0.625 | 33 | 0.216
FDA | WRh | 25_50 | 0.625 | 34 | 0.217
FDA | WRh | 25_50 | 0.625 | 35 | 0.217
FDA | WRh | 25_50 | 0.625 | 36 | 0.218
FDA | WRh | 25_50 | 0.625 | 37 | 0.219
FDA | WRh | 25_50 | 0.625 | 38 | 0.219
FDA | WRh | 25_50 | 0.625 | 39 | 0.221
FDA | WRh | 25_50 | 0.625 | 40 | 0.222
FDA | WRh | 25_50 | 0.625 | 41 | 0.223
FDA | WRh | 25_50 | 0.625 | 42 | 0.224
FDA | WRh | 25_50 | 0.625 | 43 | 0.226
FDA | WRh | 25_50 | 0.625 | 44 | 0.227
FDA | WRh | 25_50 | 0.625 | 45 | 0.228
FDA | WRh | 25_50 | 0.625 | 46 | 0.23
FDA | WRh | 25_50 | 0.625 | 47 | 0.232
FDA | WRh | 25_50 | 0.625 | 48 | 0.233
FDA | WRh | 25_50 | 0.625 | 49 | 0.235
FDA | WRh | 25_50 | 0.65 | 23 | 0.215
FDA | WRh | 25_50 | 0.65 | 24 | 0.218
FDA | WRh | 25_50 | 0.65 | 25 | 0.22
FDA | WRh | 25_50 | 0.65 | 26 | 0.221
FDA | WRh | 25_50 | 0.65 | 27 | 0.222
FDA | WRh | 25_50 | 0.65 | 28 | 0.222
FDA | WRh | 25_50 | 0.65 | 29 | 0.222
FDA | WRh | 25_50 | 0.65 | 30 | 0.223
FDA | WRh | 25_50 | 0.65 | 31 | 0.223
FDA | WRh | 25_50 | 0.65 | 32 | 0.223
FDA | WRh | 25_50 | 0.65 | 33 | 0.224
FDA | WRh | 25_50 | 0.65 | 34 | 0.224
FDA | WRh | 25_50 | 0.65 | 35 | 0.225
FDA | WRh | 25_50 | 0.65 | 36 | 0.225
FDA | WRh | 25_50 | 0.65 | 37 | 0.226
FDA | WRh | 25_50 | 0.65 | 38 | 0.227
FDA | WRh | 25_50 | 0.65 | 39 | 0.228
FDA | WRh | 25_50 | 0.65 | 40 | 0.229
FDA | WRh | 25_50 | 0.65 | 41 | 0.23
FDA | WRh | 25_50 | 0.65 | 42 | 0.231
FDA | WRh | 25_50 | 0.65 | 43 | 0.233
FDA | WRh | 25_50 | 0.65 | 44 | 0.234
FDA | WRh | 25_50 | 0.65 | 45 | 0.235
FDA | WRh | 25_50 | 0.65 | 46 | 0.237
FDA | WRh | 25_50 | 0.65 | 47 | 0.238
FDA | WRh | 25_50 | 0.65 | 48 | 0.24
FDA | WRh | 25_50 | 0.65 | 49 | 0.242
FDA | WRh | 25_50 | 0.675 | 25 | 0.228
FDA | WRh | 25_50 | 0.675 | 26 | 0.229
FDA | WRh | 25_50 | 0.675 | 27 | 0.23
FDA | WRh | 25_50 | 0.675 | 28 | 0.23
FDA | WRh | 25_50 | 0.675 | 29 | 0.23
FDA | WRh | 25_50 | 0.675 | 30 | 0.231
FDA | WRh | 25_50 | 0.675 | 31 | 0.231
FDA | WRh | 25_50 | 0.675 | 32 | 0.231
FDA | WRh | 25_50 | 0.675 | 33 | 0.231
FDA | WRh | 25_50 | 0.675 | 34 | 0.232
FDA | WRh | 25_50 | 0.675 | 35 | 0.232
FDA | WRh | 25_50 | 0.675 | 36 | 0.233
FDA | WRh | 25_50 | 0.675 | 37 | 0.233
FDA | WRh | 25_50 | 0.675 | 38 | 0.234
FDA | WRh | 25_50 | 0.675 | 39 | 0.235
FDA | WRh | 25_50 | 0.675 | 40 | 0.236
FDA | WRh | 25_50 | 0.675 | 41 | 0.237
FDA | WRh | 25_50 | 0.675 | 42 | 0.238
FDA | WRh | 25_50 | 0.675 | 43 | 0.239
FDA | WRh | 25_50 | 0.675 | 44 | 0.241
FDA | WRh | 25_50 | 0.675 | 45 | 0.242
FDA | WRh | 25_50 | 0.675 | 46 | 0.244
FDA | WRh | 25_50 | 0.675 | 47 | 0.245
FDA | WRh | 25_50 | 0.675 | 48 | 0.247
FDA | WRh | 25_50 | 0.675 | 49 | 0.248
FDA | WRh | 25_50 | 0.7 | 27 | 0.238
FDA | WRh | 25_50 | 0.7 | 28 | 0.238
FDA | WRh | 25_50 | 0.7 | 29 | 0.238
FDA | WRh | 25_50 | 0.7 | 30 | 0.238
FDA | WRh | 25_50 | 0.7 | 31 | 0.239
FDA | WRh | 25_50 | 0.7 | 32 | 0.239
FDA | WRh | 25_50 | 0.7 | 33 | 0.239
FDA | WRh | 25_50 | 0.7 | 34 | 0.239
FDA | WRh | 25_50 | 0.7 | 35 | 0.24
FDA | WRh | 25_50 | 0.7 | 36 | 0.24
FDA | WRh | 25_50 | 0.7 | 37 | 0.24
FDA | WRh | 25_50 | 0.7 | 38 | 0.241
FDA | WRh | 25_50 | 0.7 | 39 | 0.242
FDA | WRh | 25_50 | 0.7 | 40 | 0.243
FDA | WRh | 25_50 | 0.7 | 41 | 0.244
FDA | WRh | 25_50 | 0.7 | 42 | 0.245
FDA | WRh | 25_50 | 0.7 | 43 | 0.246
FDA | WRh | 25_50 | 0.7 | 44 | 0.247
FDA | WRh | 25_50 | 0.7 | 45 | 0.249
FDA | WRh | 25_50 | 0.7 | 46 | 0.25
FDA | WRh | 25_50 | 0.7 | 47 | 0.252
FDA | WRh | 25_50 | 0.7 | 48 | 0.253
FDA | WRh | 25_50 | 0.7 | 49 | 0.255
FDA | WRh | 25_50 | 0.725 | 33 | 0.246
FDA | WRh | 25_50 | 0.725 | 34 | 0.247
FDA | WRh | 25_50 | 0.725 | 35 | 0.247
FDA | WRh | 25_50 | 0.725 | 36 | 0.247
FDA | WRh | 25_50 | 0.725 | 37 | 0.248
FDA | WRh | 25_50 | 0.725 | 38 | 0.248
FDA | WRh | 25_50 | 0.725 | 39 | 0.249
FDA | WRh | 25_50 | 0.725 | 40 | 0.25
FDA | WRh | 25_50 | 0.725 | 41 | 0.251
FDA | WRh | 25_50 | 0.725 | 42 | 0.252
FDA | WRh | 25_50 | 0.725 | 43 | 0.253
FDA | WRh | 25_50 | 0.725 | 44 | 0.254
FDA | WRh | 25_50 | 0.725 | 45 | 0.255
FDA | WRh | 25_50 | 0.725 | 46 | 0.257
FDA | WRh | 25_50 | 0.725 | 47 | 0.258
FDA | WRh | 25_50 | 0.725 | 48 | 0.26
FDA | WRh | 25_50 | 0.725 | 49 | 0.262
FDA | WRh | 25_50 | 0.75 | 37 | 0.255
FDA | WRh | 25_50 | 0.75 | 38 | 0.256
FDA | WRh | 25_50 | 0.75 | 39 | 0.256
FDA | WRh | 25_50 | 0.75 | 40 | 0.257
FDA | WRh | 25_50 | 0.75 | 41 | 0.258
FDA | WRh | 25_50 | 0.75 | 42 | 0.258
FDA | WRh | 25_50 | 0.75 | 43 | 0.26
FDA | WRh | 25_50 | 0.75 | 44 | 0.261
FDA | WRh | 25_50 | 0.75 | 45 | 0.262
FDA | WRh | 25_50 | 0.75 | 46 | 0.263
FDA | WRh | 25_50 | 0.75 | 47 | 0.265
FDA | WRh | 25_50 | 0.75 | 48 | 0.266
FDA | WRh | 25_50 | 0.75 | 49 | 0.268
FDA | WRh | 25_50 | 0.775 | 43 | 0.266
FDA | WRh | 25_50 | 0.775 | 44 | 0.268
FDA | WRh | 25_50 | 0.775 | 45 | 0.269
FDA | WRh | 25_50 | 0.775 | 46 | 0.27
FDA | WRh | 25_50 | 0.775 | 47 | 0.271
FDA | WRh | 25_50 | 0.775 | 48 | 0.273
FDA | WRh | 25_50 | 0.775 | 49 | 0.275
FDA | WRh | 25_50 | 0.8 | 48 | 0.279
FDA | WRh | 25_50 | 0.8 | 49 | 0.281
US50 | MoMo | 09_25 | 0.2 | 20 | 0.054
US50 | MoMo | 09_25 | 0.2 | 21 | 0.058
US50 | MoMo | 09_25 | 0.2 | 22 | 0.061
US50 | MoMo | 09_25 | 0.2 | 23 | 0.063
US50 | MoMo | 09_25 | 0.2 | 24 | 0.065
US50 | MoMo | 09_25 | 0.2 | 25 | 0.067
US50 | MoMo | 09_25 | 0.21 | 20 | 0.059
US50 | MoMo | 09_25 | 0.21 | 21 | 0.063
US50 | MoMo | 09_25 | 0.21 | 22 | 0.065
US50 | MoMo | 09_25 | 0.21 | 23 | 0.067
US50 | MoMo | 09_25 | 0.21 | 24 | 0.069
US50 | MoMo | 09_25 | 0.21 | 25 | 0.071
US50 | MoMo | 09_25 | 0.22 | 20 | 0.063
US50 | MoMo | 09_25 | 0.22 | 21 | 0.067
US50 | MoMo | 09_25 | 0.22 | 22 | 0.07
US50 | MoMo | 09_25 | 0.22 | 23 | 0.072
US50 | MoMo | 09_25 | 0.22 | 24 | 0.074
US50 | MoMo | 09_25 | 0.22 | 25 | 0.075
US50 | MoMo | 09_25 | 0.22 | 26 | 0.077
US50 | MoMo | 09_25 | 0.23 | 20 | 0.067
US50 | MoMo | 09_25 | 0.23 | 21 | 0.072
US50 | MoMo | 09_25 | 0.23 | 22 | 0.074
US50 | MoMo | 09_25 | 0.23 | 23 | 0.076
US50 | MoMo | 09_25 | 0.23 | 24 | 0.078
US50 | MoMo | 09_25 | 0.23 | 25 | 0.08
US50 | MoMo | 09_25 | 0.23 | 26 | 0.081
US50 | MoMo | 09_25 | 0.23 | 27 | 0.083
US50 | MoMo | 09_25 | 0.23 | 28 | 0.084
US50 | MoMo | 09_25 | 0.24 | 20 | 0.072
US50 | MoMo | 09_25 | 0.24 | 21 | 0.076
US50 | MoMo | 09_25 | 0.24 | 22 | 0.079
US50 | MoMo | 09_25 | 0.24 | 23 | 0.081
US50 | MoMo | 09_25 | 0.24 | 24 | 0.082
US50 | MoMo | 09_25 | 0.24 | 25 | 0.084
US50 | MoMo | 09_25 | 0.24 | 26 | 0.085
US50 | MoMo | 09_25 | 0.24 | 27 | 0.087
US50 | MoMo | 09_25 | 0.24 | 28 | 0.088
US50 | MoMo | 09_25 | 0.24 | 29 | 0.089
US50 | MoMo | 09_25 | 0.25 | 20 | 0.076
US50 | MoMo | 09_25 | 0.25 | 21 | 0.08
US50 | MoMo | 09_25 | 0.25 | 22 | 0.083
US50 | MoMo | 09_25 | 0.25 | 23 | 0.085
US50 | MoMo | 09_25 | 0.25 | 24 | 0.087
US50 | MoMo | 09_25 | 0.25 | 25 | 0.088
US50 | MoMo | 09_25 | 0.25 | 26 | 0.089
US50 | MoMo | 09_25 | 0.25 | 27 | 0.091
US50 | MoMo | 09_25 | 0.25 | 28 | 0.092
US50 | MoMo | 09_25 | 0.25 | 29 | 0.093
US50 | MoMo | 09_25 | 0.25 | 30 | 0.094
US50 | MoMo | 09_25 | 0.25 | 31 | 0.096
US50 | MoMo | 09_25 | 0.26 | 20 | 0.08
US50 | MoMo | 09_25 | 0.26 | 21 | 0.085
US50 | MoMo | 09_25 | 0.26 | 22 | 0.087
US50 | MoMo | 09_25 | 0.26 | 23 | 0.089
US50 | MoMo | 09_25 | 0.26 | 24 | 0.091
US50 | MoMo | 09_25 | 0.26 | 25 | 0.092
US50 | MoMo | 09_25 | 0.26 | 26 | 0.093
US50 | MoMo | 09_25 | 0.26 | 27 | 0.095
US50 | MoMo | 09_25 | 0.26 | 28 | 0.096
US50 | MoMo | 09_25 | 0.26 | 29 | 0.097
US50 | MoMo | 09_25 | 0.26 | 30 | 0.098
US50 | MoMo | 09_25 | 0.26 | 31 | 0.099
US50 | MoMo | 09_25 | 0.26 | 32 | 0.1
US50 | MoMo | 09_25 | 0.27 | 20 | 0.084
US50 | MoMo | 09_25 | 0.27 | 21 | 0.089
US50 | MoMo | 09_25 | 0.27 | 22 | 0.091
US50 | MoMo | 09_25 | 0.27 | 23 | 0.093
US50 | MoMo | 09_25 | 0.27 | 24 | 0.095
US50 | MoMo | 09_25 | 0.27 | 25 | 0.096
US50 | MoMo | 09_25 | 0.27 | 26 | 0.097
US50 | MoMo | 09_25 | 0.27 | 27 | 0.098
US50 | MoMo | 09_25 | 0.27 | 28 | 0.1
US50 | MoMo | 09_25 | 0.27 | 29 | 0.101
US50 | MoMo | 09_25 | 0.27 | 30 | 0.102
US50 | MoMo | 09_25 | 0.27 | 31 | 0.103
US50 | MoMo | 09_25 | 0.27 | 32 | 0.104
US50 | MoMo | 09_25 | 0.27 | 33 | 0.105
US50 | MoMo | 09_25 | 0.27 | 34 | 0.106
US50 | MoMo | 09_25 | 0.28 | 20 | 0.089
US50 | MoMo | 09_25 | 0.28 | 21 | 0.093
US50 | MoMo | 09_25 | 0.28 | 22 | 0.096
US50 | MoMo | 09_25 | 0.28 | 23 | 0.097
US50 | MoMo | 09_25 | 0.28 | 24 | 0.099
US50 | MoMo | 09_25 | 0.28 | 25 | 0.1
US50 | MoMo | 09_25 | 0.28 | 26 | 0.101
US50 | MoMo | 09_25 | 0.28 | 27 | 0.102
US50 | MoMo | 09_25 | 0.28 | 28 | 0.103
US50 | MoMo | 09_25 | 0.28 | 29 | 0.104
US50 | MoMo | 09_25 | 0.28 | 30 | 0.105
US50 | MoMo | 09_25 | 0.28 | 31 | 0.106
US50 | MoMo | 09_25 | 0.28 | 32 | 0.107
US50 | MoMo | 09_25 | 0.28 | 33 | 0.108
US50 | MoMo | 09_25 | 0.28 | 34 | 0.109
US50 | MoMo | 09_25 | 0.28 | 35 | 0.11
US50 | MoMo | 09_25 | 0.28 | 36 | 0.111
US50 | MoMo | 09_25 | 0.29 | 20 | 0.093
US50 | MoMo | 09_25 | 0.29 | 21 | 0.097
US50 | MoMo | 09_25 | 0.29 | 22 | 0.1
US50 | MoMo | 09_25 | 0.29 | 23 | 0.101
US50 | MoMo | 09_25 | 0.29 | 24 | 0.103
US50 | MoMo | 09_25 | 0.29 | 25 | 0.104
US50 | MoMo | 09_25 | 0.29 | 26 | 0.105
US50 | MoMo | 09_25 | 0.29 | 27 | 0.106
US50 | MoMo | 09_25 | 0.29 | 28 | 0.107
US50 | MoMo | 09_25 | 0.29 | 29 | 0.108
US50 | MoMo | 09_25 | 0.29 | 30 | 0.109
US50 | MoMo | 09_25 | 0.29 | 31 | 0.11
US50 | MoMo | 09_25 | 0.29 | 32 | 0.111
US50 | MoMo | 09_25 | 0.29 | 33 | 0.112
US50 | MoMo | 09_25 | 0.29 | 34 | 0.113
US50 | MoMo | 09_25 | 0.29 | 35 | 0.114
US50 | MoMo | 09_25 | 0.29 | 36 | 0.115
US50 | MoMo | 09_25 | 0.29 | 37 | 0.115
US50 | MoMo | 09_25 | 0.29 | 38 | 0.116
US50 | MoMo | 09_25 | 0.3 | 20 | 0.097
US50 | MoMo | 09_25 | 0.3 | 21 | 0.102
US50 | MoMo | 09_25 | 0.3 | 22 | 0.104
US50 | MoMo | 09_25 | 0.3 | 23 | 0.106
US50 | MoMo | 09_25 | 0.3 | 24 | 0.107
US50 | MoMo | 09_25 | 0.3 | 25 | 0.108
US50 | MoMo | 09_25 | 0.3 | 26 | 0.109
US50 | MoMo | 09_25 | 0.3 | 27 | 0.11
US50 | MoMo | 09_25 | 0.3 | 28 | 0.111
US50 | MoMo | 09_25 | 0.3 | 29 | 0.112
US50 | MoMo | 09_25 | 0.3 | 30 | 0.112
US50 | MoMo | 09_25 | 0.3 | 31 | 0.113
US50 | MoMo | 09_25 | 0.3 | 32 | 0.114
US50 | MoMo | 09_25 | 0.3 | 33 | 0.115
US50 | MoMo | 09_25 | 0.3 | 34 | 0.116
US50 | MoMo | 09_25 | 0.3 | 35 | 0.117
US50 | MoMo | 09_25 | 0.3 | 36 | 0.118
US50 | MoMo | 09_25 | 0.3 | 37 | 0.119
US50 | MoMo | 09_25 | 0.3 | 38 | 0.12
US50 | MoMo | 09_25 | 0.3 | 39 | 0.12
US50 | MoMo | 09_25 | 0.3 | 40 | 0.121
US50 | MoMo | 09_25 | 0.3 | 41 | 0.122
US50 | MoMo | 09_25 | 0.3 | 42 | 0.122
US50 | MoMo | 09_25 | 0.31 | 20 | 0.101
US50 | MoMo | 09_25 | 0.31 | 21 | 0.106
US50 | MoMo | 09_25 | 0.31 | 22 | 0.108
US50 | MoMo | 09_25 | 0.31 | 23 | 0.11
US50 | MoMo | 09_25 | 0.31 | 24 | 0.111
US50 | MoMo | 09_25 | 0.31 | 25 | 0.112
US50 | MoMo | 09_25 | 0.31 | 26 | 0.113
US50 | MoMo | 09_25 | 0.31 | 27 | 0.113
US50 | MoMo | 09_25 | 0.31 | 28 | 0.114
US50 | MoMo | 09_25 | 0.31 | 29 | 0.115
US50 | MoMo | 09_25 | 0.31 | 30 | 0.116
US50 | MoMo | 09_25 | 0.31 | 31 | 0.117
US50 | MoMo | 09_25 | 0.31 | 32 | 0.118
US50 | MoMo | 09_25 | 0.31 | 33 | 0.119
US50 | MoMo | 09_25 | 0.31 | 34 | 0.119
US50 | MoMo | 09_25 | 0.31 | 35 | 0.12
US50 | MoMo | 09_25 | 0.31 | 36 | 0.121
US50 | MoMo | 09_25 | 0.31 | 37 | 0.122
US50 | MoMo | 09_25 | 0.31 | 38 | 0.123
US50 | MoMo | 09_25 | 0.31 | 39 | 0.123
US50 | MoMo | 09_25 | 0.31 | 40 | 0.124
US50 | MoMo | 09_25 | 0.31 | 41 | 0.125
US50 | MoMo | 09_25 | 0.31 | 42 | 0.126
US50 | MoMo | 09_25 | 0.31 | 43 | 0.126
US50 | MoMo | 09_25 | 0.32 | 20 | 0.106
US50 | MoMo | 09_25 | 0.32 | 21 | 0.11
US50 | MoMo | 09_25 | 0.32 | 22 | 0.112
US50 | MoMo | 09_25 | 0.32 | 23 | 0.114
US50 | MoMo | 09_25 | 0.32 | 24 | 0.115
US50 | MoMo | 09_25 | 0.32 | 25 | 0.116
US50 | MoMo | 09_25 | 0.32 | 26 | 0.116
US50 | MoMo | 09_25 | 0.32 | 27 | 0.117
US50 | MoMo | 09_25 | 0.32 | 28 | 0.118
US50 | MoMo | 09_25 | 0.32 | 29 | 0.119
US50 | MoMo | 09_25 | 0.32 | 30 | 0.119
US50 | MoMo | 09_25 | 0.32 | 31 | 0.12
US50 | MoMo | 09_25 | 0.32 | 32 | 0.121
US50 | MoMo | 09_25 | 0.32 | 33 | 0.122
US50 | MoMo | 09_25 | 0.32 | 34 | 0.123
US50 | MoMo | 09_25 | 0.32 | 35 | 0.124
US50 | MoMo | 09_25 | 0.32 | 36 | 0.124
US50 | MoMo | 09_25 | 0.32 | 37 | 0.125
US50 | MoMo | 09_25 | 0.32 | 38 | 0.126
US50 | MoMo | 09_25 | 0.32 | 39 | 0.127
US50 | MoMo | 09_25 | 0.32 | 40 | 0.127
US50 | MoMo | 09_25 | 0.32 | 41 | 0.128
US50 | MoMo | 09_25 | 0.32 | 42 | 0.129
US50 | MoMo | 09_25 | 0.32 | 43 | 0.129
US50 | MoMo | 09_25 | 0.32 | 44 | 0.13
US50 | MoMo | 09_25 | 0.32 | 45 | 0.13
US50 | MoMo | 09_25 | 0.32 | 46 | 0.131
US50 | MoMo | 09_25 | 0.33 | 22 | 0.116
US50 | MoMo | 09_25 | 0.33 | 23 | 0.118
US50 | MoMo | 09_25 | 0.33 | 24 | 0.119
US50 | MoMo | 09_25 | 0.33 | 25 | 0.119
US50 | MoMo | 09_25 | 0.33 | 26 | 0.12
US50 | MoMo | 09_25 | 0.33 | 27 | 0.121
US50 | MoMo | 09_25 | 0.33 | 28 | 0.121
US50 | MoMo | 09_25 | 0.33 | 29 | 0.122
US50 | MoMo | 09_25 | 0.33 | 30 | 0.123
US50 | MoMo | 09_25 | 0.33 | 31 | 0.123
US50 | MoMo | 09_25 | 0.33 | 32 | 0.124
US50 | MoMo | 09_25 | 0.33 | 33 | 0.125
US50 | MoMo | 09_25 | 0.33 | 34 | 0.126
US50 | MoMo | 09_25 | 0.33 | 35 | 0.127
US50 | MoMo | 09_25 | 0.33 | 36 | 0.128
US50 | MoMo | 09_25 | 0.33 | 37 | 0.128
US50 | MoMo | 09_25 | 0.33 | 38 | 0.129
US50 | MoMo | 09_25 | 0.33 | 39 | 0.13
US50 | MoMo | 09_25 | 0.33 | 40 | 0.131
US50 | MoMo | 09_25 | 0.33 | 41 | 0.131
US50 | MoMo | 09_25 | 0.33 | 42 | 0.132
US50 | MoMo | 09_25 | 0.33 | 43 | 0.132
US50 | MoMo | 09_25 | 0.33 | 44 | 0.133
US50 | MoMo | 09_25 | 0.33 | 45 | 0.134
US50 | MoMo | 09_25 | 0.33 | 46 | 0.134
US50 | MoMo | 09_25 | 0.33 | 47 | 0.134
US50 | MoMo | 09_25 | 0.33 | 48 | 0.135
US50 | MoMo | 09_25 | 0.33 | 49 | 0.135
US50 | MoMo | 09_25 | 0.34 | 22 | 0.12
US50 | MoMo | 09_25 | 0.34 | 23 | 0.122
US50 | MoMo | 09_25 | 0.34 | 24 | 0.122
US50 | MoMo | 09_25 | 0.34 | 25 | 0.123
US50 | MoMo | 09_25 | 0.34 | 26 | 0.124
US50 | MoMo | 09_25 | 0.34 | 27 | 0.124
US50 | MoMo | 09_25 | 0.34 | 28 | 0.125
US50 | MoMo | 09_25 | 0.34 | 29 | 0.126
US50 | MoMo | 09_25 | 0.34 | 30 | 0.126
US50 | MoMo | 09_25 | 0.34 | 31 | 0.127
US50 | MoMo | 09_25 | 0.34 | 32 | 0.128
US50 | MoMo | 09_25 | 0.34 | 33 | 0.128
US50 | MoMo | 09_25 | 0.34 | 34 | 0.129
US50 | MoMo | 09_25 | 0.34 | 35 | 0.13
US50 | MoMo | 09_25 | 0.34 | 36 | 0.131
US50 | MoMo | 09_25 | 0.34 | 37 | 0.131
US50 | MoMo | 09_25 | 0.34 | 38 | 0.132
US50 | MoMo | 09_25 | 0.34 | 39 | 0.133
US50 | MoMo | 09_25 | 0.34 | 40 | 0.134
US50 | MoMo | 09_25 | 0.34 | 41 | 0.134
US50 | MoMo | 09_25 | 0.34 | 42 | 0.135
US50 | MoMo | 09_25 | 0.34 | 43 | 0.136
US50 | MoMo | 09_25 | 0.34 | 44 | 0.136
US50 | MoMo | 09_25 | 0.34 | 45 | 0.137
US50 | MoMo | 09_25 | 0.34 | 46 | 0.137
US50 | MoMo | 09_25 | 0.34 | 47 | 0.138
US50 | MoMo | 09_25 | 0.34 | 48 | 0.138
US50 | MoMo | 09_25 | 0.34 | 49 | 0.139
US50 | MoMo | 09_25 | 0.35 | 22 | 0.125
US50 | MoMo | 09_25 | 0.35 | 23 | 0.126
US50 | MoMo | 09_25 | 0.35 | 24 | 0.126
US50 | MoMo | 09_25 | 0.35 | 25 | 0.127
US50 | MoMo | 09_25 | 0.35 | 26 | 0.127
US50 | MoMo | 09_25 | 0.35 | 27 | 0.128
US50 | MoMo | 09_25 | 0.35 | 28 | 0.128
US50 | MoMo | 09_25 | 0.35 | 29 | 0.129
US50 | MoMo | 09_25 | 0.35 | 30 | 0.13
US50 | MoMo | 09_25 | 0.35 | 31 | 0.13
US50 | MoMo | 09_25 | 0.35 | 32 | 0.131
US50 | MoMo | 09_25 | 0.35 | 33 | 0.132
US50 | MoMo | 09_25 | 0.35 | 34 | 0.132
US50 | MoMo | 09_25 | 0.35 | 35 | 0.133
US50 | MoMo | 09_25 | 0.35 | 36 | 0.134
US50 | MoMo | 09_25 | 0.35 | 37 | 0.135
US50 | MoMo | 09_25 | 0.35 | 38 | 0.135
US50 | MoMo | 09_25 | 0.35 | 39 | 0.136
US50 | MoMo | 09_25 | 0.35 | 40 | 0.137
US50 | MoMo | 09_25 | 0.35 | 41 | 0.137
US50 | MoMo | 09_25 | 0.35 | 42 | 0.138
US50 | MoMo | 09_25 | 0.35 | 43 | 0.139
US50 | MoMo | 09_25 | 0.35 | 44 | 0.139
US50 | MoMo | 09_25 | 0.35 | 45 | 0.14
US50 | MoMo | 09_25 | 0.35 | 46 | 0.14
US50 | MoMo | 09_25 | 0.35 | 47 | 0.141
US50 | MoMo | 09_25 | 0.35 | 48 | 0.141
US50 | MoMo | 09_25 | 0.35 | 49 | 0.142
US50 | MoMo | 09_25 | 0.36 | 23 | 0.13
US50 | MoMo | 09_25 | 0.36 | 24 | 0.13
US50 | MoMo | 09_25 | 0.36 | 25 | 0.131
US50 | MoMo | 09_25 | 0.36 | 26 | 0.131
US50 | MoMo | 09_25 | 0.36 | 27 | 0.132
US50 | MoMo | 09_25 | 0.36 | 28 | 0.132
US50 | MoMo | 09_25 | 0.36 | 29 | 0.133
US50 | MoMo | 09_25 | 0.36 | 30 | 0.133
US50 | MoMo | 09_25 | 0.36 | 31 | 0.134
US50 | MoMo | 09_25 | 0.36 | 32 | 0.134
US50 | MoMo | 09_25 | 0.36 | 33 | 0.135
US50 | MoMo | 09_25 | 0.36 | 34 | 0.136
US50 | MoMo | 09_25 | 0.36 | 35 | 0.136
US50 | MoMo | 09_25 | 0.36 | 36 | 0.137
US50 | MoMo | 09_25 | 0.36 | 37 | 0.138
US50 | MoMo | 09_25 | 0.36 | 38 | 0.139
US50 | MoMo | 09_25 | 0.36 | 39 | 0.139
US50 | MoMo | 09_25 | 0.36 | 40 | 0.14
US50 | MoMo | 09_25 | 0.36 | 41 | 0.141
US50 | MoMo | 09_25 | 0.36 | 42 | 0.141
US50 | MoMo | 09_25 | 0.36 | 43 | 0.142
US50 | MoMo | 09_25 | 0.36 | 44 | 0.142
US50 | MoMo | 09_25 | 0.36 | 45 | 0.143
US50 | MoMo | 09_25 | 0.36 | 46 | 0.143
US50 | MoMo | 09_25 | 0.36 | 47 | 0.144
US50 | MoMo | 09_25 | 0.36 | 48 | 0.144
US50 | MoMo | 09_25 | 0.36 | 49 | 0.145
US50 | MoMo | 09_25 | 0.37 | 24 | 0.134
US50 | MoMo | 09_25 | 0.37 | 25 | 0.135
US50 | MoMo | 09_25 | 0.37 | 26 | 0.135
US50 | MoMo | 09_25 | 0.37 | 27 | 0.135
US50 | MoMo | 09_25 | 0.37 | 28 | 0.136
US50 | MoMo | 09_25 | 0.37 | 29 | 0.136
US50 | MoMo | 09_25 | 0.37 | 30 | 0.136
US50 | MoMo | 09_25 | 0.37 | 31 | 0.137
US50 | MoMo | 09_25 | 0.37 | 32 | 0.138
US50 | MoMo | 09_25 | 0.37 | 33 | 0.138
US50 | MoMo | 09_25 | 0.37 | 34 | 0.139
US50 | MoMo | 09_25 | 0.37 | 35 | 0.14
US50 | MoMo | 09_25 | 0.37 | 36 | 0.14
US50 | MoMo | 09_25 | 0.37 | 37 | 0.141
US50 | MoMo | 09_25 | 0.37 | 38 | 0.142
US50 | MoMo | 09_25 | 0.37 | 39 | 0.142
US50 | MoMo | 09_25 | 0.37 | 40 | 0.143
US50 | MoMo | 09_25 | 0.37 | 41 | 0.144
US50 | MoMo | 09_25 | 0.37 | 42 | 0.144
US50 | MoMo | 09_25 | 0.37 | 43 | 0.145
US50 | MoMo | 09_25 | 0.37 | 44 | 0.145
US50 | MoMo | 09_25 | 0.37 | 45 | 0.146
US50 | MoMo | 09_25 | 0.37 | 46 | 0.146
US50 | MoMo | 09_25 | 0.37 | 47 | 0.147
US50 | MoMo | 09_25 | 0.37 | 48 | 0.147
US50 | MoMo | 09_25 | 0.37 | 49 | 0.148
US50 | MoMo | 09_25 | 0.38 | 26 | 0.139
US50 | MoMo | 09_25 | 0.38 | 27 | 0.139
US50 | MoMo | 09_25 | 0.38 | 28 | 0.139
US50 | MoMo | 09_25 | 0.38 | 29 | 0.14
US50 | MoMo | 09_25 | 0.38 | 30 | 0.14
US50 | MoMo | 09_25 | 0.38 | 31 | 0.14
US50 | MoMo | 09_25 | 0.38 | 32 | 0.141
US50 | MoMo | 09_25 | 0.38 | 33 | 0.141
US50 | MoMo | 09_25 | 0.38 | 34 | 0.142
US50 | MoMo | 09_25 | 0.38 | 35 | 0.143
US50 | MoMo | 09_25 | 0.38 | 36 | 0.143
US50 | MoMo | 09_25 | 0.38 | 37 | 0.144
US50 | MoMo | 09_25 | 0.38 | 38 | 0.145
US50 | MoMo | 09_25 | 0.38 | 39 | 0.145
US50 | MoMo | 09_25 | 0.38 | 40 | 0.146
US50 | MoMo | 09_25 | 0.38 | 41 | 0.147
US50 | MoMo | 09_25 | 0.38 | 42 | 0.147
US50 | MoMo | 09_25 | 0.38 | 43 | 0.148
US50 | MoMo | 09_25 | 0.38 | 44 | 0.148
US50 | MoMo | 09_25 | 0.38 | 45 | 0.149
US50 | MoMo | 09_25 | 0.38 | 46 | 0.149
US50 | MoMo | 09_25 | 0.38 | 47 | 0.15
US50 | MoMo | 09_25 | 0.38 | 48 | 0.15
US50 | MoMo | 09_25 | 0.38 | 49 | 0.151
US50 | MoMo | 09_25 | 0.39 | 26 | 0.143
US50 | MoMo | 09_25 | 0.39 | 27 | 0.143
US50 | MoMo | 09_25 | 0.39 | 28 | 0.143
US50 | MoMo | 09_25 | 0.39 | 29 | 0.143
US50 | MoMo | 09_25 | 0.39 | 30 | 0.143
US50 | MoMo | 09_25 | 0.39 | 31 | 0.144
US50 | MoMo | 09_25 | 0.39 | 32 | 0.144
US50 | MoMo | 09_25 | 0.39 | 33 | 0.145
US50 | MoMo | 09_25 | 0.39 | 34 | 0.145
US50 | MoMo | 09_25 | 0.39 | 35 | 0.146
US50 | MoMo | 09_25 | 0.39 | 36 | 0.147
US50 | MoMo | 09_25 | 0.39 | 37 | 0.147
US50 | MoMo | 09_25 | 0.39 | 38 | 0.148
US50 | MoMo | 09_25 | 0.39 | 39 | 0.148
US50 | MoMo | 09_25 | 0.39 | 40 | 0.149
US50 | MoMo | 09_25 | 0.39 | 41 | 0.15
US50 | MoMo | 09_25 | 0.39 | 42 | 0.15
US50 | MoMo | 09_25 | 0.39 | 43 | 0.151
US50 | MoMo | 09_25 | 0.39 | 44 | 0.152
US50 | MoMo | 09_25 | 0.39 | 45 | 0.152
US50 | MoMo | 09_25 | 0.39 | 46 | 0.153
US50 | MoMo | 09_25 | 0.39 | 47 | 0.153
US50 | MoMo | 09_25 | 0.39 | 48 | 0.154
US50 | MoMo | 09_25 | 0.39 | 49 | 0.154
US50 | MoMo | 09_25 | 0.4 | 27 | 0.146
US50 | MoMo | 09_25 | 0.4 | 28 | 0.147
US50 | MoMo | 09_25 | 0.4 | 29 | 0.147
US50 | MoMo | 09_25 | 0.4 | 30 | 0.147
US50 | MoMo | 09_25 | 0.4 | 31 | 0.147
US50 | MoMo | 09_25 | 0.4 | 32 | 0.148
US50 | MoMo | 09_25 | 0.4 | 33 | 0.148
US50 | MoMo | 09_25 | 0.4 | 34 | 0.149
US50 | MoMo | 09_25 | 0.4 | 35 | 0.149
US50 | MoMo | 09_25 | 0.4 | 36 | 0.15
US50 | MoMo | 09_25 | 0.4 | 37 | 0.15
US50 | MoMo | 09_25 | 0.4 | 38 | 0.151
US50 | MoMo | 09_25 | 0.4 | 39 | 0.152
US50 | MoMo | 09_25 | 0.4 | 40 | 0.152
US50 | MoMo | 09_25 | 0.4 | 41 | 0.153
US50 | MoMo | 09_25 | 0.4 | 42 | 0.153
US50 | MoMo | 09_25 | 0.4 | 43 | 0.154
US50 | MoMo | 09_25 | 0.4 | 44 | 0.155
US50 | MoMo | 09_25 | 0.4 | 45 | 0.155
US50 | MoMo | 09_25 | 0.4 | 46 | 0.156
US50 | MoMo | 09_25 | 0.4 | 47 | 0.156
US50 | MoMo | 09_25 | 0.4 | 48 | 0.157
US50 | MoMo | 09_25 | 0.4 | 49 | 0.157
US50 | MoMo | 09_25 | 0.41 | 29 | 0.15
US50 | MoMo | 09_25 | 0.41 | 30 | 0.151
US50 | MoMo | 09_25 | 0.41 | 31 | 0.151
US50 | MoMo | 09_25 | 0.41 | 32 | 0.151
US50 | MoMo | 09_25 | 0.41 | 33 | 0.152
US50 | MoMo | 09_25 | 0.41 | 34 | 0.152
US50 | MoMo | 09_25 | 0.41 | 35 | 0.153
US50 | MoMo | 09_25 | 0.41 | 36 | 0.153
US50 | MoMo | 09_25 | 0.41 | 37 | 0.154
US50 | MoMo | 09_25 | 0.41 | 38 | 0.154
US50 | MoMo | 09_25 | 0.41 | 39 | 0.155
US50 | MoMo | 09_25 | 0.41 | 40 | 0.155
US50 | MoMo | 09_25 | 0.41 | 41 | 0.156
US50 | MoMo | 09_25 | 0.41 | 42 | 0.157
US50 | MoMo | 09_25 | 0.41 | 43 | 0.157
US50 | MoMo | 09_25 | 0.41 | 44 | 0.158
US50 | MoMo | 09_25 | 0.41 | 45 | 0.158
US50 | MoMo | 09_25 | 0.41 | 46 | 0.159
US50 | MoMo | 09_25 | 0.41 | 47 | 0.159
US50 | MoMo | 09_25 | 0.41 | 48 | 0.16
US50 | MoMo | 09_25 | 0.41 | 49 | 0.16
US50 | MoMo | 09_25 | 0.42 | 30 | 0.154
US50 | MoMo | 09_25 | 0.42 | 31 | 0.154
US50 | MoMo | 09_25 | 0.42 | 32 | 0.155
US50 | MoMo | 09_25 | 0.42 | 33 | 0.155
US50 | MoMo | 09_25 | 0.42 | 34 | 0.155
US50 | MoMo | 09_25 | 0.42 | 35 | 0.156
US50 | MoMo | 09_25 | 0.42 | 36 | 0.156
US50 | MoMo | 09_25 | 0.42 | 37 | 0.157
US50 | MoMo | 09_25 | 0.42 | 38 | 0.157
US50 | MoMo | 09_25 | 0.42 | 39 | 0.158
US50 | MoMo | 09_25 | 0.42 | 40 | 0.159
US50 | MoMo | 09_25 | 0.42 | 41 | 0.159
US50 | MoMo | 09_25 | 0.42 | 42 | 0.16
US50 | MoMo | 09_25 | 0.42 | 43 | 0.16
US50 | MoMo | 09_25 | 0.42 | 44 | 0.161
US50 | MoMo | 09_25 | 0.42 | 45 | 0.161
US50 | MoMo | 09_25 | 0.42 | 46 | 0.162
US50 | MoMo | 09_25 | 0.42 | 47 | 0.162
US50 | MoMo | 09_25 | 0.42 | 48 | 0.163
US50 | MoMo | 09_25 | 0.42 | 49 | 0.163
US50 | MoMo | 09_25 | 0.43 | 32 | 0.158
US50 | MoMo | 09_25 | 0.43 | 33 | 0.159
US50 | MoMo | 09_25 | 0.43 | 34 | 0.159
US50 | MoMo | 09_25 | 0.43 | 35 | 0.159
US50 | MoMo | 09_25 | 0.43 | 36 | 0.16
US50 | MoMo | 09_25 | 0.43 | 37 | 0.16
US50 | MoMo | 09_25 | 0.43 | 38 | 0.161
US50 | MoMo | 09_25 | 0.43 | 39 | 0.161
US50 | MoMo | 09_25 | 0.43 | 40 | 0.162
US50 | MoMo | 09_25 | 0.43 | 41 | 0.162
US50 | MoMo | 09_25 | 0.43 | 42 | 0.163
US50 | MoMo | 09_25 | 0.43 | 43 | 0.164
US50 | MoMo | 09_25 | 0.43 | 44 | 0.164
US50 | MoMo | 09_25 | 0.43 | 45 | 0.165
US50 | MoMo | 09_25 | 0.43 | 46 | 0.165
US50 | MoMo | 09_25 | 0.43 | 47 | 0.166
US50 | MoMo | 09_25 | 0.43 | 48 | 0.166
US50 | MoMo | 09_25 | 0.43 | 49 | 0.167
US50 | MoMo | 09_25 | 0.44 | 33 | 0.162
US50 | MoMo | 09_25 | 0.44 | 34 | 0.162
US50 | MoMo | 09_25 | 0.44 | 35 | 0.163
US50 | MoMo | 09_25 | 0.44 | 36 | 0.163
US50 | MoMo | 09_25 | 0.44 | 37 | 0.164
US50 | MoMo | 09_25 | 0.44 | 38 | 0.164
US50 | MoMo | 09_25 | 0.44 | 39 | 0.165
US50 | MoMo | 09_25 | 0.44 | 40 | 0.165
US50 | MoMo | 09_25 | 0.44 | 41 | 0.166
US50 | MoMo | 09_25 | 0.44 | 42 | 0.166
US50 | MoMo | 09_25 | 0.44 | 43 | 0.167
US50 | MoMo | 09_25 | 0.44 | 44 | 0.167
US50 | MoMo | 09_25 | 0.44 | 45 | 0.168
US50 | MoMo | 09_25 | 0.44 | 46 | 0.168
US50 | MoMo | 09_25 | 0.44 | 47 | 0.169
US50 | MoMo | 09_25 | 0.44 | 48 | 0.17
US50 | MoMo | 09_25 | 0.44 | 49 | 0.17
US50 | MoMo | 09_25 | 0.45 | 35 | 0.166
US50 | MoMo | 09_25 | 0.45 | 36 | 0.167
US50 | MoMo | 09_25 | 0.45 | 37 | 0.167
US50 | MoMo | 09_25 | 0.45 | 38 | 0.168
US50 | MoMo | 09_25 | 0.45 | 39 | 0.168
US50 | MoMo | 09_25 | 0.45 | 40 | 0.169
US50 | MoMo | 09_25 | 0.45 | 41 | 0.169
US50 | MoMo | 09_25 | 0.45 | 42 | 0.17
US50 | MoMo | 09_25 | 0.45 | 43 | 0.17
US50 | MoMo | 09_25 | 0.45 | 44 | 0.171
US50 | MoMo | 09_25 | 0.45 | 45 | 0.171
US50 | MoMo | 09_25 | 0.45 | 46 | 0.172
US50 | MoMo | 09_25 | 0.45 | 47 | 0.172
US50 | MoMo | 09_25 | 0.45 | 48 | 0.173
US50 | MoMo | 09_25 | 0.45 | 49 | 0.173
US50 | MoMo | 09_25 | 0.46 | 37 | 0.171
US50 | MoMo | 09_25 | 0.46 | 38 | 0.171
US50 | MoMo | 09_25 | 0.46 | 39 | 0.172
US50 | MoMo | 09_25 | 0.46 | 40 | 0.172
US50 | MoMo | 09_25 | 0.46 | 41 | 0.173
US50 | MoMo | 09_25 | 0.46 | 42 | 0.173
US50 | MoMo | 09_25 | 0.46 | 43 | 0.174
US50 | MoMo | 09_25 | 0.46 | 44 | 0.174
US50 | MoMo | 09_25 | 0.46 | 45 | 0.175
US50 | MoMo | 09_25 | 0.46 | 46 | 0.175
US50 | MoMo | 09_25 | 0.46 | 47 | 0.176
US50 | MoMo | 09_25 | 0.46 | 48 | 0.177
US50 | MoMo | 09_25 | 0.46 | 49 | 0.177
US50 | MoMo | 09_25 | 0.47 | 39 | 0.175
US50 | MoMo | 09_25 | 0.47 | 40 | 0.176
US50 | MoMo | 09_25 | 0.47 | 41 | 0.176
US50 | MoMo | 09_25 | 0.47 | 42 | 0.177
US50 | MoMo | 09_25 | 0.47 | 43 | 0.177
US50 | MoMo | 09_25 | 0.47 | 44 | 0.178
US50 | MoMo | 09_25 | 0.47 | 45 | 0.179
US50 | MoMo | 09_25 | 0.47 | 46 | 0.179
US50 | MoMo | 09_25 | 0.47 | 47 | 0.18
US50 | MoMo | 09_25 | 0.47 | 48 | 0.18
US50 | MoMo | 09_25 | 0.47 | 49 | 0.181
US50 | MoMo | 09_25 | 0.48 | 43 | 0.181
US50 | MoMo | 09_25 | 0.48 | 44 | 0.182
US50 | MoMo | 09_25 | 0.48 | 45 | 0.182
US50 | MoMo | 09_25 | 0.48 | 46 | 0.183
US50 | MoMo | 09_25 | 0.48 | 47 | 0.184
US50 | MoMo | 09_25 | 0.48 | 48 | 0.184
US50 | MoMo | 09_25 | 0.48 | 49 | 0.185
US50 | MoMo | 09_25 | 0.49 | 44 | 0.186
US50 | MoMo | 09_25 | 0.49 | 45 | 0.186
US50 | MoMo | 09_25 | 0.49 | 46 | 0.187
US50 | MoMo | 09_25 | 0.49 | 47 | 0.188
US50 | MoMo | 09_25 | 0.49 | 48 | 0.188
US50 | MoMo | 09_25 | 0.49 | 49 | 0.189
US50 | MoMo | 09_25 | 0.5 | 47 | 0.192
US50 | MoMo | 09_25 | 0.5 | 48 | 0.193
US50 | MoMo | 09_25 | 0.5 | 49 | 0.193
US50 | MoRh | 09_25 | 0.195 | 20 | 0.052
US50 | MoRh | 09_25 | 0.195 | 21 | 0.056
US50 | MoRh | 09_25 | 0.195 | 22 | 0.059
US50 | MoRh | 09_25 | 0.21 | 20 | 0.059
US50 | MoRh | 09_25 | 0.21 | 21 | 0.063
US50 | MoRh | 09_25 | 0.21 | 22 | 0.066
US50 | MoRh | 09_25 | 0.21 | 23 | 0.069
US50 | MoRh | 09_25 | 0.21 | 24 | 0.071
US50 | MoRh | 09_25 | 0.225 | 20 | 0.065
US50 | MoRh | 09_25 | 0.225 | 21 | 0.069
US50 | MoRh | 09_25 | 0.225 | 22 | 0.073
US50 | MoRh | 09_25 | 0.225 | 23 | 0.076
US50 | MoRh | 09_25 | 0.225 | 24 | 0.079
US50 | MoRh | 09_25 | 0.225 | 25 | 0.081
US50 | MoRh | 09_25 | 0.225 | 26 | 0.082
US50 | MoRh | 09_25 | 0.225 | 27 | 0.084
US50 | MoRh | 09_25 | 0.24 | 20 | 0.072
US50 | MoRh | 09_25 | 0.24 | 21 | 0.076
US50 | MoRh | 09_25 | 0.24 | 22 | 0.08
US50 | MoRh | 09_25 | 0.24 | 23 | 0.083
US50 | MoRh | 09_25 | 0.24 | 24 | 0.086
US50 | MoRh | 09_25 | 0.24 | 25 | 0.088
US50 | MoRh | 09_25 | 0.24 | 26 | 0.089
US50 | MoRh | 09_25 | 0.24 | 27 | 0.09
US50 | MoRh | 09_25 | 0.255 | 20 | 0.078
US50 | MoRh | 09_25 | 0.255 | 21 | 0.082
US50 | MoRh | 09_25 | 0.255 | 22 | 0.086
US50 | MoRh | 09_25 | 0.255 | 23 | 0.089
US50 | MoRh | 09_25 | 0.255 | 24 | 0.092
US50 | MoRh | 09_25 | 0.255 | 25 | 0.094
US50 | MoRh | 09_25 | 0.255 | 26 | 0.096
US50 | MoRh | 09_25 | 0.255 | 27 | 0.097
US50 | MoRh | 09_25 | 0.255 | 28 | 0.098
US50 | MoRh | 09_25 | 0.27 | 20 | 0.084
US50 | MoRh | 09_25 | 0.27 | 21 | 0.089
US50 | MoRh | 09_25 | 0.27 | 22 | 0.093
US50 | MoRh | 09_25 | 0.27 | 23 | 0.096
US50 | MoRh | 09_25 | 0.27 | 24 | 0.099
US50 | MoRh | 09_25 | 0.27 | 25 | 0.101
US50 | MoRh | 09_25 | 0.27 | 26 | 0.102
US50 | MoRh | 09_25 | 0.27 | 27 | 0.103
US50 | MoRh | 09_25 | 0.27 | 28 | 0.104
US50 | MoRh | 09_25 | 0.27 | 29 | 0.105
US50 | MoRh | 09_25 | 0.27 | 30 | 0.106
US50 | MoRh | 09_25 | 0.27 | 31 | 0.107
US50 | MoRh | 09_25 | 0.285 | 20 | 0.091
US50 | MoRh | 09_25 | 0.285 | 21 | 0.095
US50 | MoRh | 09_25 | 0.285 | 22 | 0.099
US50 | MoRh | 09_25 | 0.285 | 23 | 0.102
US50 | MoRh | 09_25 | 0.285 | 24 | 0.105
US50 | MoRh | 09_25 | 0.285 | 25 | 0.107
US50 | MoRh | 09_25 | 0.285 | 26 | 0.109
US50 | MoRh | 09_25 | 0.285 | 27 | 0.11
US50 | MoRh | 09_25 | 0.285 | 28 | 0.111
US50 | MoRh | 09_25 | 0.285 | 29 | 0.111
US50 | MoRh | 09_25 | 0.285 | 30 | 0.112
US50 | MoRh | 09_25 | 0.285 | 31 | 0.113
US50 | MoRh | 09_25 | 0.285 | 32 | 0.114
US50 | MoRh | 09_25 | 0.285 | 33 | 0.114
US50 | MoRh | 09_25 | 0.3 | 20 | 0.097
US50 | MoRh | 09_25 | 0.3 | 21 | 0.102
US50 | MoRh | 09_25 | 0.3 | 22 | 0.105
US50 | MoRh | 09_25 | 0.3 | 23 | 0.109
US50 | MoRh | 09_25 | 0.3 | 24 | 0.111
US50 | MoRh | 09_25 | 0.3 | 25 | 0.113
US50 | MoRh | 09_25 | 0.3 | 26 | 0.115
US50 | MoRh | 09_25 | 0.3 | 27 | 0.116
US50 | MoRh | 09_25 | 0.3 | 28 | 0.117
US50 | MoRh | 09_25 | 0.3 | 29 | 0.117
US50 | MoRh | 09_25 | 0.3 | 30 | 0.118
US50 | MoRh | 09_25 | 0.3 | 31 | 0.119
US50 | MoRh | 09_25 | 0.3 | 32 | 0.119
US50 | MoRh | 09_25 | 0.3 | 33 | 0.12
US50 | MoRh | 09_25 | 0.3 | 34 | 0.121
US50 | MoRh | 09_25 | 0.3 | 35 | 0.121
US50 | MoRh | 09_25 | 0.3 | 36 | 0.122
US50 | MoRh | 09_25 | 0.3 | 37 | 0.122
US50 | MoRh | 09_25 | 0.3 | 38 | 0.123
US50 | MoRh | 09_25 | 0.3 | 39 | 0.124
US50 | MoRh | 09_25 | 0.3 | 40 | 0.124
US50 | MoRh | 09_25 | 0.3 | 41 | 0.125
US50 | MoRh | 09_25 | 0.3 | 42 | 0.125
US50 | MoRh | 09_25 | 0.315 | 20 | 0.103
US50 | MoRh | 09_25 | 0.315 | 21 | 0.108
US50 | MoRh | 09_25 | 0.315 | 22 | 0.112
US50 | MoRh | 09_25 | 0.315 | 23 | 0.115
US50 | MoRh | 09_25 | 0.315 | 24 | 0.118
US50 | MoRh | 09_25 | 0.315 | 25 | 0.12
US50 | MoRh | 09_25 | 0.315 | 26 | 0.121
US50 | MoRh | 09_25 | 0.315 | 27 | 0.122
US50 | MoRh | 09_25 | 0.315 | 28 | 0.123
US50 | MoRh | 09_25 | 0.315 | 29 | 0.123
US50 | MoRh | 09_25 | 0.315 | 30 | 0.124
US50 | MoRh | 09_25 | 0.315 | 31 | 0.124
US50 | MoRh | 09_25 | 0.315 | 32 | 0.125
US50 | MoRh | 09_25 | 0.315 | 33 | 0.126
US50 | MoRh | 09_25 | 0.315 | 34 | 0.126
US50 | MoRh | 09_25 | 0.315 | 35 | 0.127
US50 | MoRh | 09_25 | 0.315 | 36 | 0.127
US50 | MoRh | 09_25 | 0.315 | 37 | 0.128
US50 | MoRh | 09_25 | 0.315 | 38 | 0.128
US50 | MoRh | 09_25 | 0.315 | 39 | 0.129
US50 | MoRh | 09_25 | 0.315 | 40 | 0.129
US50 | MoRh | 09_25 | 0.315 | 41 | 0.13
US50 | MoRh | 09_25 | 0.315 | 42 | 0.13
US50 | MoRh | 09_25 | 0.33 | 20 | 0.11
US50 | MoRh | 09_25 | 0.33 | 21 | 0.114
US50 | MoRh | 09_25 | 0.33 | 22 | 0.118
US50 | MoRh | 09_25 | 0.33 | 23 | 0.121
US50 | MoRh | 09_25 | 0.33 | 24 | 0.124
US50 | MoRh | 09_25 | 0.33 | 25 | 0.126
US50 | MoRh | 09_25 | 0.33 | 26 | 0.127
US50 | MoRh | 09_25 | 0.33 | 27 | 0.128
US50 | MoRh | 09_25 | 0.33 | 28 | 0.128
US50 | MoRh | 09_25 | 0.33 | 29 | 0.129
US50 | MoRh | 09_25 | 0.33 | 30 | 0.129
US50 | MoRh | 09_25 | 0.33 | 31 | 0.13
US50 | MoRh | 09_25 | 0.33 | 32 | 0.131
US50 | MoRh | 09_25 | 0.33 | 33 | 0.131
US50 | MoRh | 09_25 | 0.33 | 34 | 0.132
US50 | MoRh | 09_25 | 0.33 | 35 | 0.132
US50 | MoRh | 09_25 | 0.33 | 36 | 0.133
US50 | MoRh | 09_25 | 0.33 | 37 | 0.133
US50 | MoRh | 09_25 | 0.33 | 38 | 0.134
US50 | MoRh | 09_25 | 0.33 | 39 | 0.134
US50 | MoRh | 09_25 | 0.33 | 40 | 0.135
US50 | MoRh | 09_25 | 0.33 | 41 | 0.135
US50 | MoRh | 09_25 | 0.33 | 42 | 0.136
US50 | MoRh | 09_25 | 0.345 | 20 | 0.116
US50 | MoRh | 09_25 | 0.345 | 21 | 0.121
US50 | MoRh | 09_25 | 0.345 | 22 | 0.124
US50 | MoRh | 09_25 | 0.345 | 23 | 0.127
US50 | MoRh | 09_25 | 0.345 | 24 | 0.13
US50 | MoRh | 09_25 | 0.345 | 25 | 0.132
US50 | MoRh | 09_25 | 0.345 | 26 | 0.133
US50 | MoRh | 09_25 | 0.345 | 27 | 0.134
US50 | MoRh | 09_25 | 0.345 | 28 | 0.134
US50 | MoRh | 09_25 | 0.345 | 29 | 0.135
US50 | MoRh | 09_25 | 0.345 | 30 | 0.135
US50 | MoRh | 09_25 | 0.345 | 31 | 0.136
US50 | MoRh | 09_25 | 0.345 | 32 | 0.136
US50 | MoRh | 09_25 | 0.345 | 33 | 0.137
US50 | MoRh | 09_25 | 0.345 | 34 | 0.137
US50 | MoRh | 09_25 | 0.345 | 35 | 0.138
US50 | MoRh | 09_25 | 0.345 | 36 | 0.138
US50 | MoRh | 09_25 | 0.345 | 37 | 0.139
US50 | MoRh | 09_25 | 0.345 | 38 | 0.139
US50 | MoRh | 09_25 | 0.345 | 39 | 0.14
US50 | MoRh | 09_25 | 0.345 | 40 | 0.14
US50 | MoRh | 09_25 | 0.345 | 41 | 0.14
US50 | MoRh | 09_25 | 0.345 | 42 | 0.141
US50 | MoRh | 09_25 | 0.345 | 43 | 0.141
US50 | MoRh | 09_25 | 0.345 | 44 | 0.142
US50 | MoRh | 09_25 | 0.345 | 45 | 0.142
US50 | MoRh | 09_25 | 0.345 | 46 | 0.142
US50 | MoRh | 09_25 | 0.36 | 20 | 0.123
US50 | MoRh | 09_25 | 0.36 | 21 | 0.127
US50 | MoRh | 09_25 | 0.36 | 22 | 0.13
US50 | MoRh | 09_25 | 0.36 | 23 | 0.134
US50 | MoRh | 09_25 | 0.36 | 24 | 0.136
US50 | MoRh | 09_25 | 0.36 | 25 | 0.138
US50 | MoRh | 09_25 | 0.36 | 26 | 0.139
US50 | MoRh | 09_25 | 0.36 | 27 | 0.14
US50 | MoRh | 09_25 | 0.36 | 28 | 0.14
US50 | MoRh | 09_25 | 0.36 | 29 | 0.14
US50 | MoRh | 09_25 | 0.36 | 30 | 0.141
US50 | MoRh | 09_25 | 0.36 | 31 | 0.141
US50 | MoRh | 09_25 | 0.36 | 32 | 0.142
US50 | MoRh | 09_25 | 0.36 | 33 | 0.142
US50 | MoRh | 09_25 | 0.36 | 34 | 0.143
US50 | MoRh | 09_25 | 0.36 | 35 | 0.143
US50 | MoRh | 09_25 | 0.36 | 36 | 0.143
US50 | MoRh | 09_25 | 0.36 | 37 | 0.144
US50 | MoRh | 09_25 | 0.36 | 38 | 0.144
US50 | MoRh | 09_25 | 0.36 | 39 | 0.145
US50 | MoRh | 09_25 | 0.36 | 40 | 0.145
US50 | MoRh | 09_25 | 0.36 | 41 | 0.146
US50 | MoRh | 09_25 | 0.36 | 42 | 0.146
US50 | MoRh | 09_25 | 0.36 | 43 | 0.146
US50 | MoRh | 09_25 | 0.36 | 44 | 0.147
US50 | MoRh | 09_25 | 0.36 | 45 | 0.147
US50 | MoRh | 09_25 | 0.36 | 46 | 0.147
US50 | MoRh | 09_25 | 0.36 | 47 | 0.148
US50 | MoRh | 09_25 | 0.36 | 48 | 0.148
US50 | MoRh | 09_25 | 0.36 | 49 | 0.148
US50 | MoRh | 09_25 | 0.375 | 20 | 0.129
US50 | MoRh | 09_25 | 0.375 | 21 | 0.133
US50 | MoRh | 09_25 | 0.375 | 22 | 0.137
US50 | MoRh | 09_25 | 0.375 | 23 | 0.14
US50 | MoRh | 09_25 | 0.375 | 24 | 0.142
US50 | MoRh | 09_25 | 0.375 | 25 | 0.144
US50 | MoRh | 09_25 | 0.375 | 26 | 0.145
US50 | MoRh | 09_25 | 0.375 | 27 | 0.146
US50 | MoRh | 09_25 | 0.375 | 28 | 0.146
US50 | MoRh | 09_25 | 0.375 | 29 | 0.146
US50 | MoRh | 09_25 | 0.375 | 30 | 0.146
US50 | MoRh | 09_25 | 0.375 | 31 | 0.147
US50 | MoRh | 09_25 | 0.375 | 32 | 0.147
US50 | MoRh | 09_25 | 0.375 | 33 | 0.147
US50 | MoRh | 09_25 | 0.375 | 34 | 0.148
US50 | MoRh | 09_25 | 0.375 | 35 | 0.148
US50 | MoRh | 09_25 | 0.375 | 36 | 0.149
US50 | MoRh | 09_25 | 0.375 | 37 | 0.149
US50 | MoRh | 09_25 | 0.375 | 38 | 0.15
US50 | MoRh | 09_25 | 0.375 | 39 | 0.15
US50 | MoRh | 09_25 | 0.375 | 40 | 0.15
US50 | MoRh | 09_25 | 0.375 | 41 | 0.151
US50 | MoRh | 09_25 | 0.375 | 42 | 0.151
US50 | MoRh | 09_25 | 0.375 | 43 | 0.152
US50 | MoRh | 09_25 | 0.375 | 44 | 0.152
US50 | MoRh | 09_25 | 0.375 | 45 | 0.152
US50 | MoRh | 09_25 | 0.375 | 46 | 0.153
US50 | MoRh | 09_25 | 0.375 | 47 | 0.153
US50 | MoRh | 09_25 | 0.375 | 48 | 0.153
US50 | MoRh | 09_25 | 0.375 | 49 | 0.154
US50 | MoRh | 09_25 | 0.39 | 20 | 0.136
US50 | MoRh | 09_25 | 0.39 | 21 | 0.14
US50 | MoRh | 09_25 | 0.39 | 22 | 0.143
US50 | MoRh | 09_25 | 0.39 | 23 | 0.146
US50 | MoRh | 09_25 | 0.39 | 24 | 0.148
US50 | MoRh | 09_25 | 0.39 | 25 | 0.15
US50 | MoRh | 09_25 | 0.39 | 26 | 0.151
US50 | MoRh | 09_25 | 0.39 | 27 | 0.151
US50 | MoRh | 09_25 | 0.39 | 28 | 0.152
US50 | MoRh | 09_25 | 0.39 | 29 | 0.152
US50 | MoRh | 09_25 | 0.39 | 30 | 0.152
US50 | MoRh | 09_25 | 0.39 | 31 | 0.152
US50 | MoRh | 09_25 | 0.39 | 32 | 0.153
US50 | MoRh | 09_25 | 0.39 | 33 | 0.153
US50 | MoRh | 09_25 | 0.39 | 34 | 0.153
US50 | MoRh | 09_25 | 0.39 | 35 | 0.154
US50 | MoRh | 09_25 | 0.39 | 36 | 0.154
US50 | MoRh | 09_25 | 0.39 | 37 | 0.154
US50 | MoRh | 09_25 | 0.39 | 38 | 0.155
US50 | MoRh | 09_25 | 0.39 | 39 | 0.155
US50 | MoRh | 09_25 | 0.39 | 40 | 0.156
US50 | MoRh | 09_25 | 0.39 | 41 | 0.156
US50 | MoRh | 09_25 | 0.39 | 42 | 0.156
US50 | MoRh | 09_25 | 0.39 | 43 | 0.157
US50 | MoRh | 09_25 | 0.39 | 44 | 0.157
US50 | MoRh | 09_25 | 0.39 | 45 | 0.157
US50 | MoRh | 09_25 | 0.39 | 46 | 0.158
US50 | MoRh | 09_25 | 0.39 | 47 | 0.158
US50 | MoRh | 09_25 | 0.39 | 48 | 0.158
US50 | MoRh | 09_25 | 0.39 | 49 | 0.159
US50 | MoRh | 09_25 | 0.405 | 20 | 0.142
US50 | MoRh | 09_25 | 0.405 | 21 | 0.146
US50 | MoRh | 09_25 | 0.405 | 22 | 0.149
US50 | MoRh | 09_25 | 0.405 | 23 | 0.152
US50 | MoRh | 09_25 | 0.405 | 24 | 0.155
US50 | MoRh | 09_25 | 0.405 | 25 | 0.156
US50 | MoRh | 09_25 | 0.405 | 26 | 0.157
US50 | MoRh | 09_25 | 0.405 | 27 | 0.157
US50 | MoRh | 09_25 | 0.405 | 28 | 0.158
US50 | MoRh | 09_25 | 0.405 | 29 | 0.158
US50 | MoRh | 09_25 | 0.405 | 30 | 0.158
US50 | MoRh | 09_25 | 0.405 | 31 | 0.158
US50 | MoRh | 09_25 | 0.405 | 32 | 0.158
US50 | MoRh | 09_25 | 0.405 | 33 | 0.158
US50 | MoRh | 09_25 | 0.405 | 34 | 0.159
US50 | MoRh | 09_25 | 0.405 | 35 | 0.159
US50 | MoRh | 09_25 | 0.405 | 36 | 0.159
US50 | MoRh | 09_25 | 0.405 | 37 | 0.16
US50 | MoRh | 09_25 | 0.405 | 38 | 0.16
US50 | MoRh | 09_25 | 0.405 | 39 | 0.16
US50 | MoRh | 09_25 | 0.405 | 40 | 0.161
US50 | MoRh | 09_25 | 0.405 | 41 | 0.161
US50 | MoRh | 09_25 | 0.405 | 42 | 0.161
US50 | MoRh | 09_25 | 0.405 | 43 | 0.162
US50 | MoRh | 09_25 | 0.405 | 44 | 0.162
US50 | MoRh | 09_25 | 0.405 | 45 | 0.162
US50 | MoRh | 09_25 | 0.405 | 46 | 0.163
US50 | MoRh | 09_25 | 0.405 | 47 | 0.163
US50 | MoRh | 09_25 | 0.405 | 48 | 0.163
US50 | MoRh | 09_25 | 0.405 | 49 | 0.164
US50 | MoRh | 09_25 | 0.42 | 23 | 0.158
US50 | MoRh | 09_25 | 0.42 | 24 | 0.161
US50 | MoRh | 09_25 | 0.42 | 25 | 0.162
US50 | MoRh | 09_25 | 0.42 | 26 | 0.163
US50 | MoRh | 09_25 | 0.42 | 27 | 0.163
US50 | MoRh | 09_25 | 0.42 | 28 | 0.163
US50 | MoRh | 09_25 | 0.42 | 29 | 0.164
US50 | MoRh | 09_25 | 0.42 | 30 | 0.164
US50 | MoRh | 09_25 | 0.42 | 31 | 0.164
US50 | MoRh | 09_25 | 0.42 | 32 | 0.164
US50 | MoRh | 09_25 | 0.42 | 33 | 0.164
US50 | MoRh | 09_25 | 0.42 | 34 | 0.164
US50 | MoRh | 09_25 | 0.42 | 35 | 0.164
US50 | MoRh | 09_25 | 0.42 | 36 | 0.165
US50 | MoRh | 09_25 | 0.42 | 37 | 0.165
US50 | MoRh | 09_25 | 0.42 | 38 | 0.165
US50 | MoRh | 09_25 | 0.42 | 39 | 0.166
US50 | MoRh | 09_25 | 0.42 | 40 | 0.166
US50 | MoRh | 09_25 | 0.42 | 41 | 0.166
US50 | MoRh | 09_25 | 0.42 | 42 | 0.167
US50 | MoRh | 09_25 | 0.42 | 43 | 0.167
US50 | MoRh | 09_25 | 0.42 | 44 | 0.167
US50 | MoRh | 09_25 | 0.42 | 45 | 0.168
US50 | MoRh | 09_25 | 0.42 | 46 | 0.168
US50 | MoRh | 09_25 | 0.42 | 47 | 0.168
US50 | MoRh | 09_25 | 0.42 | 48 | 0.169
US50 | MoRh | 09_25 | 0.42 | 49 | 0.169
US50 | MoRh | 09_25 | 0.435 | 23 | 0.164
US50 | MoRh | 09_25 | 0.435 | 24 | 0.167
US50 | MoRh | 09_25 | 0.435 | 25 | 0.168
US50 | MoRh | 09_25 | 0.435 | 26 | 0.169
US50 | MoRh | 09_25 | 0.435 | 27 | 0.169
US50 | MoRh | 09_25 | 0.435 | 28 | 0.169
US50 | MoRh | 09_25 | 0.435 | 29 | 0.169
US50 | MoRh | 09_25 | 0.435 | 30 | 0.169
US50 | MoRh | 09_25 | 0.435 | 31 | 0.169
US50 | MoRh | 09_25 | 0.435 | 32 | 0.169
US50 | MoRh | 09_25 | 0.435 | 33 | 0.17
US50 | MoRh | 09_25 | 0.435 | 34 | 0.17
US50 | MoRh | 09_25 | 0.435 | 35 | 0.17
US50 | MoRh | 09_25 | 0.435 | 36 | 0.17
US50 | MoRh | 09_25 | 0.435 | 37 | 0.17
US50 | MoRh | 09_25 | 0.435 | 38 | 0.171
US50 | MoRh | 09_25 | 0.435 | 39 | 0.171
US50 | MoRh | 09_25 | 0.435 | 40 | 0.171
US50 | MoRh | 09_25 | 0.435 | 41 | 0.172
US50 | MoRh | 09_25 | 0.435 | 42 | 0.172
US50 | MoRh | 09_25 | 0.435 | 43 | 0.172
US50 | MoRh | 09_25 | 0.435 | 44 | 0.173
US50 | MoRh | 09_25 | 0.435 | 45 | 0.173
US50 | MoRh | 09_25 | 0.435 | 46 | 0.173
US50 | MoRh | 09_25 | 0.435 | 47 | 0.173
US50 | MoRh | 09_25 | 0.435 | 48 | 0.174
US50 | MoRh | 09_25 | 0.435 | 49 | 0.174
US50 | MoRh | 09_25 | 0.45 | 23 | 0.171
US50 | MoRh | 09_25 | 0.45 | 24 | 0.173
US50 | MoRh | 09_25 | 0.45 | 25 | 0.174
US50 | MoRh | 09_25 | 0.45 | 26 | 0.175
US50 | MoRh | 09_25 | 0.45 | 27 | 0.175
US50 | MoRh | 09_25 | 0.45 | 28 | 0.175
US50 | MoRh | 09_25 | 0.45 | 29 | 0.175
US50 | MoRh | 09_25 | 0.45 | 30 | 0.175
US50 | MoRh | 09_25 | 0.45 | 31 | 0.175
US50 | MoRh | 09_25 | 0.45 | 32 | 0.175
US50 | MoRh | 09_25 | 0.45 | 33 | 0.175
US50 | MoRh | 09_25 | 0.45 | 34 | 0.175
US50 | MoRh | 09_25 | 0.45 | 35 | 0.176
US50 | MoRh | 09_25 | 0.45 | 36 | 0.176
US50 | MoRh | 09_25 | 0.45 | 37 | 0.176
US50 | MoRh | 09_25 | 0.45 | 38 | 0.176
US50 | MoRh | 09_25 | 0.45 | 39 | 0.176
US50 | MoRh | 09_25 | 0.45 | 40 | 0.177
US50 | MoRh | 09_25 | 0.45 | 41 | 0.177
US50 | MoRh | 09_25 | 0.45 | 42 | 0.177
US50 | MoRh | 09_25 | 0.45 | 43 | 0.178
US50 | MoRh | 09_25 | 0.45 | 44 | 0.178
US50 | MoRh | 09_25 | 0.45 | 45 | 0.178
US50 | MoRh | 09_25 | 0.45 | 46 | 0.178
US50 | MoRh | 09_25 | 0.45 | 47 | 0.179
US50 | MoRh | 09_25 | 0.45 | 48 | 0.179
US50 | MoRh | 09_25 | 0.45 | 49 | 0.179
US50 | MoRh | 09_25 | 0.465 | 23 | 0.177
US50 | MoRh | 09_25 | 0.465 | 24 | 0.179
US50 | MoRh | 09_25 | 0.465 | 25 | 0.181
US50 | MoRh | 09_25 | 0.465 | 26 | 0.181
US50 | MoRh | 09_25 | 0.465 | 27 | 0.181
US50 | MoRh | 09_25 | 0.465 | 28 | 0.181
US50 | MoRh | 09_25 | 0.465 | 29 | 0.181
US50 | MoRh | 09_25 | 0.465 | 30 | 0.181
US50 | MoRh | 09_25 | 0.465 | 31 | 0.181
US50 | MoRh | 09_25 | 0.465 | 32 | 0.181
US50 | MoRh | 09_25 | 0.465 | 33 | 0.181
US50 | MoRh | 09_25 | 0.465 | 34 | 0.181
US50 | MoRh | 09_25 | 0.465 | 35 | 0.181
US50 | MoRh | 09_25 | 0.465 | 36 | 0.181
US50 | MoRh | 09_25 | 0.465 | 37 | 0.182
US50 | MoRh | 09_25 | 0.465 | 38 | 0.182
US50 | MoRh | 09_25 | 0.465 | 39 | 0.182
US50 | MoRh | 09_25 | 0.465 | 40 | 0.182
US50 | MoRh | 09_25 | 0.465 | 41 | 0.182
US50 | MoRh | 09_25 | 0.465 | 42 | 0.183
US50 | MoRh | 09_25 | 0.465 | 43 | 0.183
US50 | MoRh | 09_25 | 0.465 | 44 | 0.183
US50 | MoRh | 09_25 | 0.465 | 45 | 0.184
US50 | MoRh | 09_25 | 0.465 | 46 | 0.184
US50 | MoRh | 09_25 | 0.465 | 47 | 0.184
US50 | MoRh | 09_25 | 0.465 | 48 | 0.185
US50 | MoRh | 09_25 | 0.465 | 49 | 0.185
US50 | MoRh | 09_25 | 0.48 | 25 | 0.187
US50 | MoRh | 09_25 | 0.48 | 26 | 0.187
US50 | MoRh | 09_25 | 0.48 | 27 | 0.187
US50 | MoRh | 09_25 | 0.48 | 28 | 0.187
US50 | MoRh | 09_25 | 0.48 | 29 | 0.187
US50 | MoRh | 09_25 | 0.48 | 30 | 0.187
US50 | MoRh | 09_25 | 0.48 | 31 | 0.187
US50 | MoRh | 09_25 | 0.48 | 32 | 0.187
US50 | MoRh | 09_25 | 0.48 | 33 | 0.187
US50 | MoRh | 09_25 | 0.48 | 34 | 0.187
US50 | MoRh | 09_25 | 0.48 | 35 | 0.187
US50 | MoRh | 09_25 | 0.48 | 36 | 0.187
US50 | MoRh | 09_25 | 0.48 | 37 | 0.187
US50 | MoRh | 09_25 | 0.48 | 38 | 0.187
US50 | MoRh | 09_25 | 0.48 | 39 | 0.188
US50 | MoRh | 09_25 | 0.48 | 40 | 0.188
US50 | MoRh | 09_25 | 0.48 | 41 | 0.188
US50 | MoRh | 09_25 | 0.48 | 42 | 0.188
US50 | MoRh | 09_25 | 0.48 | 43 | 0.189
US50 | MoRh | 09_25 | 0.48 | 44 | 0.189
US50 | MoRh | 09_25 | 0.48 | 45 | 0.189
US50 | MoRh | 09_25 | 0.48 | 46 | 0.19
US50 | MoRh | 09_25 | 0.48 | 47 | 0.19
US50 | MoRh | 09_25 | 0.48 | 48 | 0.19
US50 | MoRh | 09_25 | 0.48 | 49 | 0.19
US50 | MoRh | 09_25 | 0.495 | 28 | 0.194
US50 | MoRh | 09_25 | 0.495 | 29 | 0.193
US50 | MoRh | 09_25 | 0.495 | 30 | 0.193
US50 | MoRh | 09_25 | 0.495 | 31 | 0.193
US50 | MoRh | 09_25 | 0.495 | 32 | 0.193
US50 | MoRh | 09_25 | 0.495 | 33 | 0.193
US50 | MoRh | 09_25 | 0.495 | 34 | 0.193
US50 | MoRh | 09_25 | 0.495 | 35 | 0.193
US50 | MoRh | 09_25 | 0.495 | 36 | 0.193
US50 | MoRh | 09_25 | 0.495 | 37 | 0.193
US50 | MoRh | 09_25 | 0.495 | 38 | 0.193
US50 | MoRh | 09_25 | 0.495 | 39 | 0.193
US50 | MoRh | 09_25 | 0.495 | 40 | 0.194
US50 | MoRh | 09_25 | 0.495 | 41 | 0.194
US50 | MoRh | 09_25 | 0.495 | 42 | 0.194
US50 | MoRh | 09_25 | 0.495 | 43 | 0.195
US50 | MoRh | 09_25 | 0.495 | 44 | 0.195
US50 | MoRh | 09_25 | 0.495 | 45 | 0.195
US50 | MoRh | 09_25 | 0.495 | 46 | 0.195
US50 | MoRh | 09_25 | 0.495 | 47 | 0.196
US50 | MoRh | 09_25 | 0.495 | 48 | 0.196
US50 | MoRh | 09_25 | 0.495 | 49 | 0.196
US50 | MoRh | 09_25 | 0.51 | 28 | 0.2
US50 | MoRh | 09_25 | 0.51 | 29 | 0.2
US50 | MoRh | 09_25 | 0.51 | 30 | 0.199
US50 | MoRh | 09_25 | 0.51 | 31 | 0.199
US50 | MoRh | 09_25 | 0.51 | 32 | 0.199
US50 | MoRh | 09_25 | 0.51 | 33 | 0.199
US50 | MoRh | 09_25 | 0.51 | 34 | 0.199
US50 | MoRh | 09_25 | 0.51 | 35 | 0.199
US50 | MoRh | 09_25 | 0.51 | 36 | 0.199
US50 | MoRh | 09_25 | 0.51 | 37 | 0.199
US50 | MoRh | 09_25 | 0.51 | 38 | 0.199
US50 | MoRh | 09_25 | 0.51 | 39 | 0.2
US50 | MoRh | 09_25 | 0.51 | 40 | 0.2
US50 | MoRh | 09_25 | 0.51 | 41 | 0.2
US50 | MoRh | 09_25 | 0.51 | 42 | 0.2
US50 | MoRh | 09_25 | 0.51 | 43 | 0.201
US50 | MoRh | 09_25 | 0.51 | 44 | 0.201
US50 | MoRh | 09_25 | 0.51 | 45 | 0.201
US50 | MoRh | 09_25 | 0.51 | 46 | 0.201
US50 | MoRh | 09_25 | 0.51 | 47 | 0.202
US50 | MoRh | 09_25 | 0.51 | 48 | 0.202
US50 | MoRh | 09_25 | 0.51 | 49 | 0.202
US50 | MoRh | 09_25 | 0.525 | 29 | 0.206
US50 | MoRh | 09_25 | 0.525 | 30 | 0.206
US50 | MoRh | 09_25 | 0.525 | 31 | 0.205
US50 | MoRh | 09_25 | 0.525 | 32 | 0.205
US50 | MoRh | 09_25 | 0.525 | 33 | 0.205
US50 | MoRh | 09_25 | 0.525 | 34 | 0.205
US50 | MoRh | 09_25 | 0.525 | 35 | 0.205
US50 | MoRh | 09_25 | 0.525 | 36 | 0.205
US50 | MoRh | 09_25 | 0.525 | 37 | 0.205
US50 | MoRh | 09_25 | 0.525 | 38 | 0.206
US50 | MoRh | 09_25 | 0.525 | 39 | 0.206
US50 | MoRh | 09_25 | 0.525 | 40 | 0.206
US50 | MoRh | 09_25 | 0.525 | 41 | 0.206
US50 | MoRh | 09_25 | 0.525 | 42 | 0.206
US50 | MoRh | 09_25 | 0.525 | 43 | 0.207
US50 | MoRh | 09_25 | 0.525 | 44 | 0.207
US50 | MoRh | 09_25 | 0.525 | 45 | 0.207
US50 | MoRh | 09_25 | 0.525 | 46 | 0.208
US50 | MoRh | 09_25 | 0.525 | 47 | 0.208
US50 | MoRh | 09_25 | 0.525 | 48 | 0.208
US50 | MoRh | 09_25 | 0.525 | 49 | 0.209
US50 | MoRh | 09_25 | 0.54 | 32 | 0.212
US50 | MoRh | 09_25 | 0.54 | 33 | 0.212
US50 | MoRh | 09_25 | 0.54 | 34 | 0.212
US50 | MoRh | 09_25 | 0.54 | 35 | 0.212
US50 | MoRh | 09_25 | 0.54 | 36 | 0.212
US50 | MoRh | 09_25 | 0.54 | 37 | 0.212
US50 | MoRh | 09_25 | 0.54 | 38 | 0.212
US50 | MoRh | 09_25 | 0.54 | 39 | 0.212
US50 | MoRh | 09_25 | 0.54 | 40 | 0.212
US50 | MoRh | 09_25 | 0.54 | 41 | 0.212
US50 | MoRh | 09_25 | 0.54 | 42 | 0.213
US50 | MoRh | 09_25 | 0.54 | 43 | 0.213
US50 | MoRh | 09_25 | 0.54 | 44 | 0.213
US50 | MoRh | 09_25 | 0.54 | 45 | 0.214
US50 | MoRh | 09_25 | 0.54 | 46 | 0.214
US50 | MoRh | 09_25 | 0.54 | 47 | 0.214
US50 | MoRh | 09_25 | 0.54 | 48 | 0.215
US50 | MoRh | 09_25 | 0.54 | 49 | 0.215
US50 | MoRh | 09_25 | 0.555 | 34 | 0.218
US50 | MoRh | 09_25 | 0.555 | 35 | 0.218
US50 | MoRh | 09_25 | 0.555 | 36 | 0.218
US50 | MoRh | 09_25 | 0.555 | 37 | 0.218
US50 | MoRh | 09_25 | 0.555 | 38 | 0.218
US50 | MoRh | 09_25 | 0.555 | 39 | 0.218
US50 | MoRh | 09_25 | 0.555 | 40 | 0.218
US50 | MoRh | 09_25 | 0.555 | 41 | 0.219
US50 | MoRh | 09_25 | 0.555 | 42 | 0.219
US50 | MoRh | 09_25 | 0.555 | 43 | 0.219
US50 | MoRh | 09_25 | 0.555 | 44 | 0.22
US50 | MoRh | 09_25 | 0.555 | 45 | 0.22
US50 | MoRh | 09_25 | 0.555 | 46 | 0.22
US50 | MoRh | 09_25 | 0.555 | 47 | 0.221
US50 | MoRh | 09_25 | 0.555 | 48 | 0.221
US50 | MoRh | 09_25 | 0.555 | 49 | 0.222
US50 | MoRh | 09_25 | 0.57 | 43 | 0.226
US50 | MoRh | 09_25 | 0.57 | 44 | 0.226
US50 | MoRh | 09_25 | 0.57 | 45 | 0.226
US50 | MoRh | 09_25 | 0.57 | 46 | 0.227
US50 | MoRh | 09_25 | 0.57 | 47 | 0.227
US50 | MoRh | 09_25 | 0.57 | 48 | 0.228
US50 | MoRh | 09_25 | 0.57 | 49 | 0.228
US50 | MoRh | 09_25 | 0.585 | 43 | 0.232
US50 | MoRh | 09_25 | 0.585 | 44 | 0.232
US50 | MoRh | 09_25 | 0.585 | 45 | 0.233
US50 | MoRh | 09_25 | 0.585 | 46 | 0.233
US50 | MoRh | 09_25 | 0.585 | 47 | 0.233
US50 | MoRh | 09_25 | 0.585 | 48 | 0.234
US50 | MoRh | 09_25 | 0.585 | 49 | 0.234
US50 | MoRh | 09_25 | 0.6 | 43 | 0.237
US50 | MoRh | 09_25 | 0.6 | 44 | 0.238
US50 | MoRh | 09_25 | 0.6 | 45 | 0.239
US50 | MoRh | 09_25 | 0.6 | 46 | 0.239
US50 | MoRh | 09_25 | 0.6 | 47 | 0.24
US50 | MoRh | 09_25 | 0.6 | 48 | 0.24
US50 | MoRh | 09_25 | 0.6 | 49 | 0.241
US50 | RhAg | 09_25 | 0.2 | 20 | 0.054
US50 | RhAg | 09_25 | 0.2 | 21 | 0.057
US50 | RhAg | 09_25 | 0.2 | 22 | 0.06
US50 | RhAg | 09_25 | 0.2 | 23 | 0.062
US50 | RhAg | 09_25 | 0.2 | 24 | 0.066
US50 | RhAg | 09_25 | 0.2 | 25 | 0.069
US50 | RhAg | 09_25 | 0.22 | 20 | 0.063
US50 | RhAg | 09_25 | 0.22 | 21 | 0.067
US50 | RhAg | 09_25 | 0.22 | 22 | 0.069
US50 | RhAg | 09_25 | 0.22 | 23 | 0.072
US50 | RhAg | 09_25 | 0.22 | 24 | 0.076
US50 | RhAg | 09_25 | 0.22 | 25 | 0.079
US50 | RhAg | 09_25 | 0.22 | 26 | 0.082
US50 | RhAg | 09_25 | 0.22 | 27 | 0.085
US50 | RhAg | 09_25 | 0.24 | 20 | 0.072
US50 | RhAg | 09_25 | 0.24 | 21 | 0.075
US50 | RhAg | 09_25 | 0.24 | 22 | 0.078
US50 | RhAg | 09_25 | 0.24 | 23 | 0.081
US50 | RhAg | 09_25 | 0.24 | 24 | 0.085
US50 | RhAg | 09_25 | 0.24 | 25 | 0.089
US50 | RhAg | 09_25 | 0.24 | 26 | 0.092
US50 | RhAg | 09_25 | 0.24 | 27 | 0.095
US50 | RhAg | 09_25 | 0.24 | 28 | 0.098
US50 | RhAg | 09_25 | 0.24 | 29 | 0.1
US50 | RhAg | 09_25 | 0.26 | 20 | 0.08
US50 | RhAg | 09_25 | 0.26 | 21 | 0.084
US50 | RhAg | 09_25 | 0.26 | 22 | 0.087
US50 | RhAg | 09_25 | 0.26 | 23 | 0.091
US50 | RhAg | 09_25 | 0.26 | 24 | 0.095
US50 | RhAg | 09_25 | 0.26 | 25 | 0.099
US50 | RhAg | 09_25 | 0.26 | 26 | 0.102
US50 | RhAg | 09_25 | 0.26 | 27 | 0.105
US50 | RhAg | 09_25 | 0.26 | 28 | 0.108
US50 | RhAg | 09_25 | 0.26 | 29 | 0.11
US50 | RhAg | 09_25 | 0.26 | 30 | 0.112
US50 | RhAg | 09_25 | 0.26 | 31 | 0.113
US50 | RhAg | 09_25 | 0.26 | 32 | 0.115
US50 | RhAg | 09_25 | 0.28 | 20 | 0.088
US50 | RhAg | 09_25 | 0.28 | 21 | 0.093
US50 | RhAg | 09_25 | 0.28 | 22 | 0.096
US50 | RhAg | 09_25 | 0.28 | 23 | 0.1
US50 | RhAg | 09_25 | 0.28 | 24 | 0.104
US50 | RhAg | 09_25 | 0.28 | 25 | 0.108
US50 | RhAg | 09_25 | 0.28 | 26 | 0.112
US50 | RhAg | 09_25 | 0.28 | 27 | 0.115
US50 | RhAg | 09_25 | 0.28 | 28 | 0.117
US50 | RhAg | 09_25 | 0.28 | 29 | 0.119
US50 | RhAg | 09_25 | 0.28 | 30 | 0.121
US50 | RhAg | 09_25 | 0.28 | 31 | 0.123
US50 | RhAg | 09_25 | 0.28 | 32 | 0.125
US50 | RhAg | 09_25 | 0.28 | 33 | 0.126
US50 | RhAg | 09_25 | 0.28 | 34 | 0.127
US50 | RhAg | 09_25 | 0.3 | 20 | 0.097
US50 | RhAg | 09_25 | 0.3 | 21 | 0.102
US50 | RhAg | 09_25 | 0.3 | 22 | 0.105
US50 | RhAg | 09_25 | 0.3 | 23 | 0.108
US50 | RhAg | 09_25 | 0.3 | 24 | 0.113
US50 | RhAg | 09_25 | 0.3 | 25 | 0.117
US50 | RhAg | 09_25 | 0.3 | 26 | 0.121
US50 | RhAg | 09_25 | 0.3 | 27 | 0.124
US50 | RhAg | 09_25 | 0.3 | 28 | 0.126
US50 | RhAg | 09_25 | 0.3 | 29 | 0.128
US50 | RhAg | 09_25 | 0.3 | 30 | 0.13
US50 | RhAg | 09_25 | 0.3 | 31 | 0.132
US50 | RhAg | 09_25 | 0.3 | 32 | 0.133
US50 | RhAg | 09_25 | 0.3 | 33 | 0.135
US50 | RhAg | 09_25 | 0.3 | 34 | 0.136
US50 | RhAg | 09_25 | 0.3 | 35 | 0.137
US50 | RhAg | 09_25 | 0.3 | 36 | 0.139
US50 | RhAg | 09_25 | 0.3 | 37 | 0.14
US50 | RhAg | 09_25 | 0.32 | 20 | 0.105
US50 | RhAg | 09_25 | 0.32 | 21 | 0.11
US50 | RhAg | 09_25 | 0.32 | 22 | 0.113
US50 | RhAg | 09_25 | 0.32 | 23 | 0.117
US50 | RhAg | 09_25 | 0.32 | 24 | 0.122
US50 | RhAg | 09_25 | 0.32 | 25 | 0.126
US50 | RhAg | 09_25 | 0.32 | 26 | 0.13
US50 | RhAg | 09_25 | 0.32 | 27 | 0.133
US50 | RhAg | 09_25 | 0.32 | 28 | 0.135
US50 | RhAg | 09_25 | 0.32 | 29 | 0.137
US50 | RhAg | 09_25 | 0.32 | 30 | 0.139
US50 | RhAg | 09_25 | 0.32 | 31 | 0.141
US50 | RhAg | 09_25 | 0.32 | 32 | 0.142
US50 | RhAg | 09_25 | 0.32 | 33 | 0.143
US50 | RhAg | 09_25 | 0.32 | 34 | 0.145
US50 | RhAg | 09_25 | 0.32 | 35 | 0.146
US50 | RhAg | 09_25 | 0.32 | 36 | 0.147
US50 | RhAg | 09_25 | 0.32 | 37 | 0.148
US50 | RhAg | 09_25 | 0.32 | 38 | 0.149
US50 | RhAg | 09_25 | 0.32 | 39 | 0.15
US50 | RhAg | 09_25 | 0.32 | 40 | 0.151
US50 | RhAg | 09_25 | 0.32 | 41 | 0.152
US50 | RhAg | 09_25 | 0.32 | 42 | 0.153
US50 | RhAg | 09_25 | 0.32 | 43 | 0.154
US50 | RhAg | 09_25 | 0.34 | 20 | 0.114
US50 | RhAg | 09_25 | 0.34 | 21 | 0.119
US50 | RhAg | 09_25 | 0.34 | 22 | 0.122
US50 | RhAg | 09_25 | 0.34 | 23 | 0.126
US50 | RhAg | 09_25 | 0.34 | 24 | 0.131
US50 | RhAg | 09_25 | 0.34 | 25 | 0.135
US50 | RhAg | 09_25 | 0.34 | 26 | 0.138
US50 | RhAg | 09_25 | 0.34 | 27 | 0.141
US50 | RhAg | 09_25 | 0.34 | 28 | 0.144
US50 | RhAg | 09_25 | 0.34 | 29 | 0.146
US50 | RhAg | 09_25 | 0.34 | 30 | 0.147
US50 | RhAg | 09_25 | 0.34 | 31 | 0.149
US50 | RhAg | 09_25 | 0.34 | 32 | 0.15
US50 | RhAg | 09_25 | 0.34 | 33 | 0.152
US50 | RhAg | 09_25 | 0.34 | 34 | 0.153
US50 | RhAg | 09_25 | 0.34 | 35 | 0.154
US50 | RhAg | 09_25 | 0.34 | 36 | 0.155
US50 | RhAg | 09_25 | 0.34 | 37 | 0.156
US50 | RhAg | 09_25 | 0.34 | 38 | 0.157
US50 | RhAg | 09_25 | 0.34 | 39 | 0.158
US50 | RhAg | 09_25 | 0.34 | 40 | 0.159
US50 | RhAg | 09_25 | 0.34 | 41 | 0.16
US50 | RhAg | 09_25 | 0.34 | 42 | 0.161
US50 | RhAg | 09_25 | 0.34 | 43 | 0.161
US50 | RhAg | 09_25 | 0.34 | 44 | 0.162
US50 | RhAg | 09_25 | 0.34 | 45 | 0.163
US50 | RhAg | 09_25 | 0.34 | 46 | 0.163
US50 | RhAg | 09_25 | 0.36 | 20 | 0.123
US50 | RhAg | 09_25 | 0.36 | 21 | 0.127
US50 | RhAg | 09_25 | 0.36 | 22 | 0.131
US50 | RhAg | 09_25 | 0.36 | 23 | 0.135
US50 | RhAg | 09_25 | 0.36 | 24 | 0.139
US50 | RhAg | 09_25 | 0.36 | 25 | 0.143
US50 | RhAg | 09_25 | 0.36 | 26 | 0.147
US50 | RhAg | 09_25 | 0.36 | 27 | 0.15
US50 | RhAg | 09_25 | 0.36 | 28 | 0.152
US50 | RhAg | 09_25 | 0.36 | 29 | 0.154
US50 | RhAg | 09_25 | 0.36 | 30 | 0.156
US50 | RhAg | 09_25 | 0.36 | 31 | 0.157
US50 | RhAg | 09_25 | 0.36 | 32 | 0.158
US50 | RhAg | 09_25 | 0.36 | 33 | 0.16
US50 | RhAg | 09_25 | 0.36 | 34 | 0.161
US50 | RhAg | 09_25 | 0.36 | 35 | 0.162
US50 | RhAg | 09_25 | 0.36 | 36 | 0.163
US50 | RhAg | 09_25 | 0.36 | 37 | 0.164
US50 | RhAg | 09_25 | 0.36 | 38 | 0.165
US50 | RhAg | 09_25 | 0.36 | 39 | 0.166
US50 | RhAg | 09_25 | 0.36 | 40 | 0.167
US50 | RhAg | 09_25 | 0.36 | 41 | 0.167
US50 | RhAg | 09_25 | 0.36 | 42 | 0.168
US50 | RhAg | 09_25 | 0.36 | 43 | 0.169
US50 | RhAg | 09_25 | 0.36 | 44 | 0.17
US50 | RhAg | 09_25 | 0.36 | 45 | 0.17
US50 | RhAg | 09_25 | 0.36 | 46 | 0.171
US50 | RhAg | 09_25 | 0.36 | 47 | 0.172
US50 | RhAg | 09_25 | 0.36 | 48 | 0.172
US50 | RhAg | 09_25 | 0.36 | 49 | 0.173
US50 | RhAg | 09_25 | 0.38 | 20 | 0.131
US50 | RhAg | 09_25 | 0.38 | 21 | 0.136
US50 | RhAg | 09_25 | 0.38 | 22 | 0.139
US50 | RhAg | 09_25 | 0.38 | 23 | 0.143
US50 | RhAg | 09_25 | 0.38 | 24 | 0.148
US50 | RhAg | 09_25 | 0.38 | 25 | 0.152
US50 | RhAg | 09_25 | 0.38 | 26 | 0.155
US50 | RhAg | 09_25 | 0.38 | 27 | 0.158
US50 | RhAg | 09_25 | 0.38 | 28 | 0.16
US50 | RhAg | 09_25 | 0.38 | 29 | 0.162
US50 | RhAg | 09_25 | 0.38 | 30 | 0.163
US50 | RhAg | 09_25 | 0.38 | 31 | 0.165
US50 | RhAg | 09_25 | 0.38 | 32 | 0.166
US50 | RhAg | 09_25 | 0.38 | 33 | 0.167
US50 | RhAg | 09_25 | 0.38 | 34 | 0.168
US50 | RhAg | 09_25 | 0.38 | 35 | 0.169
US50 | RhAg | 09_25 | 0.38 | 36 | 0.17
US50 | RhAg | 09_25 | 0.38 | 37 | 0.171
US50 | RhAg | 09_25 | 0.38 | 38 | 0.172
US50 | RhAg | 09_25 | 0.38 | 39 | 0.173
US50 | RhAg | 09_25 | 0.38 | 40 | 0.174
US50 | RhAg | 09_25 | 0.38 | 41 | 0.175
US50 | RhAg | 09_25 | 0.38 | 42 | 0.175
US50 | RhAg | 09_25 | 0.38 | 43 | 0.176
US50 | RhAg | 09_25 | 0.38 | 44 | 0.177
US50 | RhAg | 09_25 | 0.38 | 45 | 0.177
US50 | RhAg | 09_25 | 0.38 | 46 | 0.178
US50 | RhAg | 09_25 | 0.38 | 47 | 0.179
US50 | RhAg | 09_25 | 0.38 | 48 | 0.179
US50 | RhAg | 09_25 | 0.38 | 49 | 0.18
US50 | RhAg | 09_25 | 0.4 | 20 | 0.14
US50 | RhAg | 09_25 | 0.4 | 21 | 0.144
US50 | RhAg | 09_25 | 0.4 | 22 | 0.148
US50 | RhAg | 09_25 | 0.4 | 23 | 0.152
US50 | RhAg | 09_25 | 0.4 | 24 | 0.156
US50 | RhAg | 09_25 | 0.4 | 25 | 0.16
US50 | RhAg | 09_25 | 0.4 | 26 | 0.163
US50 | RhAg | 09_25 | 0.4 | 27 | 0.166
US50 | RhAg | 09_25 | 0.4 | 28 | 0.168
US50 | RhAg | 09_25 | 0.4 | 29 | 0.17
US50 | RhAg | 09_25 | 0.4 | 30 | 0.171
US50 | RhAg | 09_25 | 0.4 | 31 | 0.173
US50 | RhAg | 09_25 | 0.4 | 32 | 0.174
US50 | RhAg | 09_25 | 0.4 | 33 | 0.175
US50 | RhAg | 09_25 | 0.4 | 34 | 0.176
US50 | RhAg | 09_25 | 0.4 | 35 | 0.177
US50 | RhAg | 09_25 | 0.4 | 36 | 0.178
US50 | RhAg | 09_25 | 0.4 | 37 | 0.179
US50 | RhAg | 09_25 | 0.4 | 38 | 0.179
US50 | RhAg | 09_25 | 0.4 | 39 | 0.18
US50 | RhAg | 09_25 | 0.4 | 40 | 0.181
US50 | RhAg | 09_25 | 0.4 | 41 | 0.182
US50 | RhAg | 09_25 | 0.4 | 42 | 0.183
US50 | RhAg | 09_25 | 0.4 | 43 | 0.183
US50 | RhAg | 09_25 | 0.4 | 44 | 0.184
US50 | RhAg | 09_25 | 0.4 | 45 | 0.185
US50 | RhAg | 09_25 | 0.4 | 46 | 0.185
US50 | RhAg | 09_25 | 0.4 | 47 | 0.186
US50 | RhAg | 09_25 | 0.4 | 48 | 0.186
US50 | RhAg | 09_25 | 0.4 | 49 | 0.187
US50 | RhAg | 09_25 | 0.42 | 21 | 0.153
US50 | RhAg | 09_25 | 0.42 | 22 | 0.156
US50 | RhAg | 09_25 | 0.42 | 23 | 0.16
US50 | RhAg | 09_25 | 0.42 | 24 | 0.165
US50 | RhAg | 09_25 | 0.42 | 25 | 0.168
US50 | RhAg | 09_25 | 0.42 | 26 | 0.172
US50 | RhAg | 09_25 | 0.42 | 27 | 0.174
US50 | RhAg | 09_25 | 0.42 | 28 | 0.176
US50 | RhAg | 09_25 | 0.42 | 29 | 0.178
US50 | RhAg | 09_25 | 0.42 | 30 | 0.179
US50 | RhAg | 09_25 | 0.42 | 31 | 0.18
US50 | RhAg | 09_25 | 0.42 | 32 | 0.181
US50 | RhAg | 09_25 | 0.42 | 33 | 0.182
US50 | RhAg | 09_25 | 0.42 | 34 | 0.183
US50 | RhAg | 09_25 | 0.42 | 35 | 0.184
US50 | RhAg | 09_25 | 0.42 | 36 | 0.185
US50 | RhAg | 09_25 | 0.42 | 37 | 0.186
US50 | RhAg | 09_25 | 0.42 | 38 | 0.186
US50 | RhAg | 09_25 | 0.42 | 39 | 0.187
US50 | RhAg | 09_25 | 0.42 | 40 | 0.188
US50 | RhAg | 09_25 | 0.42 | 41 | 0.189
US50 | RhAg | 09_25 | 0.42 | 42 | 0.189
US50 | RhAg | 09_25 | 0.42 | 43 | 0.19
US50 | RhAg | 09_25 | 0.42 | 44 | 0.191
US50 | RhAg | 09_25 | 0.42 | 45 | 0.191
US50 | RhAg | 09_25 | 0.42 | 46 | 0.192
US50 | RhAg | 09_25 | 0.42 | 47 | 0.192
US50 | RhAg | 09_25 | 0.42 | 48 | 0.193
US50 | RhAg | 09_25 | 0.42 | 49 | 0.194
US50 | RhAg | 09_25 | 0.44 | 21 | 0.162
US50 | RhAg | 09_25 | 0.44 | 22 | 0.165
US50 | RhAg | 09_25 | 0.44 | 23 | 0.169
US50 | RhAg | 09_25 | 0.44 | 24 | 0.173
US50 | RhAg | 09_25 | 0.44 | 25 | 0.176
US50 | RhAg | 09_25 | 0.44 | 26 | 0.18
US50 | RhAg | 09_25 | 0.44 | 27 | 0.182
US50 | RhAg | 09_25 | 0.44 | 28 | 0.184
US50 | RhAg | 09_25 | 0.44 | 29 | 0.185
US50 | RhAg | 09_25 | 0.44 | 30 | 0.186
US50 | RhAg | 09_25 | 0.44 | 31 | 0.188
US50 | RhAg | 09_25 | 0.44 | 32 | 0.188
US50 | RhAg | 09_25 | 0.44 | 33 | 0.189
US50 | RhAg | 09_25 | 0.44 | 34 | 0.19
US50 | RhAg | 09_25 | 0.44 | 35 | 0.191
US50 | RhAg | 09_25 | 0.44 | 36 | 0.192
US50 | RhAg | 09_25 | 0.44 | 37 | 0.193
US50 | RhAg | 09_25 | 0.44 | 38 | 0.193
US50 | RhAg | 09_25 | 0.44 | 39 | 0.194
US50 | RhAg | 09_25 | 0.44 | 40 | 0.195
US50 | RhAg | 09_25 | 0.44 | 41 | 0.195
US50 | RhAg | 09_25 | 0.44 | 42 | 0.196
US50 | RhAg | 09_25 | 0.44 | 43 | 0.197
US50 | RhAg | 09_25 | 0.44 | 44 | 0.197
US50 | RhAg | 09_25 | 0.44 | 45 | 0.198
US50 | RhAg | 09_25 | 0.44 | 46 | 0.199
US50 | RhAg | 09_25 | 0.44 | 47 | 0.199
US50 | RhAg | 09_25 | 0.44 | 48 | 0.2
US50 | RhAg | 09_25 | 0.44 | 49 | 0.2
US50 | RhAg | 09_25 | 0.46 | 22 | 0.173
US50 | RhAg | 09_25 | 0.46 | 23 | 0.177
US50 | RhAg | 09_25 | 0.46 | 24 | 0.181
US50 | RhAg | 09_25 | 0.46 | 25 | 0.184
US50 | RhAg | 09_25 | 0.46 | 26 | 0.187
US50 | RhAg | 09_25 | 0.46 | 27 | 0.19
US50 | RhAg | 09_25 | 0.46 | 28 | 0.192
US50 | RhAg | 09_25 | 0.46 | 29 | 0.193
US50 | RhAg | 09_25 | 0.46 | 30 | 0.194
US50 | RhAg | 09_25 | 0.46 | 31 | 0.195
US50 | RhAg | 09_25 | 0.46 | 32 | 0.196
US50 | RhAg | 09_25 | 0.46 | 33 | 0.196
US50 | RhAg | 09_25 | 0.46 | 34 | 0.197
US50 | RhAg | 09_25 | 0.46 | 35 | 0.198
US50 | RhAg | 09_25 | 0.46 | 36 | 0.199
US50 | RhAg | 09_25 | 0.46 | 37 | 0.199
US50 | RhAg | 09_25 | 0.46 | 38 | 0.2
US50 | RhAg | 09_25 | 0.46 | 39 | 0.201
US50 | RhAg | 09_25 | 0.46 | 40 | 0.201
US50 | RhAg | 09_25 | 0.46 | 41 | 0.202
US50 | RhAg | 09_25 | 0.46 | 42 | 0.203
US50 | RhAg | 09_25 | 0.46 | 43 | 0.203
US50 | RhAg | 09_25 | 0.46 | 44 | 0.204
US50 | RhAg | 09_25 | 0.46 | 45 | 0.205
US50 | RhAg | 09_25 | 0.46 | 46 | 0.205
US50 | RhAg | 09_25 | 0.46 | 47 | 0.206
US50 | RhAg | 09_25 | 0.46 | 48 | 0.206
US50 | RhAg | 09_25 | 0.46 | 49 | 0.207
US50 | RhAg | 09_25 | 0.48 | 22 | 0.182
US50 | RhAg | 09_25 | 0.48 | 23 | 0.185
US50 | RhAg | 09_25 | 0.48 | 24 | 0.189
US50 | RhAg | 09_25 | 0.48 | 25 | 0.192
US50 | RhAg | 09_25 | 0.48 | 26 | 0.195
US50 | RhAg | 09_25 | 0.48 | 27 | 0.197
US50 | RhAg | 09_25 | 0.48 | 28 | 0.199
US50 | RhAg | 09_25 | 0.48 | 29 | 0.2
US50 | RhAg | 09_25 | 0.48 | 30 | 0.201
US50 | RhAg | 09_25 | 0.48 | 31 | 0.202
US50 | RhAg | 09_25 | 0.48 | 32 | 0.203
US50 | RhAg | 09_25 | 0.48 | 33 | 0.204
US50 | RhAg | 09_25 | 0.48 | 34 | 0.204
US50 | RhAg | 09_25 | 0.48 | 35 | 0.205
US50 | RhAg | 09_25 | 0.48 | 36 | 0.205
US50 | RhAg | 09_25 | 0.48 | 37 | 0.206
US50 | RhAg | 09_25 | 0.48 | 38 | 0.207
US50 | RhAg | 09_25 | 0.48 | 39 | 0.207
US50 | RhAg | 09_25 | 0.48 | 40 | 0.208
US50 | RhAg | 09_25 | 0.48 | 41 | 0.209
US50 | RhAg | 09_25 | 0.48 | 42 | 0.209
US50 | RhAg | 09_25 | 0.48 | 43 | 0.21
US50 | RhAg | 09_25 | 0.48 | 44 | 0.21
US50 | RhAg | 09_25 | 0.48 | 45 | 0.211
US50 | RhAg | 09_25 | 0.48 | 46 | 0.212
US50 | RhAg | 09_25 | 0.48 | 47 | 0.212
US50 | RhAg | 09_25 | 0.48 | 48 | 0.212
US50 | RhAg | 09_25 | 0.48 | 49 | 0.213
US50 | RhAg | 09_25 | 0.5 | 23 | 0.194
US50 | RhAg | 09_25 | 0.5 | 24 | 0.197
US50 | RhAg | 09_25 | 0.5 | 25 | 0.2
US50 | RhAg | 09_25 | 0.5 | 26 | 0.203
US50 | RhAg | 09_25 | 0.5 | 27 | 0.205
US50 | RhAg | 09_25 | 0.5 | 28 | 0.207
US50 | RhAg | 09_25 | 0.5 | 29 | 0.208
US50 | RhAg | 09_25 | 0.5 | 30 | 0.209
US50 | RhAg | 09_25 | 0.5 | 31 | 0.209
US50 | RhAg | 09_25 | 0.5 | 32 | 0.21
US50 | RhAg | 09_25 | 0.5 | 33 | 0.21
US50 | RhAg | 09_25 | 0.5 | 34 | 0.211
US50 | RhAg | 09_25 | 0.5 | 35 | 0.212
US50 | RhAg | 09_25 | 0.5 | 36 | 0.212
US50 | RhAg | 09_25 | 0.5 | 37 | 0.213
US50 | RhAg | 09_25 | 0.5 | 38 | 0.213
US50 | RhAg | 09_25 | 0.5 | 39 | 0.214
US50 | RhAg | 09_25 | 0.5 | 40 | 0.215
US50 | RhAg | 09_25 | 0.5 | 41 | 0.215
US50 | RhAg | 09_25 | 0.5 | 42 | 0.216
US50 | RhAg | 09_25 | 0.5 | 43 | 0.216
US50 | RhAg | 09_25 | 0.5 | 44 | 0.217
US50 | RhAg | 09_25 | 0.5 | 45 | 0.217
US50 | RhAg | 09_25 | 0.5 | 46 | 0.218
US50 | RhAg | 09_25 | 0.5 | 47 | 0.218
US50 | RhAg | 09_25 | 0.5 | 48 | 0.219
US50 | RhAg | 09_25 | 0.5 | 49 | 0.219
US50 | RhAg | 09_25 | 0.52 | 23 | 0.202
US50 | RhAg | 09_25 | 0.52 | 24 | 0.205
US50 | RhAg | 09_25 | 0.52 | 25 | 0.208
US50 | RhAg | 09_25 | 0.52 | 26 | 0.211
US50 | RhAg | 09_25 | 0.52 | 27 | 0.213
US50 | RhAg | 09_25 | 0.52 | 28 | 0.214
US50 | RhAg | 09_25 | 0.52 | 29 | 0.215
US50 | RhAg | 09_25 | 0.52 | 30 | 0.216
US50 | RhAg | 09_25 | 0.52 | 31 | 0.216
US50 | RhAg | 09_25 | 0.52 | 32 | 0.217
US50 | RhAg | 09_25 | 0.52 | 33 | 0.217
US50 | RhAg | 09_25 | 0.52 | 34 | 0.218
US50 | RhAg | 09_25 | 0.52 | 35 | 0.218
US50 | RhAg | 09_25 | 0.52 | 36 | 0.219
US50 | RhAg | 09_25 | 0.52 | 37 | 0.219
US50 | RhAg | 09_25 | 0.52 | 38 | 0.22
US50 | RhAg | 09_25 | 0.52 | 39 | 0.221
US50 | RhAg | 09_25 | 0.52 | 40 | 0.221
US50 | RhAg | 09_25 | 0.52 | 41 | 0.222
US50 | RhAg | 09_25 | 0.52 | 42 | 0.222
US50 | RhAg | 09_25 | 0.52 | 43 | 0.223
US50 | RhAg | 09_25 | 0.52 | 44 | 0.223
US50 | RhAg | 09_25 | 0.52 | 45 | 0.224
US50 | RhAg | 09_25 | 0.52 | 46 | 0.224
US50 | RhAg | 09_25 | 0.52 | 47 | 0.225
US50 | RhAg | 09_25 | 0.52 | 48 | 0.225
US50 | RhAg | 09_25 | 0.52 | 49 | 0.226
US50 | RhAg | 09_25 | 0.54 | 24 | 0.213
US50 | RhAg | 09_25 | 0.54 | 25 | 0.216
US50 | RhAg | 09_25 | 0.54 | 26 | 0.218
US50 | RhAg | 09_25 | 0.54 | 27 | 0.22
US50 | RhAg | 09_25 | 0.54 | 28 | 0.221
US50 | RhAg | 09_25 | 0.54 | 29 | 0.222
US50 | RhAg | 09_25 | 0.54 | 30 | 0.223
US50 | RhAg | 09_25 | 0.54 | 31 | 0.223
US50 | RhAg | 09_25 | 0.54 | 32 | 0.224
US50 | RhAg | 09_25 | 0.54 | 33 | 0.224
US50 | RhAg | 09_25 | 0.54 | 34 | 0.225
US50 | RhAg | 09_25 | 0.54 | 35 | 0.225
US50 | RhAg | 09_25 | 0.54 | 36 | 0.225
US50 | RhAg | 09_25 | 0.54 | 37 | 0.226
US50 | RhAg | 09_25 | 0.54 | 38 | 0.226
US50 | RhAg | 09_25 | 0.54 | 39 | 0.227
US50 | RhAg | 09_25 | 0.54 | 40 | 0.227
US50 | RhAg | 09_25 | 0.54 | 41 | 0.228
US50 | RhAg | 09_25 | 0.54 | 42 | 0.229
US50 | RhAg | 09_25 | 0.54 | 43 | 0.229
US50 | RhAg | 09_25 | 0.54 | 44 | 0.229
US50 | RhAg | 09_25 | 0.54 | 45 | 0.23
US50 | RhAg | 09_25 | 0.54 | 46 | 0.23
US50 | RhAg | 09_25 | 0.54 | 47 | 0.231
US50 | RhAg | 09_25 | 0.54 | 48 | 0.231
US50 | RhAg | 09_25 | 0.54 | 49 | 0.232
US50 | RhAg | 09_25 | 0.56 | 26 | 0.226
US50 | RhAg | 09_25 | 0.56 | 27 | 0.227
US50 | RhAg | 09_25 | 0.56 | 28 | 0.228
US50 | RhAg | 09_25 | 0.56 | 29 | 0.229
US50 | RhAg | 09_25 | 0.56 | 30 | 0.23
US50 | RhAg | 09_25 | 0.56 | 31 | 0.23
US50 | RhAg | 09_25 | 0.56 | 32 | 0.231
US50 | RhAg | 09_25 | 0.56 | 33 | 0.231
US50 | RhAg | 09_25 | 0.56 | 34 | 0.231
US50 | RhAg | 09_25 | 0.56 | 35 | 0.232
US50 | RhAg | 09_25 | 0.56 | 36 | 0.232
US50 | RhAg | 09_25 | 0.56 | 37 | 0.232
US50 | RhAg | 09_25 | 0.56 | 38 | 0.233
US50 | RhAg | 09_25 | 0.56 | 39 | 0.233
US50 | RhAg | 09_25 | 0.56 | 40 | 0.234
US50 | RhAg | 09_25 | 0.56 | 41 | 0.234
US50 | RhAg | 09_25 | 0.56 | 42 | 0.235
US50 | RhAg | 09_25 | 0.56 | 43 | 0.235
US50 | RhAg | 09_25 | 0.56 | 44 | 0.236
US50 | RhAg | 09_25 | 0.56 | 45 | 0.236
US50 | RhAg | 09_25 | 0.56 | 46 | 0.237
US50 | RhAg | 09_25 | 0.56 | 47 | 0.237
US50 | RhAg | 09_25 | 0.56 | 48 | 0.237
US50 | RhAg | 09_25 | 0.56 | 49 | 0.238
US50 | RhAg | 09_25 | 0.58 | 28 | 0.236
US50 | RhAg | 09_25 | 0.58 | 29 | 0.236
US50 | RhAg | 09_25 | 0.58 | 30 | 0.237
US50 | RhAg | 09_25 | 0.58 | 31 | 0.237
US50 | RhAg | 09_25 | 0.58 | 32 | 0.237
US50 | RhAg | 09_25 | 0.58 | 33 | 0.238
US50 | RhAg | 09_25 | 0.58 | 34 | 0.238
US50 | RhAg | 09_25 | 0.58 | 35 | 0.238
US50 | RhAg | 09_25 | 0.58 | 36 | 0.239
US50 | RhAg | 09_25 | 0.58 | 37 | 0.239
US50 | RhAg | 09_25 | 0.58 | 38 | 0.239
US50 | RhAg | 09_25 | 0.58 | 39 | 0.24
US50 | RhAg | 09_25 | 0.58 | 40 | 0.24
US50 | RhAg | 09_25 | 0.58 | 41 | 0.241
US50 | RhAg | 09_25 | 0.58 | 42 | 0.241
US50 | RhAg | 09_25 | 0.58 | 43 | 0.241
US50 | RhAg | 09_25 | 0.58 | 44 | 0.242
US50 | RhAg | 09_25 | 0.58 | 45 | 0.242
US50 | RhAg | 09_25 | 0.58 | 46 | 0.243
US50 | RhAg | 09_25 | 0.58 | 47 | 0.243
US50 | RhAg | 09_25 | 0.58 | 48 | 0.244
US50 | RhAg | 09_25 | 0.58 | 49 | 0.244
US50 | RhAg | 09_25 | 0.6 | 30 | 0.244
US50 | RhAg | 09_25 | 0.6 | 31 | 0.244
US50 | RhAg | 09_25 | 0.6 | 32 | 0.244
US50 | RhAg | 09_25 | 0.6 | 33 | 0.244
US50 | RhAg | 09_25 | 0.6 | 34 | 0.245
US50 | RhAg | 09_25 | 0.6 | 35 | 0.245
US50 | RhAg | 09_25 | 0.6 | 36 | 0.245
US50 | RhAg | 09_25 | 0.6 | 37 | 0.245
US50 | RhAg | 09_25 | 0.6 | 38 | 0.246
US50 | RhAg | 09_25 | 0.6 | 39 | 0.246
US50 | RhAg | 09_25 | 0.6 | 40 | 0.246
US50 | RhAg | 09_25 | 0.6 | 41 | 0.247
US50 | RhAg | 09_25 | 0.6 | 42 | 0.247
US50 | RhAg | 09_25 | 0.6 | 43 | 0.248
US50 | RhAg | 09_25 | 0.6 | 44 | 0.248
US50 | RhAg | 09_25 | 0.6 | 45 | 0.248
US50 | RhAg | 09_25 | 0.6 | 46 | 0.249
US50 | RhAg | 09_25 | 0.6 | 47 | 0.249
US50 | RhAg | 09_25 | 0.6 | 48 | 0.25
US50 | RhAg | 09_25 | 0.6 | 49 | 0.25
US50 | RhAg | 09_25 | 0.62 | 33 | 0.251
US50 | RhAg | 09_25 | 0.62 | 34 | 0.251
US50 | RhAg | 09_25 | 0.62 | 35 | 0.251
US50 | RhAg | 09_25 | 0.62 | 36 | 0.252
US50 | RhAg | 09_25 | 0.62 | 37 | 0.252
US50 | RhAg | 09_25 | 0.62 | 38 | 0.252
US50 | RhAg | 09_25 | 0.62 | 39 | 0.252
US50 | RhAg | 09_25 | 0.62 | 40 | 0.253
US50 | RhAg | 09_25 | 0.62 | 41 | 0.253
US50 | RhAg | 09_25 | 0.62 | 42 | 0.253
US50 | RhAg | 09_25 | 0.62 | 43 | 0.254
US50 | RhAg | 09_25 | 0.62 | 44 | 0.254
US50 | RhAg | 09_25 | 0.62 | 45 | 0.254
US50 | RhAg | 09_25 | 0.62 | 46 | 0.255
US50 | RhAg | 09_25 | 0.62 | 47 | 0.255
US50 | RhAg | 09_25 | 0.62 | 48 | 0.256
US50 | RhAg | 09_25 | 0.62 | 49 | 0.256
US50 | RhAg | 09_25 | 0.64 | 35 | 0.258
US50 | RhAg | 09_25 | 0.64 | 36 | 0.258
US50 | RhAg | 09_25 | 0.64 | 37 | 0.258
US50 | RhAg | 09_25 | 0.64 | 38 | 0.258
US50 | RhAg | 09_25 | 0.64 | 39 | 0.259
US50 | RhAg | 09_25 | 0.64 | 40 | 0.259
US50 | RhAg | 09_25 | 0.64 | 41 | 0.259
US50 | RhAg | 09_25 | 0.64 | 42 | 0.26
US50 | RhAg | 09_25 | 0.64 | 43 | 0.26
US50 | RhAg | 09_25 | 0.64 | 44 | 0.26
US50 | RhAg | 09_25 | 0.64 | 45 | 0.261
US50 | RhAg | 09_25 | 0.64 | 46 | 0.261
US50 | RhAg | 09_25 | 0.64 | 47 | 0.261
US50 | RhAg | 09_25 | 0.64 | 48 | 0.262
US50 | RhAg | 09_25 | 0.64 | 49 | 0.262
US50 | RhAg | 09_25 | 0.66 | 38 | 0.265
US50 | RhAg | 09_25 | 0.66 | 39 | 0.265
US50 | RhAg | 09_25 | 0.66 | 40 | 0.265
US50 | RhAg | 09_25 | 0.66 | 41 | 0.265
US50 | RhAg | 09_25 | 0.66 | 42 | 0.266
US50 | RhAg | 09_25 | 0.66 | 43 | 0.266
US50 | RhAg | 09_25 | 0.66 | 44 | 0.266
US50 | RhAg | 09_25 | 0.66 | 45 | 0.267
US50 | RhAg | 09_25 | 0.66 | 46 | 0.267
US50 | RhAg | 09_25 | 0.66 | 47 | 0.267
US50 | RhAg | 09_25 | 0.66 | 48 | 0.268
US50 | RhAg | 09_25 | 0.66 | 49 | 0.268
US50 | RhAg | 09_25 | 0.68 | 44 | 0.272
US50 | RhAg | 09_25 | 0.68 | 45 | 0.273
US50 | RhAg | 09_25 | 0.68 | 46 | 0.273
US50 | RhAg | 09_25 | 0.68 | 47 | 0.273
US50 | RhAg | 09_25 | 0.68 | 48 | 0.274
US50 | RhAg | 09_25 | 0.68 | 49 | 0.274
US50 | RhAg | 09_25 | 0.7 | 47 | 0.28
US50 | RhAg | 09_25 | 0.7 | 48 | 0.28
US50 | RhAg | 09_25 | 0.7 | 49 | 0.28
US50 | RhRh | 09_25 | 0.2 | 20 | 0.054
US50 | RhRh | 09_25 | 0.2 | 21 | 0.057
US50 | RhRh | 09_25 | 0.2 | 22 | 0.06
US50 | RhRh | 09_25 | 0.2 | 23 | 0.062
US50 | RhRh | 09_25 | 0.2 | 24 | 0.065
US50 | RhRh | 09_25 | 0.2 | 25 | 0.068
US50 | RhRh | 09_25 | 0.22 | 20 | 0.063
US50 | RhRh | 09_25 | 0.22 | 21 | 0.067
US50 | RhRh | 09_25 | 0.22 | 22 | 0.069
US50 | RhRh | 09_25 | 0.22 | 23 | 0.072
US50 | RhRh | 09_25 | 0.22 | 24 | 0.076
US50 | RhRh | 09_25 | 0.22 | 25 | 0.079
US50 | RhRh | 09_25 | 0.22 | 26 | 0.082
US50 | RhRh | 09_25 | 0.22 | 27 | 0.084
US50 | RhRh | 09_25 | 0.24 | 20 | 0.072
US50 | RhRh | 09_25 | 0.24 | 21 | 0.075
US50 | RhRh | 09_25 | 0.24 | 22 | 0.078
US50 | RhRh | 09_25 | 0.24 | 23 | 0.081
US50 | RhRh | 09_25 | 0.24 | 24 | 0.085
US50 | RhRh | 09_25 | 0.24 | 25 | 0.089
US50 | RhRh | 09_25 | 0.24 | 26 | 0.092
US50 | RhRh | 09_25 | 0.24 | 27 | 0.094
US50 | RhRh | 09_25 | 0.24 | 28 | 0.097
US50 | RhRh | 09_25 | 0.24 | 29 | 0.099
US50 | RhRh | 09_25 | 0.26 | 20 | 0.08
US50 | RhRh | 09_25 | 0.26 | 21 | 0.084
US50 | RhRh | 09_25 | 0.26 | 22 | 0.087
US50 | RhRh | 09_25 | 0.26 | 23 | 0.09
US50 | RhRh | 09_25 | 0.26 | 24 | 0.095
US50 | RhRh | 09_25 | 0.26 | 25 | 0.098
US50 | RhRh | 09_25 | 0.26 | 26 | 0.101
US50 | RhRh | 09_25 | 0.26 | 27 | 0.104
US50 | RhRh | 09_25 | 0.26 | 28 | 0.106
US50 | RhRh | 09_25 | 0.26 | 29 | 0.108
US50 | RhRh | 09_25 | 0.26 | 30 | 0.11
US50 | RhRh | 09_25 | 0.26 | 31 | 0.112
US50 | RhRh | 09_25 | 0.26 | 32 | 0.114
US50 | RhRh | 09_25 | 0.28 | 20 | 0.089
US50 | RhRh | 09_25 | 0.28 | 21 | 0.093
US50 | RhRh | 09_25 | 0.28 | 22 | 0.096
US50 | RhRh | 09_25 | 0.28 | 23 | 0.1
US50 | RhRh | 09_25 | 0.28 | 24 | 0.104
US50 | RhRh | 09_25 | 0.28 | 25 | 0.108
US50 | RhRh | 09_25 | 0.28 | 26 | 0.111
US50 | RhRh | 09_25 | 0.28 | 27 | 0.113
US50 | RhRh | 09_25 | 0.28 | 28 | 0.116
US50 | RhRh | 09_25 | 0.28 | 29 | 0.118
US50 | RhRh | 09_25 | 0.28 | 30 | 0.119
US50 | RhRh | 09_25 | 0.28 | 31 | 0.121
US50 | RhRh | 09_25 | 0.28 | 32 | 0.123
US50 | RhRh | 09_25 | 0.28 | 33 | 0.124
US50 | RhRh | 09_25 | 0.28 | 34 | 0.126
US50 | RhRh | 09_25 | 0.3 | 20 | 0.097
US50 | RhRh | 09_25 | 0.3 | 21 | 0.102
US50 | RhRh | 09_25 | 0.3 | 22 | 0.105
US50 | RhRh | 09_25 | 0.3 | 23 | 0.108
US50 | RhRh | 09_25 | 0.3 | 24 | 0.113
US50 | RhRh | 09_25 | 0.3 | 25 | 0.117
US50 | RhRh | 09_25 | 0.3 | 26 | 0.12
US50 | RhRh | 09_25 | 0.3 | 27 | 0.122
US50 | RhRh | 09_25 | 0.3 | 28 | 0.124
US50 | RhRh | 09_25 | 0.3 | 29 | 0.126
US50 | RhRh | 09_25 | 0.3 | 30 | 0.128
US50 | RhRh | 09_25 | 0.3 | 31 | 0.13
US50 | RhRh | 09_25 | 0.3 | 32 | 0.132
US50 | RhRh | 09_25 | 0.3 | 33 | 0.133
US50 | RhRh | 09_25 | 0.3 | 34 | 0.134
US50 | RhRh | 09_25 | 0.3 | 35 | 0.136
US50 | RhRh | 09_25 | 0.3 | 36 | 0.137
US50 | RhRh | 09_25 | 0.3 | 37 | 0.138
US50 | RhRh | 09_25 | 0.32 | 20 | 0.105
US50 | RhRh | 09_25 | 0.32 | 21 | 0.11
US50 | RhRh | 09_25 | 0.32 | 22 | 0.113
US50 | RhRh | 09_25 | 0.32 | 23 | 0.117
US50 | RhRh | 09_25 | 0.32 | 24 | 0.122
US50 | RhRh | 09_25 | 0.32 | 25 | 0.126
US50 | RhRh | 09_25 | 0.32 | 26 | 0.129
US50 | RhRh | 09_25 | 0.32 | 27 | 0.131
US50 | RhRh | 09_25 | 0.32 | 28 | 0.133
US50 | RhRh | 09_25 | 0.32 | 29 | 0.135
US50 | RhRh | 09_25 | 0.32 | 30 | 0.137
US50 | RhRh | 09_25 | 0.32 | 31 | 0.138
US50 | RhRh | 09_25 | 0.32 | 32 | 0.14
US50 | RhRh | 09_25 | 0.32 | 33 | 0.141
US50 | RhRh | 09_25 | 0.32 | 34 | 0.143
US50 | RhRh | 09_25 | 0.32 | 35 | 0.144
US50 | RhRh | 09_25 | 0.32 | 36 | 0.145
US50 | RhRh | 09_25 | 0.32 | 37 | 0.146
US50 | RhRh | 09_25 | 0.32 | 38 | 0.148
US50 | RhRh | 09_25 | 0.32 | 39 | 0.149
US50 | RhRh | 09_25 | 0.34 | 20 | 0.114
US50 | RhRh | 09_25 | 0.34 | 21 | 0.119
US50 | RhRh | 09_25 | 0.34 | 22 | 0.122
US50 | RhRh | 09_25 | 0.34 | 23 | 0.126
US50 | RhRh | 09_25 | 0.34 | 24 | 0.131
US50 | RhRh | 09_25 | 0.34 | 25 | 0.134
US50 | RhRh | 09_25 | 0.34 | 26 | 0.137
US50 | RhRh | 09_25 | 0.34 | 27 | 0.14
US50 | RhRh | 09_25 | 0.34 | 28 | 0.141
US50 | RhRh | 09_25 | 0.34 | 29 | 0.143
US50 | RhRh | 09_25 | 0.34 | 30 | 0.145
US50 | RhRh | 09_25 | 0.34 | 31 | 0.146
US50 | RhRh | 09_25 | 0.34 | 32 | 0.148
US50 | RhRh | 09_25 | 0.34 | 33 | 0.149
US50 | RhRh | 09_25 | 0.34 | 34 | 0.151
US50 | RhRh | 09_25 | 0.34 | 35 | 0.152
US50 | RhRh | 09_25 | 0.34 | 36 | 0.153
US50 | RhRh | 09_25 | 0.34 | 37 | 0.154
US50 | RhRh | 09_25 | 0.34 | 38 | 0.155
US50 | RhRh | 09_25 | 0.34 | 39 | 0.156
US50 | RhRh | 09_25 | 0.34 | 40 | 0.157
US50 | RhRh | 09_25 | 0.34 | 41 | 0.158
US50 | RhRh | 09_25 | 0.34 | 42 | 0.159
US50 | RhRh | 09_25 | 0.36 | 20 | 0.123
US50 | RhRh | 09_25 | 0.36 | 21 | 0.127
US50 | RhRh | 09_25 | 0.36 | 22 | 0.131
US50 | RhRh | 09_25 | 0.36 | 23 | 0.135
US50 | RhRh | 09_25 | 0.36 | 24 | 0.139
US50 | RhRh | 09_25 | 0.36 | 25 | 0.143
US50 | RhRh | 09_25 | 0.36 | 26 | 0.146
US50 | RhRh | 09_25 | 0.36 | 27 | 0.148
US50 | RhRh | 09_25 | 0.36 | 28 | 0.15
US50 | RhRh | 09_25 | 0.36 | 29 | 0.151
US50 | RhRh | 09_25 | 0.36 | 30 | 0.153
US50 | RhRh | 09_25 | 0.36 | 31 | 0.154
US50 | RhRh | 09_25 | 0.36 | 32 | 0.156
US50 | RhRh | 09_25 | 0.36 | 33 | 0.157
US50 | RhRh | 09_25 | 0.36 | 34 | 0.158
US50 | RhRh | 09_25 | 0.36 | 35 | 0.159
US50 | RhRh | 09_25 | 0.36 | 36 | 0.16
US50 | RhRh | 09_25 | 0.36 | 37 | 0.162
US50 | RhRh | 09_25 | 0.36 | 38 | 0.163
US50 | RhRh | 09_25 | 0.36 | 39 | 0.164
US50 | RhRh | 09_25 | 0.36 | 40 | 0.165
US50 | RhRh | 09_25 | 0.36 | 41 | 0.166
US50 | RhRh | 09_25 | 0.36 | 42 | 0.167
US50 | RhRh | 09_25 | 0.36 | 43 | 0.167
US50 | RhRh | 09_25 | 0.36 | 44 | 0.168
US50 | RhRh | 09_25 | 0.36 | 45 | 0.169
US50 | RhRh | 09_25 | 0.36 | 46 | 0.17
US50 | RhRh | 09_25 | 0.38 | 20 | 0.131
US50 | RhRh | 09_25 | 0.38 | 21 | 0.136
US50 | RhRh | 09_25 | 0.38 | 22 | 0.139
US50 | RhRh | 09_25 | 0.38 | 23 | 0.143
US50 | RhRh | 09_25 | 0.38 | 24 | 0.148
US50 | RhRh | 09_25 | 0.38 | 25 | 0.151
US50 | RhRh | 09_25 | 0.38 | 26 | 0.154
US50 | RhRh | 09_25 | 0.38 | 27 | 0.156
US50 | RhRh | 09_25 | 0.38 | 28 | 0.158
US50 | RhRh | 09_25 | 0.38 | 29 | 0.159
US50 | RhRh | 09_25 | 0.38 | 30 | 0.161
US50 | RhRh | 09_25 | 0.38 | 31 | 0.162
US50 | RhRh | 09_25 | 0.38 | 32 | 0.163
US50 | RhRh | 09_25 | 0.38 | 33 | 0.164
US50 | RhRh | 09_25 | 0.38 | 34 | 0.166
US50 | RhRh | 09_25 | 0.38 | 35 | 0.167
US50 | RhRh | 09_25 | 0.38 | 36 | 0.168
US50 | RhRh | 09_25 | 0.38 | 37 | 0.169
US50 | RhRh | 09_25 | 0.38 | 38 | 0.17
US50 | RhRh | 09_25 | 0.38 | 39 | 0.171
US50 | RhRh | 09_25 | 0.38 | 40 | 0.172
US50 | RhRh | 09_25 | 0.38 | 41 | 0.173
US50 | RhRh | 09_25 | 0.38 | 42 | 0.174
US50 | RhRh | 09_25 | 0.38 | 43 | 0.174
US50 | RhRh | 09_25 | 0.38 | 44 | 0.175
US50 | RhRh | 09_25 | 0.38 | 45 | 0.176
US50 | RhRh | 09_25 | 0.38 | 46 | 0.177
US50 | RhRh | 09_25 | 0.38 | 47 | 0.177
US50 | RhRh | 09_25 | 0.38 | 48 | 0.178
US50 | RhRh | 09_25 | 0.38 | 49 | 0.178
US50 | RhRh | 09_25 | 0.4 | 20 | 0.14
US50 | RhRh | 09_25 | 0.4 | 21 | 0.144
US50 | RhRh | 09_25 | 0.4 | 22 | 0.148
US50 | RhRh | 09_25 | 0.4 | 23 | 0.152
US50 | RhRh | 09_25 | 0.4 | 24 | 0.156
US50 | RhRh | 09_25 | 0.4 | 25 | 0.16
US50 | RhRh | 09_25 | 0.4 | 26 | 0.162
US50 | RhRh | 09_25 | 0.4 | 27 | 0.164
US50 | RhRh | 09_25 | 0.4 | 28 | 0.166
US50 | RhRh | 09_25 | 0.4 | 29 | 0.167
US50 | RhRh | 09_25 | 0.4 | 30 | 0.168
US50 | RhRh | 09_25 | 0.4 | 31 | 0.17
US50 | RhRh | 09_25 | 0.4 | 32 | 0.171
US50 | RhRh | 09_25 | 0.4 | 33 | 0.172
US50 | RhRh | 09_25 | 0.4 | 34 | 0.173
US50 | RhRh | 09_25 | 0.4 | 35 | 0.174
US50 | RhRh | 09_25 | 0.4 | 36 | 0.175
US50 | RhRh | 09_25 | 0.4 | 37 | 0.176
US50 | RhRh | 09_25 | 0.4 | 38 | 0.177
US50 | RhRh | 09_25 | 0.4 | 39 | 0.178
US50 | RhRh | 09_25 | 0.4 | 40 | 0.179
US50 | RhRh | 09_25 | 0.4 | 41 | 0.18
US50 | RhRh | 09_25 | 0.4 | 42 | 0.181
US50 | RhRh | 09_25 | 0.4 | 43 | 0.181
US50 | RhRh | 09_25 | 0.4 | 44 | 0.182
US50 | RhRh | 09_25 | 0.4 | 45 | 0.183
US50 | RhRh | 09_25 | 0.4 | 46 | 0.183
US50 | RhRh | 09_25 | 0.4 | 47 | 0.184
US50 | RhRh | 09_25 | 0.4 | 48 | 0.185
US50 | RhRh | 09_25 | 0.4 | 49 | 0.185
US50 | RhRh | 09_25 | 0.42 | 21 | 0.153
US50 | RhRh | 09_25 | 0.42 | 22 | 0.156
US50 | RhRh | 09_25 | 0.42 | 23 | 0.16
US50 | RhRh | 09_25 | 0.42 | 24 | 0.164
US50 | RhRh | 09_25 | 0.42 | 25 | 0.168
US50 | RhRh | 09_25 | 0.42 | 26 | 0.17
US50 | RhRh | 09_25 | 0.42 | 27 | 0.172
US50 | RhRh | 09_25 | 0.42 | 28 | 0.173
US50 | RhRh | 09_25 | 0.42 | 29 | 0.175
US50 | RhRh | 09_25 | 0.42 | 30 | 0.176
US50 | RhRh | 09_25 | 0.42 | 31 | 0.177
US50 | RhRh | 09_25 | 0.42 | 32 | 0.178
US50 | RhRh | 09_25 | 0.42 | 33 | 0.179
US50 | RhRh | 09_25 | 0.42 | 34 | 0.18
US50 | RhRh | 09_25 | 0.42 | 35 | 0.181
US50 | RhRh | 09_25 | 0.42 | 36 | 0.182
US50 | RhRh | 09_25 | 0.42 | 37 | 0.183
US50 | RhRh | 09_25 | 0.42 | 38 | 0.184
US50 | RhRh | 09_25 | 0.42 | 39 | 0.185
US50 | RhRh | 09_25 | 0.42 | 40 | 0.186
US50 | RhRh | 09_25 | 0.42 | 41 | 0.186
US50 | RhRh | 09_25 | 0.42 | 42 | 0.187
US50 | RhRh | 09_25 | 0.42 | 43 | 0.188
US50 | RhRh | 09_25 | 0.42 | 44 | 0.189
US50 | RhRh | 09_25 | 0.42 | 45 | 0.189
US50 | RhRh | 09_25 | 0.42 | 46 | 0.19
US50 | RhRh | 09_25 | 0.42 | 47 | 0.191
US50 | RhRh | 09_25 | 0.42 | 48 | 0.191
US50 | RhRh | 09_25 | 0.42 | 49 | 0.192
US50 | RhRh | 09_25 | 0.44 | 21 | 0.162
US50 | RhRh | 09_25 | 0.44 | 22 | 0.165
US50 | RhRh | 09_25 | 0.44 | 23 | 0.169
US50 | RhRh | 09_25 | 0.44 | 24 | 0.173
US50 | RhRh | 09_25 | 0.44 | 25 | 0.176
US50 | RhRh | 09_25 | 0.44 | 26 | 0.178
US50 | RhRh | 09_25 | 0.44 | 27 | 0.18
US50 | RhRh | 09_25 | 0.44 | 28 | 0.181
US50 | RhRh | 09_25 | 0.44 | 29 | 0.182
US50 | RhRh | 09_25 | 0.44 | 30 | 0.183
US50 | RhRh | 09_25 | 0.44 | 31 | 0.184
US50 | RhRh | 09_25 | 0.44 | 32 | 0.185
US50 | RhRh | 09_25 | 0.44 | 33 | 0.186
US50 | RhRh | 09_25 | 0.44 | 34 | 0.187
US50 | RhRh | 09_25 | 0.44 | 35 | 0.188
US50 | RhRh | 09_25 | 0.44 | 36 | 0.189
US50 | RhRh | 09_25 | 0.44 | 37 | 0.19
US50 | RhRh | 09_25 | 0.44 | 38 | 0.191
US50 | RhRh | 09_25 | 0.44 | 39 | 0.191
US50 | RhRh | 09_25 | 0.44 | 40 | 0.192
US50 | RhRh | 09_25 | 0.44 | 41 | 0.193
US50 | RhRh | 09_25 | 0.44 | 42 | 0.194
US50 | RhRh | 09_25 | 0.44 | 43 | 0.195
US50 | RhRh | 09_25 | 0.44 | 44 | 0.195
US50 | RhRh | 09_25 | 0.44 | 45 | 0.196
US50 | RhRh | 09_25 | 0.44 | 46 | 0.197
US50 | RhRh | 09_25 | 0.44 | 47 | 0.197
US50 | RhRh | 09_25 | 0.44 | 48 | 0.198
US50 | RhRh | 09_25 | 0.44 | 49 | 0.198
US50 | RhRh | 09_25 | 0.46 | 22 | 0.173
US50 | RhRh | 09_25 | 0.46 | 23 | 0.177
US50 | RhRh | 09_25 | 0.46 | 24 | 0.181
US50 | RhRh | 09_25 | 0.46 | 25 | 0.184
US50 | RhRh | 09_25 | 0.46 | 26 | 0.186
US50 | RhRh | 09_25 | 0.46 | 27 | 0.187
US50 | RhRh | 09_25 | 0.46 | 28 | 0.189
US50 | RhRh | 09_25 | 0.46 | 29 | 0.19
US50 | RhRh | 09_25 | 0.46 | 30 | 0.19
US50 | RhRh | 09_25 | 0.46 | 31 | 0.191
US50 | RhRh | 09_25 | 0.46 | 32 | 0.192
US50 | RhRh | 09_25 | 0.46 | 33 | 0.193
US50 | RhRh | 09_25 | 0.46 | 34 | 0.194
US50 | RhRh | 09_25 | 0.46 | 35 | 0.195
US50 | RhRh | 09_25 | 0.46 | 36 | 0.196
US50 | RhRh | 09_25 | 0.46 | 37 | 0.196
US50 | RhRh | 09_25 | 0.46 | 38 | 0.197
US50 | RhRh | 09_25 | 0.46 | 39 | 0.198
US50 | RhRh | 09_25 | 0.46 | 40 | 0.199
US50 | RhRh | 09_25 | 0.46 | 41 | 0.2
US50 | RhRh | 09_25 | 0.46 | 42 | 0.2
US50 | RhRh | 09_25 | 0.46 | 43 | 0.201
US50 | RhRh | 09_25 | 0.46 | 44 | 0.202
US50 | RhRh | 09_25 | 0.46 | 45 | 0.202
US50 | RhRh | 09_25 | 0.46 | 46 | 0.203
US50 | RhRh | 09_25 | 0.46 | 47 | 0.204
US50 | RhRh | 09_25 | 0.46 | 48 | 0.204
US50 | RhRh | 09_25 | 0.46 | 49 | 0.205
US50 | RhRh | 09_25 | 0.48 | 22 | 0.182
US50 | RhRh | 09_25 | 0.48 | 23 | 0.185
US50 | RhRh | 09_25 | 0.48 | 24 | 0.189
US50 | RhRh | 09_25 | 0.48 | 25 | 0.192
US50 | RhRh | 09_25 | 0.48 | 26 | 0.194
US50 | RhRh | 09_25 | 0.48 | 27 | 0.195
US50 | RhRh | 09_25 | 0.48 | 28 | 0.196
US50 | RhRh | 09_25 | 0.48 | 29 | 0.197
US50 | RhRh | 09_25 | 0.48 | 30 | 0.198
US50 | RhRh | 09_25 | 0.48 | 31 | 0.199
US50 | RhRh | 09_25 | 0.48 | 32 | 0.199
US50 | RhRh | 09_25 | 0.48 | 33 | 0.2
US50 | RhRh | 09_25 | 0.48 | 34 | 0.201
US50 | RhRh | 09_25 | 0.48 | 35 | 0.202
US50 | RhRh | 09_25 | 0.48 | 36 | 0.202
US50 | RhRh | 09_25 | 0.48 | 37 | 0.203
US50 | RhRh | 09_25 | 0.48 | 38 | 0.204
US50 | RhRh | 09_25 | 0.48 | 39 | 0.205
US50 | RhRh | 09_25 | 0.48 | 40 | 0.205
US50 | RhRh | 09_25 | 0.48 | 41 | 0.206
US50 | RhRh | 09_25 | 0.48 | 42 | 0.207
US50 | RhRh | 09_25 | 0.48 | 43 | 0.208
US50 | RhRh | 09_25 | 0.48 | 44 | 0.208
US50 | RhRh | 09_25 | 0.48 | 45 | 0.209
US50 | RhRh | 09_25 | 0.48 | 46 | 0.209
US50 | RhRh | 09_25 | 0.48 | 47 | 0.21
US50 | RhRh | 09_25 | 0.48 | 48 | 0.21
US50 | RhRh | 09_25 | 0.48 | 49 | 0.211
US50 | RhRh | 09_25 | 0.5 | 23 | 0.194
US50 | RhRh | 09_25 | 0.5 | 24 | 0.197
US50 | RhRh | 09_25 | 0.5 | 25 | 0.2
US50 | RhRh | 09_25 | 0.5 | 26 | 0.201
US50 | RhRh | 09_25 | 0.5 | 27 | 0.203
US50 | RhRh | 09_25 | 0.5 | 28 | 0.203
US50 | RhRh | 09_25 | 0.5 | 29 | 0.204
US50 | RhRh | 09_25 | 0.5 | 30 | 0.205
US50 | RhRh | 09_25 | 0.5 | 31 | 0.206
US50 | RhRh | 09_25 | 0.5 | 32 | 0.206
US50 | RhRh | 09_25 | 0.5 | 33 | 0.207
US50 | RhRh | 09_25 | 0.5 | 34 | 0.208
US50 | RhRh | 09_25 | 0.5 | 35 | 0.208
US50 | RhRh | 09_25 | 0.5 | 36 | 0.209
US50 | RhRh | 09_25 | 0.5 | 37 | 0.21
US50 | RhRh | 09_25 | 0.5 | 38 | 0.21
US50 | RhRh | 09_25 | 0.5 | 39 | 0.211
US50 | RhRh | 09_25 | 0.5 | 40 | 0.212
US50 | RhRh | 09_25 | 0.5 | 41 | 0.212
US50 | RhRh | 09_25 | 0.5 | 42 | 0.213
US50 | RhRh | 09_25 | 0.5 | 43 | 0.214
US50 | RhRh | 09_25 | 0.5 | 44 | 0.214
US50 | RhRh | 09_25 | 0.5 | 45 | 0.215
US50 | RhRh | 09_25 | 0.5 | 46 | 0.216
US50 | RhRh | 09_25 | 0.5 | 47 | 0.216
US50 | RhRh | 09_25 | 0.5 | 48 | 0.217
US50 | RhRh | 09_25 | 0.5 | 49 | 0.217
US50 | RhRh | 09_25 | 0.52 | 23 | 0.202
US50 | RhRh | 09_25 | 0.52 | 24 | 0.205
US50 | RhRh | 09_25 | 0.52 | 25 | 0.208
US50 | RhRh | 09_25 | 0.52 | 26 | 0.209
US50 | RhRh | 09_25 | 0.52 | 27 | 0.21
US50 | RhRh | 09_25 | 0.52 | 28 | 0.211
US50 | RhRh | 09_25 | 0.52 | 29 | 0.211
US50 | RhRh | 09_25 | 0.52 | 30 | 0.212
US50 | RhRh | 09_25 | 0.52 | 31 | 0.213
US50 | RhRh | 09_25 | 0.52 | 32 | 0.213
US50 | RhRh | 09_25 | 0.52 | 33 | 0.214
US50 | RhRh | 09_25 | 0.52 | 34 | 0.214
US50 | RhRh | 09_25 | 0.52 | 35 | 0.215
US50 | RhRh | 09_25 | 0.52 | 36 | 0.215
US50 | RhRh | 09_25 | 0.52 | 37 | 0.216
US50 | RhRh | 09_25 | 0.52 | 38 | 0.217
US50 | RhRh | 09_25 | 0.52 | 39 | 0.217
US50 | RhRh | 09_25 | 0.52 | 40 | 0.218
US50 | RhRh | 09_25 | 0.52 | 41 | 0.219
US50 | RhRh | 09_25 | 0.52 | 42 | 0.219
US50 | RhRh | 09_25 | 0.52 | 43 | 0.22
US50 | RhRh | 09_25 | 0.52 | 44 | 0.221
US50 | RhRh | 09_25 | 0.52 | 45 | 0.221
US50 | RhRh | 09_25 | 0.52 | 46 | 0.222
US50 | RhRh | 09_25 | 0.52 | 47 | 0.222
US50 | RhRh | 09_25 | 0.52 | 48 | 0.223
US50 | RhRh | 09_25 | 0.52 | 49 | 0.223
US50 | RhRh | 09_25 | 0.54 | 24 | 0.213
US50 | RhRh | 09_25 | 0.54 | 25 | 0.215
US50 | RhRh | 09_25 | 0.54 | 26 | 0.216
US50 | RhRh | 09_25 | 0.54 | 27 | 0.217
US50 | RhRh | 09_25 | 0.54 | 28 | 0.218
US50 | RhRh | 09_25 | 0.54 | 29 | 0.219
US50 | RhRh | 09_25 | 0.54 | 30 | 0.219
US50 | RhRh | 09_25 | 0.54 | 31 | 0.22
US50 | RhRh | 09_25 | 0.54 | 32 | 0.22
US50 | RhRh | 09_25 | 0.54 | 33 | 0.22
US50 | RhRh | 09_25 | 0.54 | 34 | 0.221
US50 | RhRh | 09_25 | 0.54 | 35 | 0.221
US50 | RhRh | 09_25 | 0.54 | 36 | 0.222
US50 | RhRh | 09_25 | 0.54 | 37 | 0.223
US50 | RhRh | 09_25 | 0.54 | 38 | 0.223
US50 | RhRh | 09_25 | 0.54 | 39 | 0.224
US50 | RhRh | 09_25 | 0.54 | 40 | 0.224
US50 | RhRh | 09_25 | 0.54 | 41 | 0.225
US50 | RhRh | 09_25 | 0.54 | 42 | 0.226
US50 | RhRh | 09_25 | 0.54 | 43 | 0.226
US50 | RhRh | 09_25 | 0.54 | 44 | 0.227
US50 | RhRh | 09_25 | 0.54 | 45 | 0.227
US50 | RhRh | 09_25 | 0.54 | 46 | 0.228
US50 | RhRh | 09_25 | 0.54 | 47 | 0.228
US50 | RhRh | 09_25 | 0.54 | 48 | 0.229
US50 | RhRh | 09_25 | 0.54 | 49 | 0.229
US50 | RhRh | 09_25 | 0.56 | 26 | 0.224
US50 | RhRh | 09_25 | 0.56 | 27 | 0.225
US50 | RhRh | 09_25 | 0.56 | 28 | 0.225
US50 | RhRh | 09_25 | 0.56 | 29 | 0.226
US50 | RhRh | 09_25 | 0.56 | 30 | 0.226
US50 | RhRh | 09_25 | 0.56 | 31 | 0.226
US50 | RhRh | 09_25 | 0.56 | 32 | 0.227
US50 | RhRh | 09_25 | 0.56 | 33 | 0.227
US50 | RhRh | 09_25 | 0.56 | 34 | 0.228
US50 | RhRh | 09_25 | 0.56 | 35 | 0.228
US50 | RhRh | 09_25 | 0.56 | 36 | 0.229
US50 | RhRh | 09_25 | 0.56 | 37 | 0.229
US50 | RhRh | 09_25 | 0.56 | 38 | 0.23
US50 | RhRh | 09_25 | 0.56 | 39 | 0.23
US50 | RhRh | 09_25 | 0.56 | 40 | 0.231
US50 | RhRh | 09_25 | 0.56 | 41 | 0.231
US50 | RhRh | 09_25 | 0.56 | 42 | 0.232
US50 | RhRh | 09_25 | 0.56 | 43 | 0.232
US50 | RhRh | 09_25 | 0.56 | 44 | 0.233
US50 | RhRh | 09_25 | 0.56 | 45 | 0.233
US50 | RhRh | 09_25 | 0.56 | 46 | 0.234
US50 | RhRh | 09_25 | 0.56 | 47 | 0.234
US50 | RhRh | 09_25 | 0.56 | 48 | 0.235
US50 | RhRh | 09_25 | 0.56 | 49 | 0.235
US50 | RhRh | 09_25 | 0.58 | 28 | 0.232
US50 | RhRh | 09_25 | 0.58 | 29 | 0.233
US50 | RhRh | 09_25 | 0.58 | 30 | 0.233
US50 | RhRh | 09_25 | 0.58 | 31 | 0.233
US50 | RhRh | 09_25 | 0.58 | 32 | 0.234
US50 | RhRh | 09_25 | 0.58 | 33 | 0.234
US50 | RhRh | 09_25 | 0.58 | 34 | 0.234
US50 | RhRh | 09_25 | 0.58 | 35 | 0.235
US50 | RhRh | 09_25 | 0.58 | 36 | 0.235
US50 | RhRh | 09_25 | 0.58 | 37 | 0.236
US50 | RhRh | 09_25 | 0.58 | 38 | 0.236
US50 | RhRh | 09_25 | 0.58 | 39 | 0.236
US50 | RhRh | 09_25 | 0.58 | 40 | 0.237
US50 | RhRh | 09_25 | 0.58 | 41 | 0.237
US50 | RhRh | 09_25 | 0.58 | 42 | 0.238
US50 | RhRh | 09_25 | 0.58 | 43 | 0.239
US50 | RhRh | 09_25 | 0.58 | 44 | 0.239
US50 | RhRh | 09_25 | 0.58 | 45 | 0.239
US50 | RhRh | 09_25 | 0.58 | 46 | 0.24
US50 | RhRh | 09_25 | 0.58 | 47 | 0.241
US50 | RhRh | 09_25 | 0.58 | 48 | 0.241
US50 | RhRh | 09_25 | 0.58 | 49 | 0.241
US50 | RhRh | 09_25 | 0.6 | 30 | 0.24
US50 | RhRh | 09_25 | 0.6 | 31 | 0.24
US50 | RhRh | 09_25 | 0.6 | 32 | 0.24
US50 | RhRh | 09_25 | 0.6 | 33 | 0.241
US50 | RhRh | 09_25 | 0.6 | 34 | 0.241
US50 | RhRh | 09_25 | 0.6 | 35 | 0.241
US50 | RhRh | 09_25 | 0.6 | 36 | 0.242
US50 | RhRh | 09_25 | 0.6 | 37 | 0.242
US50 | RhRh | 09_25 | 0.6 | 38 | 0.242
US50 | RhRh | 09_25 | 0.6 | 39 | 0.243
US50 | RhRh | 09_25 | 0.6 | 40 | 0.243
US50 | RhRh | 09_25 | 0.6 | 41 | 0.244
US50 | RhRh | 09_25 | 0.6 | 42 | 0.244
US50 | RhRh | 09_25 | 0.6 | 43 | 0.245
US50 | RhRh | 09_25 | 0.6 | 44 | 0.245
US50 | RhRh | 09_25 | 0.6 | 45 | 0.246
US50 | RhRh | 09_25 | 0.6 | 46 | 0.246
US50 | RhRh | 09_25 | 0.6 | 47 | 0.247
US50 | RhRh | 09_25 | 0.6 | 48 | 0.247
US50 | RhRh | 09_25 | 0.6 | 49 | 0.247
US50 | RhRh | 09_25 | 0.62 | 33 | 0.247
US50 | RhRh | 09_25 | 0.62 | 34 | 0.247
US50 | RhRh | 09_25 | 0.62 | 35 | 0.248
US50 | RhRh | 09_25 | 0.62 | 36 | 0.248
US50 | RhRh | 09_25 | 0.62 | 37 | 0.248
US50 | RhRh | 09_25 | 0.62 | 38 | 0.249
US50 | RhRh | 09_25 | 0.62 | 39 | 0.249
US50 | RhRh | 09_25 | 0.62 | 40 | 0.249
US50 | RhRh | 09_25 | 0.62 | 41 | 0.25
US50 | RhRh | 09_25 | 0.62 | 42 | 0.25
US50 | RhRh | 09_25 | 0.62 | 43 | 0.251
US50 | RhRh | 09_25 | 0.62 | 44 | 0.251
US50 | RhRh | 09_25 | 0.62 | 45 | 0.252
US50 | RhRh | 09_25 | 0.62 | 46 | 0.252
US50 | RhRh | 09_25 | 0.62 | 47 | 0.252
US50 | RhRh | 09_25 | 0.62 | 48 | 0.253
US50 | RhRh | 09_25 | 0.62 | 49 | 0.253
US50 | RhRh | 09_25 | 0.64 | 35 | 0.254
US50 | RhRh | 09_25 | 0.64 | 36 | 0.254
US50 | RhRh | 09_25 | 0.64 | 37 | 0.255
US50 | RhRh | 09_25 | 0.64 | 38 | 0.255
US50 | RhRh | 09_25 | 0.64 | 39 | 0.255
US50 | RhRh | 09_25 | 0.64 | 40 | 0.256
US50 | RhRh | 09_25 | 0.64 | 41 | 0.256
US50 | RhRh | 09_25 | 0.64 | 42 | 0.257
US50 | RhRh | 09_25 | 0.64 | 43 | 0.257
US50 | RhRh | 09_25 | 0.64 | 44 | 0.257
US50 | RhRh | 09_25 | 0.64 | 45 | 0.258
US50 | RhRh | 09_25 | 0.64 | 46 | 0.258
US50 | RhRh | 09_25 | 0.64 | 47 | 0.259
US50 | RhRh | 09_25 | 0.64 | 48 | 0.259
US50 | RhRh | 09_25 | 0.64 | 49 | 0.259
US50 | RhRh | 09_25 | 0.66 | 38 | 0.261
US50 | RhRh | 09_25 | 0.66 | 39 | 0.262
US50 | RhRh | 09_25 | 0.66 | 40 | 0.262
US50 | RhRh | 09_25 | 0.66 | 41 | 0.262
US50 | RhRh | 09_25 | 0.66 | 42 | 0.263
US50 | RhRh | 09_25 | 0.66 | 43 | 0.263
US50 | RhRh | 09_25 | 0.66 | 44 | 0.263
US50 | RhRh | 09_25 | 0.66 | 45 | 0.264
US50 | RhRh | 09_25 | 0.66 | 46 | 0.264
US50 | RhRh | 09_25 | 0.66 | 47 | 0.265
US50 | RhRh | 09_25 | 0.66 | 48 | 0.265
US50 | RhRh | 09_25 | 0.66 | 49 | 0.265
US50 | RhRh | 09_25 | 0.68 | 40 | 0.268
US50 | RhRh | 09_25 | 0.68 | 41 | 0.268
US50 | RhRh | 09_25 | 0.68 | 42 | 0.269
US50 | RhRh | 09_25 | 0.68 | 43 | 0.269
US50 | RhRh | 09_25 | 0.68 | 44 | 0.27
US50 | RhRh | 09_25 | 0.68 | 45 | 0.27
US50 | RhRh | 09_25 | 0.68 | 46 | 0.27
US50 | RhRh | 09_25 | 0.68 | 47 | 0.271
US50 | RhRh | 09_25 | 0.68 | 48 | 0.271
US50 | RhRh | 09_25 | 0.68 | 49 | 0.271
US50 | RhRh | 09_25 | 0.7 | 43 | 0.275
US50 | RhRh | 09_25 | 0.7 | 44 | 0.276
US50 | RhRh | 09_25 | 0.7 | 45 | 0.276
US50 | RhRh | 09_25 | 0.7 | 46 | 0.276
US50 | RhRh | 09_25 | 0.7 | 47 | 0.277
US50 | RhRh | 09_25 | 0.7 | 48 | 0.277
US50 | RhRh | 09_25 | 0.7 | 49 | 0.278
US50 | RhRh | 09_25 | 0.72 | 47 | 0.283
US50 | RhRh | 09_25 | 0.72 | 48 | 0.283
US50 | RhRh | 09_25 | 0.72 | 49 | 0.284
US50 | WAl | 15_15 | 0.2 | 20 | 0.057
US50 | WAl | 15_15 | 0.2 | 21 | 0.06
US50 | WAl | 15_15 | 0.25 | 20 | 0.077
US50 | WAl | 15_15 | 0.25 | 21 | 0.082
US50 | WAl | 15_15 | 0.25 | 22 | 0.086
US50 | WAl | 15_15 | 0.25 | 23 | 0.09
US50 | WAl | 15_15 | 0.25 | 24 | 0.094
US50 | WAl | 15_15 | 0.25 | 25 | 0.097
US50 | WAl | 15_15 | 0.3 | 20 | 0.098
US50 | WAl | 15_15 | 0.3 | 21 | 0.103
US50 | WAl | 15_15 | 0.3 | 22 | 0.107
US50 | WAl | 15_15 | 0.3 | 23 | 0.111
US50 | WAl | 15_15 | 0.3 | 24 | 0.115
US50 | WAl | 15_15 | 0.3 | 25 | 0.119
US50 | WAl | 15_15 | 0.3 | 26 | 0.122
US50 | WAl | 15_15 | 0.3 | 27 | 0.125
US50 | WAl | 15_15 | 0.3 | 28 | 0.129
US50 | WAl | 15_15 | 0.3 | 29 | 0.132
US50 | WAl | 15_15 | 0.3 | 30 | 0.134
US50 | WAl | 15_15 | 0.35 | 20 | 0.119
US50 | WAl | 15_15 | 0.35 | 21 | 0.124
US50 | WAl | 15_15 | 0.35 | 22 | 0.128
US50 | WAl | 15_15 | 0.35 | 23 | 0.132
US50 | WAl | 15_15 | 0.35 | 24 | 0.136
US50 | WAl | 15_15 | 0.35 | 25 | 0.14
US50 | WAl | 15_15 | 0.35 | 26 | 0.143
US50 | WAl | 15_15 | 0.35 | 27 | 0.146
US50 | WAl | 15_15 | 0.35 | 28 | 0.15
US50 | WAl | 15_15 | 0.35 | 29 | 0.153
US50 | WAl | 15_15 | 0.35 | 30 | 0.155
US50 | WAl | 15_15 | 0.35 | 31 | 0.158
US50 | WAl | 15_15 | 0.35 | 32 | 0.161
US50 | WAl | 15_15 | 0.35 | 33 | 0.163
US50 | WAl | 15_15 | 0.35 | 34 | 0.166
US50 | WAl | 15_15 | 0.35 | 35 | 0.168
US50 | WAl | 15_15 | 0.4 | 20 | 0.141
US50 | WAl | 15_15 | 0.4 | 21 | 0.145
US50 | WAl | 15_15 | 0.4 | 22 | 0.149
US50 | WAl | 15_15 | 0.4 | 23 | 0.153
US50 | WAl | 15_15 | 0.4 | 24 | 0.157
US50 | WAl | 15_15 | 0.4 | 25 | 0.16
US50 | WAl | 15_15 | 0.4 | 26 | 0.163
US50 | WAl | 15_15 | 0.4 | 27 | 0.167
US50 | WAl | 15_15 | 0.4 | 28 | 0.17
US50 | WAl | 15_15 | 0.4 | 29 | 0.173
US50 | WAl | 15_15 | 0.4 | 30 | 0.175
US50 | WAl | 15_15 | 0.4 | 31 | 0.178
US50 | WAl | 15_15 | 0.4 | 32 | 0.181
US50 | WAl | 15_15 | 0.4 | 33 | 0.183
US50 | WAl | 15_15 | 0.4 | 34 | 0.185
US50 | WAl | 15_15 | 0.4 | 35 | 0.187
US50 | WAl | 15_15 | 0.4 | 36 | 0.19
US50 | WAl | 15_15 | 0.4 | 37 | 0.192
US50 | WAl | 15_15 | 0.4 | 38 | 0.193
US50 | WAl | 15_15 | 0.4 | 39 | 0.195
US50 | WAl | 15_15 | 0.4 | 40 | 0.197
US50 | WAl | 15_15 | 0.45 | 20 | 0.162
US50 | WAl | 15_15 | 0.45 | 21 | 0.166
US50 | WAl | 15_15 | 0.45 | 22 | 0.17
US50 | WAl | 15_15 | 0.45 | 23 | 0.174
US50 | WAl | 15_15 | 0.45 | 24 | 0.177
US50 | WAl | 15_15 | 0.45 | 25 | 0.181
US50 | WAl | 15_15 | 0.45 | 26 | 0.184
US50 | WAl | 15_15 | 0.45 | 27 | 0.187
US50 | WAl | 15_15 | 0.45 | 28 | 0.19
US50 | WAl | 15_15 | 0.45 | 29 | 0.192
US50 | WAl | 15_15 | 0.45 | 30 | 0.195
US50 | WAl | 15_15 | 0.45 | 31 | 0.197
US50 | WAl | 15_15 | 0.45 | 32 | 0.2
US50 | WAl | 15_15 | 0.45 | 33 | 0.202
US50 | WAl | 15_15 | 0.45 | 34 | 0.204
US50 | WAl | 15_15 | 0.45 | 35 | 0.206
US50 | WAl | 15_15 | 0.45 | 36 | 0.208
US50 | WAl | 15_15 | 0.45 | 37 | 0.21
US50 | WAl | 15_15 | 0.45 | 38 | 0.212
US50 | WAl | 15_15 | 0.45 | 39 | 0.214
US50 | WAl | 15_15 | 0.45 | 40 | 0.215
US50 | WAl | 15_15 | 0.45 | 41 | 0.217
US50 | WAl | 15_15 | 0.45 | 42 | 0.219
US50 | WAl | 15_15 | 0.45 | 43 | 0.22
US50 | WAl | 15_15 | 0.45 | 44 | 0.222
US50 | WAl | 15_15 | 0.5 | 21 | 0.188
US50 | WAl | 15_15 | 0.5 | 22 | 0.191
US50 | WAl | 15_15 | 0.5 | 23 | 0.194
US50 | WAl | 15_15 | 0.5 | 24 | 0.198
US50 | WAl | 15_15 | 0.5 | 25 | 0.201
US50 | WAl | 15_15 | 0.5 | 26 | 0.203
US50 | WAl | 15_15 | 0.5 | 27 | 0.206
US50 | WAl | 15_15 | 0.5 | 28 | 0.209
US50 | WAl | 15_15 | 0.5 | 29 | 0.212
US50 | WAl | 15_15 | 0.5 | 30 | 0.214
US50 | WAl | 15_15 | 0.5 | 31 | 0.216
US50 | WAl | 15_15 | 0.5 | 32 | 0.218
US50 | WAl | 15_15 | 0.5 | 33 | 0.221
US50 | WAl | 15_15 | 0.5 | 34 | 0.223
US50 | WAl | 15_15 | 0.5 | 35 | 0.224
US50 | WAl | 15_15 | 0.5 | 36 | 0.226
US50 | WAl | 15_15 | 0.5 | 37 | 0.228
US50 | WAl | 15_15 | 0.5 | 38 | 0.23
US50 | WAl | 15_15 | 0.5 | 39 | 0.231
US50 | WAl | 15_15 | 0.5 | 40 | 0.233
US50 | WAl | 15_15 | 0.5 | 41 | 0.235
US50 | WAl | 15_15 | 0.5 | 42 | 0.236
US50 | WAl | 15_15 | 0.5 | 43 | 0.237
US50 | WAl | 15_15 | 0.5 | 44 | 0.239
US50 | WAl | 15_15 | 0.5 | 45 | 0.24
US50 | WAl | 15_15 | 0.5 | 46 | 0.241
US50 | WAl | 15_15 | 0.5 | 47 | 0.243
US50 | WAl | 15_15 | 0.5 | 48 | 0.244
US50 | WAl | 15_15 | 0.5 | 49 | 0.245
US50 | WAl | 15_15 | 0.55 | 22 | 0.212
US50 | WAl | 15_15 | 0.55 | 23 | 0.214
US50 | WAl | 15_15 | 0.55 | 24 | 0.218
US50 | WAl | 15_15 | 0.55 | 25 | 0.22
US50 | WAl | 15_15 | 0.55 | 26 | 0.223
US50 | WAl | 15_15 | 0.55 | 27 | 0.225
US50 | WAl | 15_15 | 0.55 | 28 | 0.228
US50 | WAl | 15_15 | 0.55 | 29 | 0.23
US50 | WAl | 15_15 | 0.55 | 30 | 0.232
US50 | WAl | 15_15 | 0.55 | 31 | 0.235
US50 | WAl | 15_15 | 0.55 | 32 | 0.237
US50 | WAl | 15_15 | 0.55 | 33 | 0.238
US50 | WAl | 15_15 | 0.55 | 34 | 0.24
US50 | WAl | 15_15 | 0.55 | 35 | 0.242
US50 | WAl | 15_15 | 0.55 | 36 | 0.244
US50 | WAl | 15_15 | 0.55 | 37 | 0.245
US50 | WAl | 15_15 | 0.55 | 38 | 0.247
US50 | WAl | 15_15 | 0.55 | 39 | 0.248
US50 | WAl | 15_15 | 0.55 | 40 | 0.25
US50 | WAl | 15_15 | 0.55 | 41 | 0.251
US50 | WAl | 15_15 | 0.55 | 42 | 0.253
US50 | WAl | 15_15 | 0.55 | 43 | 0.254
US50 | WAl | 15_15 | 0.55 | 44 | 0.255
US50 | WAl | 15_15 | 0.55 | 45 | 0.257
US50 | WAl | 15_15 | 0.55 | 46 | 0.258
US50 | WAl | 15_15 | 0.55 | 47 | 0.259
US50 | WAl | 15_15 | 0.55 | 48 | 0.26
US50 | WAl | 15_15 | 0.55 | 49 | 0.261
US50 | WAl | 15_15 | 0.6 | 23 | 0.234
US50 | WAl | 15_15 | 0.6 | 24 | 0.237
US50 | WAl | 15_15 | 0.6 | 25 | 0.239
US50 | WAl | 15_15 | 0.6 | 26 | 0.241
US50 | WAl | 15_15 | 0.6 | 27 | 0.244
US50 | WAl | 15_15 | 0.6 | 28 | 0.246
US50 | WAl | 15_15 | 0.6 | 29 | 0.248
US50 | WAl | 15_15 | 0.6 | 30 | 0.25
US50 | WAl | 15_15 | 0.6 | 31 | 0.252
US50 | WAl | 15_15 | 0.6 | 32 | 0.254
US50 | WAl | 15_15 | 0.6 | 33 | 0.256
US50 | WAl | 15_15 | 0.6 | 34 | 0.258
US50 | WAl | 15_15 | 0.6 | 35 | 0.259
US50 | WAl | 15_15 | 0.6 | 36 | 0.261
US50 | WAl | 15_15 | 0.6 | 37 | 0.262
US50 | WAl | 15_15 | 0.6 | 38 | 0.264
US50 | WAl | 15_15 | 0.6 | 39 | 0.265
US50 | WAl | 15_15 | 0.6 | 40 | 0.266
US50 | WAl | 15_15 | 0.6 | 41 | 0.268
US50 | WAl | 15_15 | 0.6 | 42 | 0.269
US50 | WAl | 15_15 | 0.6 | 43 | 0.27
US50 | WAl | 15_15 | 0.6 | 44 | 0.271
US50 | WAl | 15_15 | 0.6 | 45 | 0.272
US50 | WAl | 15_15 | 0.6 | 46 | 0.273
US50 | WAl | 15_15 | 0.6 | 47 | 0.274
US50 | WAl | 15_15 | 0.6 | 48 | 0.276
US50 | WAl | 15_15 | 0.6 | 49 | 0.277
US50 | WAl | 15_15 | 0.65 | 24 | 0.255
US50 | WAl | 15_15 | 0.65 | 25 | 0.257
US50 | WAl | 15_15 | 0.65 | 26 | 0.259
US50 | WAl | 15_15 | 0.65 | 27 | 0.262
US50 | WAl | 15_15 | 0.65 | 28 | 0.264
US50 | WAl | 15_15 | 0.65 | 29 | 0.266
US50 | WAl | 15_15 | 0.65 | 30 | 0.268
US50 | WAl | 15_15 | 0.65 | 31 | 0.269
US50 | WAl | 15_15 | 0.65 | 32 | 0.271
US50 | WAl | 15_15 | 0.65 | 33 | 0.273
US50 | WAl | 15_15 | 0.65 | 34 | 0.274
US50 | WAl | 15_15 | 0.65 | 35 | 0.276
US50 | WAl | 15_15 | 0.65 | 36 | 0.277
US50 | WAl | 15_15 | 0.65 | 37 | 0.279
US50 | WAl | 15_15 | 0.65 | 38 | 0.28
US50 | WAl | 15_15 | 0.65 | 39 | 0.281
US50 | WAl | 15_15 | 0.65 | 40 | 0.282
US50 | WAl | 15_15 | 0.65 | 41 | 0.284
US50 | WAl | 15_15 | 0.65 | 42 | 0.285
US50 | WAl | 15_15 | 0.65 | 43 | 0.286
US50 | WAl | 15_15 | 0.65 | 44 | 0.287
US50 | WAl | 15_15 | 0.65 | 45 | 0.288
US50 | WAl | 15_15 | 0.65 | 46 | 0.289
US50 | WAl | 15_15 | 0.65 | 47 | 0.29
US50 | WAl | 15_15 | 0.65 | 48 | 0.291
US50 | WAl | 15_15 | 0.65 | 49 | 0.292
US50 | WAl | 15_15 | 0.7 | 25 | 0.275
US50 | WAl | 15_15 | 0.7 | 26 | 0.277
US50 | WAl | 15_15 | 0.7 | 27 | 0.279
US50 | WAl | 15_15 | 0.7 | 28 | 0.281
US50 | WAl | 15_15 | 0.7 | 29 | 0.283
US50 | WAl | 15_15 | 0.7 | 30 | 0.284
US50 | WAl | 15_15 | 0.7 | 31 | 0.286
US50 | WAl | 15_15 | 0.7 | 32 | 0.288
US50 | WAl | 15_15 | 0.7 | 33 | 0.289
US50 | WAl | 15_15 | 0.7 | 34 | 0.29
US50 | WAl | 15_15 | 0.7 | 35 | 0.292
US50 | WAl | 15_15 | 0.7 | 36 | 0.293
US50 | WAl | 15_15 | 0.7 | 37 | 0.294
US50 | WAl | 15_15 | 0.7 | 38 | 0.295
US50 | WAl | 15_15 | 0.7 | 39 | 0.297
US50 | WAl | 15_15 | 0.7 | 40 | 0.298
US50 | WAl | 15_15 | 0.7 | 41 | 0.299
US50 | WAl | 15_15 | 0.7 | 42 | 0.3
US50 | WAl | 15_15 | 0.7 | 43 | 0.301
US50 | WAl | 15_15 | 0.7 | 44 | 0.302
US50 | WAl | 15_15 | 0.7 | 45 | 0.303
US50 | WAl | 15_15 | 0.7 | 46 | 0.303
US50 | WAl | 15_15 | 0.7 | 47 | 0.304
US50 | WAl | 15_15 | 0.7 | 48 | 0.305
US50 | WAl | 15_15 | 0.7 | 49 | 0.306
US50 | WAl | 15_15 | 0.75 | 26 | 0.294
US50 | WAl | 15_15 | 0.75 | 27 | 0.295
US50 | WAl | 15_15 | 0.75 | 28 | 0.297
US50 | WAl | 15_15 | 0.75 | 29 | 0.299
US50 | WAl | 15_15 | 0.75 | 30 | 0.3
US50 | WAl | 15_15 | 0.75 | 31 | 0.302
US50 | WAl | 15_15 | 0.75 | 32 | 0.303
US50 | WAl | 15_15 | 0.75 | 33 | 0.305
US50 | WAl | 15_15 | 0.75 | 34 | 0.306
US50 | WAl | 15_15 | 0.75 | 35 | 0.307
US50 | WAl | 15_15 | 0.75 | 36 | 0.308
US50 | WAl | 15_15 | 0.75 | 37 | 0.309
US50 | WAl | 15_15 | 0.75 | 38 | 0.311
US50 | WAl | 15_15 | 0.75 | 39 | 0.312
US50 | WAl | 15_15 | 0.75 | 40 | 0.313
US50 | WAl | 15_15 | 0.75 | 41 | 0.313
US50 | WAl | 15_15 | 0.75 | 42 | 0.314
US50 | WAl | 15_15 | 0.75 | 43 | 0.315
US50 | WAl | 15_15 | 0.75 | 44 | 0.316
US50 | WAl | 15_15 | 0.75 | 45 | 0.317
US50 | WAl | 15_15 | 0.75 | 46 | 0.318
US50 | WAl | 15_15 | 0.75 | 47 | 0.319
US50 | WAl | 15_15 | 0.75 | 48 | 0.319
US50 | WAl | 15_15 | 0.75 | 49 | 0.32
US50 | WAl | 15_15 | 0.8 | 27 | 0.311
US50 | WAl | 15_15 | 0.8 | 28 | 0.313
US50 | WAl | 15_15 | 0.8 | 29 | 0.314
US50 | WAl | 15_15 | 0.8 | 30 | 0.316
US50 | WAl | 15_15 | 0.8 | 31 | 0.317
US50 | WAl | 15_15 | 0.8 | 32 | 0.318
US50 | WAl | 15_15 | 0.8 | 33 | 0.32
US50 | WAl | 15_15 | 0.8 | 34 | 0.321
US50 | WAl | 15_15 | 0.8 | 35 | 0.322
US50 | WAl | 15_15 | 0.8 | 36 | 0.323
US50 | WAl | 15_15 | 0.8 | 37 | 0.324
US50 | WAl | 15_15 | 0.8 | 38 | 0.325
US50 | WAl | 15_15 | 0.8 | 39 | 0.326
US50 | WAl | 15_15 | 0.8 | 40 | 0.327
US50 | WAl | 15_15 | 0.8 | 41 | 0.328
US50 | WAl | 15_15 | 0.8 | 42 | 0.328
US50 | WAl | 15_15 | 0.8 | 43 | 0.329
US50 | WAl | 15_15 | 0.8 | 44 | 0.33
US50 | WAl | 15_15 | 0.8 | 45 | 0.331
US50 | WAl | 15_15 | 0.8 | 46 | 0.331
US50 | WAl | 15_15 | 0.8 | 47 | 0.332
US50 | WAl | 15_15 | 0.8 | 48 | 0.333
US50 | WAl | 15_15 | 0.8 | 49 | 0.334
US50 | WAl | 15_15 | 0.85 | 28 | 0.328
US50 | WAl | 15_15 | 0.85 | 29 | 0.329
US50 | WAl | 15_15 | 0.85 | 30 | 0.33
US50 | WAl | 15_15 | 0.85 | 31 | 0.332
US50 | WAl | 15_15 | 0.85 | 32 | 0.333
US50 | WAl | 15_15 | 0.85 | 33 | 0.334
US50 | WAl | 15_15 | 0.85 | 34 | 0.335
US50 | WAl | 15_15 | 0.85 | 35 | 0.336
US50 | WAl | 15_15 | 0.85 | 36 | 0.337
US50 | WAl | 15_15 | 0.85 | 37 | 0.338
US50 | WAl | 15_15 | 0.85 | 38 | 0.339
US50 | WAl | 15_15 | 0.85 | 39 | 0.34
US50 | WAl | 15_15 | 0.85 | 40 | 0.341
US50 | WAl | 15_15 | 0.85 | 41 | 0.341
US50 | WAl | 15_15 | 0.85 | 42 | 0.342
US50 | WAl | 15_15 | 0.85 | 43 | 0.343
US50 | WAl | 15_15 | 0.85 | 44 | 0.343
US50 | WAl | 15_15 | 0.85 | 45 | 0.344
US50 | WAl | 15_15 | 0.85 | 46 | 0.345
US50 | WAl | 15_15 | 0.85 | 47 | 0.345
US50 | WAl | 15_15 | 0.85 | 48 | 0.346
US50 | WAl | 15_15 | 0.85 | 49 | 0.346
US50 | WAl | 15_15 | 0.9 | 28 | 0.342
US50 | WAl | 15_15 | 0.9 | 29 | 0.344
US50 | WAl | 15_15 | 0.9 | 30 | 0.345
US50 | WAl | 15_15 | 0.9 | 31 | 0.346
US50 | WAl | 15_15 | 0.9 | 32 | 0.347
US50 | WAl | 15_15 | 0.9 | 33 | 0.348
US50 | WAl | 15_15 | 0.9 | 34 | 0.349
US50 | WAl | 15_15 | 0.9 | 35 | 0.35
US50 | WAl | 15_15 | 0.9 | 36 | 0.351
US50 | WAl | 15_15 | 0.9 | 37 | 0.351
US50 | WAl | 15_15 | 0.9 | 38 | 0.352
US50 | WAl | 15_15 | 0.9 | 39 | 0.353
US50 | WAl | 15_15 | 0.9 | 40 | 0.354
US50 | WAl | 15_15 | 0.9 | 41 | 0.354
US50 | WAl | 15_15 | 0.9 | 42 | 0.355
US50 | WAl | 15_15 | 0.9 | 43 | 0.356
US50 | WAl | 15_15 | 0.9 | 44 | 0.356
US50 | WAl | 15_15 | 0.9 | 45 | 0.357
US50 | WAl | 15_15 | 0.9 | 46 | 0.357
US50 | WAl | 15_15 | 0.9 | 47 | 0.358
US50 | WAl | 15_15 | 0.9 | 48 | 0.359
US50 | WAl | 15_15 | 0.9 | 49 | 0.359
US50 | WAl | 15_15 | 0.95 | 29 | 0.357
US50 | WAl | 15_15 | 0.95 | 30 | 0.358
US50 | WAl | 15_15 | 0.95 | 31 | 0.359
US50 | WAl | 15_15 | 0.95 | 32 | 0.36
US50 | WAl | 15_15 | 0.95 | 33 | 0.361
US50 | WAl | 15_15 | 0.95 | 34 | 0.362
US50 | WAl | 15_15 | 0.95 | 35 | 0.363
US50 | WAl | 15_15 | 0.95 | 36 | 0.364
US50 | WAl | 15_15 | 0.95 | 37 | 0.364
US50 | WAl | 15_15 | 0.95 | 38 | 0.365
US50 | WAl | 15_15 | 0.95 | 39 | 0.366
US50 | WAl | 15_15 | 0.95 | 40 | 0.366
US50 | WAl | 15_15 | 0.95 | 41 | 0.367
US50 | WAl | 15_15 | 0.95 | 42 | 0.368
US50 | WAl | 15_15 | 0.95 | 43 | 0.368
US50 | WAl | 15_15 | 0.95 | 44 | 0.369
US50 | WAl | 15_15 | 0.95 | 45 | 0.369
US50 | WAl | 15_15 | 0.95 | 46 | 0.37
US50 | WAl | 15_15 | 0.95 | 47 | 0.37
US50 | WAl | 15_15 | 0.95 | 48 | 0.371
US50 | WAl | 15_15 | 0.95 | 49 | 0.371
US50 | WAl | 15_15 | 1 | 30 | 0.371
US50 | WAl | 15_15 | 1 | 31 | 0.372
US50 | WAl | 15_15 | 1 | 32 | 0.373
US50 | WAl | 15_15 | 1 | 33 | 0.374
US50 | WAl | 15_15 | 1 | 34 | 0.375
US50 | WAl | 15_15 | 1 | 35 | 0.375
US50 | WAl | 15_15 | 1 | 36 | 0.376
US50 | WAl | 15_15 | 1 | 37 | 0.377
US50 | WAl | 15_15 | 1 | 38 | 0.377
US50 | WAl | 15_15 | 1 | 39 | 0.378
US50 | WAl | 15_15 | 1 | 40 | 0.378
US50 | WAl | 15_15 | 1 | 41 | 0.379
US50 | WAl | 15_15 | 1 | 42 | 0.38
US50 | WAl | 15_15 | 1 | 43 | 0.38
US50 | WAl | 15_15 | 1 | 44 | 0.381
US50 | WAl | 15_15 | 1 | 45 | 0.381
US50 | WAl | 15_15 | 1 | 46 | 0.381
US50 | WAl | 15_15 | 1 | 47 | 0.382
US50 | WAl | 15_15 | 1 | 48 | 0.382
US50 | WAl | 15_15 | 1 | 49 | 0.383
US50 | WAl | 15_15 | 1.05 | 31 | 0.384
US50 | WAl | 15_15 | 1.05 | 32 | 0.385
US50 | WAl | 15_15 | 1.05 | 33 | 0.386
US50 | WAl | 15_15 | 1.05 | 34 | 0.387
US50 | WAl | 15_15 | 1.05 | 35 | 0.387
US50 | WAl | 15_15 | 1.05 | 36 | 0.388
US50 | WAl | 15_15 | 1.05 | 37 | 0.389
US50 | WAl | 15_15 | 1.05 | 38 | 0.389
US50 | WAl | 15_15 | 1.05 | 39 | 0.39
US50 | WAl | 15_15 | 1.05 | 40 | 0.39
US50 | WAl | 15_15 | 1.05 | 41 | 0.391
US50 | WAl | 15_15 | 1.05 | 42 | 0.391
US50 | WAl | 15_15 | 1.05 | 43 | 0.392
US50 | WAl | 15_15 | 1.05 | 44 | 0.392
US50 | WAl | 15_15 | 1.05 | 45 | 0.392
US50 | WAl | 15_15 | 1.05 | 46 | 0.393
US50 | WAl | 15_15 | 1.05 | 47 | 0.393
US50 | WAl | 15_15 | 1.05 | 48 | 0.394
US50 | WAl | 15_15 | 1.05 | 49 | 0.394
US50 | WAl | 15_15 | 1.1 | 32 | 0.397
US50 | WAl | 15_15 | 1.1 | 33 | 0.398
US50 | WAl | 15_15 | 1.1 | 34 | 0.398
US50 | WAl | 15_15 | 1.1 | 35 | 0.399
US50 | WAl | 15_15 | 1.1 | 36 | 0.399
US50 | WAl | 15_15 | 1.1 | 37 | 0.4
US50 | WAl | 15_15 | 1.1 | 38 | 0.4
US50 | WAl | 15_15 | 1.1 | 39 | 0.401
US50 | WAl | 15_15 | 1.1 | 40 | 0.401
US50 | WAl | 15_15 | 1.1 | 41 | 0.402
US50 | WAl | 15_15 | 1.1 | 42 | 0.402
US50 | WAl | 15_15 | 1.1 | 43 | 0.403
US50 | WAl | 15_15 | 1.1 | 44 | 0.403
US50 | WAl | 15_15 | 1.1 | 45 | 0.403
US50 | WAl | 15_15 | 1.1 | 46 | 0.404
US50 | WAl | 15_15 | 1.1 | 47 | 0.404
US50 | WAl | 15_15 | 1.1 | 48 | 0.404
US50 | WAl | 15_15 | 1.1 | 49 | 0.405
US50 | WAl | 15_15 | 1.15 | 33 | 0.409
US50 | WAl | 15_15 | 1.15 | 34 | 0.41
US50 | WAl | 15_15 | 1.15 | 35 | 0.41
US50 | WAl | 15_15 | 1.15 | 36 | 0.41
US50 | WAl | 15_15 | 1.15 | 37 | 0.411
US50 | WAl | 15_15 | 1.15 | 38 | 0.411
US50 | WAl | 15_15 | 1.15 | 39 | 0.412
US50 | WAl | 15_15 | 1.15 | 40 | 0.412
US50 | WAl | 15_15 | 1.15 | 41 | 0.413
US50 | WAl | 15_15 | 1.15 | 42 | 0.413
US50 | WAl | 15_15 | 1.15 | 43 | 0.413
US50 | WAl | 15_15 | 1.15 | 44 | 0.414
US50 | WAl | 15_15 | 1.15 | 45 | 0.414
US50 | WAl | 15_15 | 1.15 | 46 | 0.414
US50 | WAl | 15_15 | 1.15 | 47 | 0.414
US50 | WAl | 15_15 | 1.15 | 48 | 0.415
US50 | WAl | 15_15 | 1.15 | 49 | 0.415
US50 | WAl | 15_15 | 1.2 | 34 | 0.42
US50 | WAl | 15_15 | 1.2 | 35 | 0.421
US50 | WAl | 15_15 | 1.2 | 36 | 0.421
US50 | WAl | 15_15 | 1.2 | 37 | 0.421
US50 | WAl | 15_15 | 1.2 | 38 | 0.422
US50 | WAl | 15_15 | 1.2 | 39 | 0.422
US50 | WAl | 15_15 | 1.2 | 40 | 0.422
US50 | WAl | 15_15 | 1.2 | 41 | 0.423
US50 | WAl | 15_15 | 1.2 | 42 | 0.423
US50 | WAl | 15_15 | 1.2 | 43 | 0.423
US50 | WAl | 15_15 | 1.2 | 44 | 0.424
US50 | WAl | 15_15 | 1.2 | 45 | 0.424
US50 | WAl | 15_15 | 1.2 | 46 | 0.424
US50 | WAl | 15_15 | 1.2 | 47 | 0.424
US50 | WAl | 15_15 | 1.2 | 48 | 0.425
US50 | WAl | 15_15 | 1.2 | 49 | 0.425
US50 | WAl | 15_15 | 1.25 | 36 | 0.431
US50 | WAl | 15_15 | 1.25 | 37 | 0.432
US50 | WAl | 15_15 | 1.25 | 38 | 0.432
US50 | WAl | 15_15 | 1.25 | 39 | 0.432
US50 | WAl | 15_15 | 1.25 | 40 | 0.432
US50 | WAl | 15_15 | 1.25 | 41 | 0.433
US50 | WAl | 15_15 | 1.25 | 42 | 0.433
US50 | WAl | 15_15 | 1.25 | 43 | 0.433
US50 | WAl | 15_15 | 1.25 | 44 | 0.433
US50 | WAl | 15_15 | 1.25 | 45 | 0.434
US50 | WAl | 15_15 | 1.25 | 46 | 0.434
US50 | WAl | 15_15 | 1.25 | 47 | 0.434
US50 | WAl | 15_15 | 1.25 | 48 | 0.434
US50 | WAl | 15_15 | 1.25 | 49 | 0.435
US50 | WAl | 15_15 | 1.3 | 41 | 0.442
US50 | WAl | 15_15 | 1.3 | 42 | 0.443
US50 | WAl | 15_15 | 1.3 | 43 | 0.443
US50 | WAl | 15_15 | 1.3 | 44 | 0.443
US50 | WAl | 15_15 | 1.3 | 45 | 0.443
US50 | WAl | 15_15 | 1.3 | 46 | 0.443
US50 | WAl | 15_15 | 1.3 | 47 | 0.443
US50 | WAl | 15_15 | 1.3 | 48 | 0.444
US50 | WAl | 15_15 | 1.3 | 49 | 0.444
US50 | WAl | 15_15 | 1.35 | 45 | 0.452
US50 | WAl | 15_15 | 1.35 | 46 | 0.452
US50 | WAl | 15_15 | 1.35 | 47 | 0.452
US50 | WAl | 15_15 | 1.35 | 48 | 0.453
US50 | WAl | 15_15 | 1.35 | 49 | 0.453
US50 | WAl | 30_15 | 0.2 | 20 | 0.057
US50 | WAl | 30_15 | 0.2 | 21 | 0.06
US50 | WAl | 30_15 | 0.25 | 20 | 0.077
US50 | WAl | 30_15 | 0.25 | 21 | 0.082
US50 | WAl | 30_15 | 0.25 | 22 | 0.086
US50 | WAl | 30_15 | 0.25 | 23 | 0.09
US50 | WAl | 30_15 | 0.25 | 24 | 0.094
US50 | WAl | 30_15 | 0.25 | 25 | 0.097
US50 | WAl | 30_15 | 0.3 | 20 | 0.098
US50 | WAl | 30_15 | 0.3 | 21 | 0.103
US50 | WAl | 30_15 | 0.3 | 22 | 0.107
US50 | WAl | 30_15 | 0.3 | 23 | 0.111
US50 | WAl | 30_15 | 0.3 | 24 | 0.115
US50 | WAl | 30_15 | 0.3 | 25 | 0.119
US50 | WAl | 30_15 | 0.3 | 26 | 0.122
US50 | WAl | 30_15 | 0.3 | 27 | 0.125
US50 | WAl | 30_15 | 0.3 | 28 | 0.129
US50 | WAl | 30_15 | 0.3 | 29 | 0.132
US50 | WAl | 30_15 | 0.3 | 30 | 0.134
US50 | WAl | 30_15 | 0.35 | 20 | 0.119
US50 | WAl | 30_15 | 0.35 | 21 | 0.124
US50 | WAl | 30_15 | 0.35 | 22 | 0.128
US50 | WAl | 30_15 | 0.35 | 23 | 0.132
US50 | WAl | 30_15 | 0.35 | 24 | 0.136
US50 | WAl | 30_15 | 0.35 | 25 | 0.14
US50 | WAl | 30_15 | 0.35 | 26 | 0.143
US50 | WAl | 30_15 | 0.35 | 27 | 0.146
US50 | WAl | 30_15 | 0.35 | 28 | 0.15
US50 | WAl | 30_15 | 0.35 | 29 | 0.153
US50 | WAl | 30_15 | 0.35 | 30 | 0.155
US50 | WAl | 30_15 | 0.35 | 31 | 0.158
US50 | WAl | 30_15 | 0.35 | 32 | 0.161
US50 | WAl | 30_15 | 0.35 | 33 | 0.163
US50 | WAl | 30_15 | 0.35 | 34 | 0.166
US50 | WAl | 30_15 | 0.35 | 35 | 0.168
US50 | WAl | 30_15 | 0.4 | 20 | 0.141
US50 | WAl | 30_15 | 0.4 | 21 | 0.145
US50 | WAl | 30_15 | 0.4 | 22 | 0.149
US50 | WAl | 30_15 | 0.4 | 23 | 0.153
US50 | WAl | 30_15 | 0.4 | 24 | 0.157
US50 | WAl | 30_15 | 0.4 | 25 | 0.16
US50 | WAl | 30_15 | 0.4 | 26 | 0.163
US50 | WAl | 30_15 | 0.4 | 27 | 0.167
US50 | WAl | 30_15 | 0.4 | 28 | 0.17
US50 | WAl | 30_15 | 0.4 | 29 | 0.173
US50 | WAl | 30_15 | 0.4 | 30 | 0.176
US50 | WAl | 30_15 | 0.4 | 31 | 0.178
US50 | WAl | 30_15 | 0.4 | 32 | 0.181
US50 | WAl | 30_15 | 0.4 | 33 | 0.183
US50 | WAl | 30_15 | 0.4 | 34 | 0.185
US50 | WAl | 30_15 | 0.4 | 35 | 0.188
US50 | WAl | 30_15 | 0.4 | 36 | 0.19
US50 | WAl | 30_15 | 0.4 | 37 | 0.192
US50 | WAl | 30_15 | 0.4 | 38 | 0.194
US50 | WAl | 30_15 | 0.4 | 39 | 0.195
US50 | WAl | 30_15 | 0.4 | 40 | 0.197
US50 | WAl | 30_15 | 0.45 | 20 | 0.162
US50 | WAl | 30_15 | 0.45 | 21 | 0.166
US50 | WAl | 30_15 | 0.45 | 22 | 0.17
US50 | WAl | 30_15 | 0.45 | 23 | 0.174
US50 | WAl | 30_15 | 0.45 | 24 | 0.178
US50 | WAl | 30_15 | 0.45 | 25 | 0.181
US50 | WAl | 30_15 | 0.45 | 26 | 0.184
US50 | WAl | 30_15 | 0.45 | 27 | 0.187
US50 | WAl | 30_15 | 0.45 | 28 | 0.19
US50 | WAl | 30_15 | 0.45 | 29 | 0.193
US50 | WAl | 30_15 | 0.45 | 30 | 0.195
US50 | WAl | 30_15 | 0.45 | 31 | 0.197
US50 | WAl | 30_15 | 0.45 | 32 | 0.2
US50 | WAl | 30_15 | 0.45 | 33 | 0.202
US50 | WAl | 30_15 | 0.45 | 34 | 0.204
US50 | WAl | 30_15 | 0.45 | 35 | 0.206
US50 | WAl | 30_15 | 0.45 | 36 | 0.208
US50 | WAl | 30_15 | 0.45 | 37 | 0.21
US50 | WAl | 30_15 | 0.45 | 38 | 0.212
US50 | WAl | 30_15 | 0.45 | 39 | 0.214
US50 | WAl | 30_15 | 0.45 | 40 | 0.215
US50 | WAl | 30_15 | 0.45 | 41 | 0.217
US50 | WAl | 30_15 | 0.45 | 42 | 0.219
US50 | WAl | 30_15 | 0.45 | 43 | 0.22
US50 | WAl | 30_15 | 0.45 | 44 | 0.222
US50 | WAl | 30_15 | 0.5 | 21 | 0.188
US50 | WAl | 30_15 | 0.5 | 22 | 0.191
US50 | WAl | 30_15 | 0.5 | 23 | 0.195
US50 | WAl | 30_15 | 0.5 | 24 | 0.198
US50 | WAl | 30_15 | 0.5 | 25 | 0.201
US50 | WAl | 30_15 | 0.5 | 26 | 0.203
US50 | WAl | 30_15 | 0.5 | 27 | 0.206
US50 | WAl | 30_15 | 0.5 | 28 | 0.209
US50 | WAl | 30_15 | 0.5 | 29 | 0.212
US50 | WAl | 30_15 | 0.5 | 30 | 0.214
US50 | WAl | 30_15 | 0.5 | 31 | 0.216
US50 | WAl | 30_15 | 0.5 | 32 | 0.218
US50 | WAl | 30_15 | 0.5 | 33 | 0.221
US50 | WAl | 30_15 | 0.5 | 34 | 0.223
US50 | WAl | 30_15 | 0.5 | 35 | 0.224
US50 | WAl | 30_15 | 0.5 | 36 | 0.226
US50 | WAl | 30_15 | 0.5 | 37 | 0.228
US50 | WAl | 30_15 | 0.5 | 38 | 0.23
US50 | WAl | 30_15 | 0.5 | 39 | 0.231
US50 | WAl | 30_15 | 0.5 | 40 | 0.233
US50 | WAl | 30_15 | 0.5 | 41 | 0.235
US50 | WAl | 30_15 | 0.5 | 42 | 0.236
US50 | WAl | 30_15 | 0.5 | 43 | 0.237
US50 | WAl | 30_15 | 0.5 | 44 | 0.239
US50 | WAl | 30_15 | 0.5 | 45 | 0.24
US50 | WAl | 30_15 | 0.5 | 46 | 0.241
US50 | WAl | 30_15 | 0.5 | 47 | 0.243
US50 | WAl | 30_15 | 0.5 | 48 | 0.244
US50 | WAl | 30_15 | 0.5 | 49 | 0.245
US50 | WAl | 30_15 | 0.55 | 22 | 0.212
US50 | WAl | 30_15 | 0.55 | 23 | 0.215
US50 | WAl | 30_15 | 0.55 | 24 | 0.218
US50 | WAl | 30_15 | 0.55 | 25 | 0.22
US50 | WAl | 30_15 | 0.55 | 26 | 0.223
US50 | WAl | 30_15 | 0.55 | 27 | 0.225
US50 | WAl | 30_15 | 0.55 | 28 | 0.228
US50 | WAl | 30_15 | 0.55 | 29 | 0.23
US50 | WAl | 30_15 | 0.55 | 30 | 0.233
US50 | WAl | 30_15 | 0.55 | 31 | 0.235
US50 | WAl | 30_15 | 0.55 | 32 | 0.237
US50 | WAl | 30_15 | 0.55 | 33 | 0.239
US50 | WAl | 30_15 | 0.55 | 34 | 0.24
US50 | WAl | 30_15 | 0.55 | 35 | 0.242
US50 | WAl | 30_15 | 0.55 | 36 | 0.244
US50 | WAl | 30_15 | 0.55 | 37 | 0.245
US50 | WAl | 30_15 | 0.55 | 38 | 0.247
US50 | WAl | 30_15 | 0.55 | 39 | 0.248
US50 | WAl | 30_15 | 0.55 | 40 | 0.25
US50 | WAl | 30_15 | 0.55 | 41 | 0.251
US50 | WAl | 30_15 | 0.55 | 42 | 0.253
US50 | WAl | 30_15 | 0.55 | 43 | 0.254
US50 | WAl | 30_15 | 0.55 | 44 | 0.255
US50 | WAl | 30_15 | 0.55 | 45 | 0.257
US50 | WAl | 30_15 | 0.55 | 46 | 0.258
US50 | WAl | 30_15 | 0.55 | 47 | 0.259
US50 | WAl | 30_15 | 0.55 | 48 | 0.26
US50 | WAl | 30_15 | 0.55 | 49 | 0.261
US50 | WAl | 30_15 | 0.6 | 23 | 0.234
US50 | WAl | 30_15 | 0.6 | 24 | 0.237
US50 | WAl | 30_15 | 0.6 | 25 | 0.239
US50 | WAl | 30_15 | 0.6 | 26 | 0.241
US50 | WAl | 30_15 | 0.6 | 27 | 0.244
US50 | WAl | 30_15 | 0.6 | 28 | 0.246
US50 | WAl | 30_15 | 0.6 | 29 | 0.248
US50 | WAl | 30_15 | 0.6 | 30 | 0.25
US50 | WAl | 30_15 | 0.6 | 31 | 0.252
US50 | WAl | 30_15 | 0.6 | 32 | 0.254
US50 | WAl | 30_15 | 0.6 | 33 | 0.256
US50 | WAl | 30_15 | 0.6 | 34 | 0.258
US50 | WAl | 30_15 | 0.6 | 35 | 0.259
US50 | WAl | 30_15 | 0.6 | 36 | 0.261
US50 | WAl | 30_15 | 0.6 | 37 | 0.262
US50 | WAl | 30_15 | 0.6 | 38 | 0.264
US50 | WAl | 30_15 | 0.6 | 39 | 0.265
US50 | WAl | 30_15 | 0.6 | 40 | 0.266
US50 | WAl | 30_15 | 0.6 | 41 | 0.268
US50 | WAl | 30_15 | 0.6 | 42 | 0.269
US50 | WAl | 30_15 | 0.6 | 43 | 0.27
US50 | WAl | 30_15 | 0.6 | 44 | 0.271
US50 | WAl | 30_15 | 0.6 | 45 | 0.272
US50 | WAl | 30_15 | 0.6 | 46 | 0.273
US50 | WAl | 30_15 | 0.6 | 47 | 0.274
US50 | WAl | 30_15 | 0.6 | 48 | 0.276
US50 | WAl | 30_15 | 0.6 | 49 | 0.277
US50 | WAl | 30_15 | 0.65 | 24 | 0.255
US50 | WAl | 30_15 | 0.65 | 25 | 0.257
US50 | WAl | 30_15 | 0.65 | 26 | 0.259
US50 | WAl | 30_15 | 0.65 | 27 | 0.262
US50 | WAl | 30_15 | 0.65 | 28 | 0.264
US50 | WAl | 30_15 | 0.65 | 29 | 0.266
US50 | WAl | 30_15 | 0.65 | 30 | 0.268
US50 | WAl | 30_15 | 0.65 | 31 | 0.269
US50 | WAl | 30_15 | 0.65 | 32 | 0.271
US50 | WAl | 30_15 | 0.65 | 33 | 0.273
US50 | WAl | 30_15 | 0.65 | 34 | 0.274
US50 | WAl | 30_15 | 0.65 | 35 | 0.276
US50 | WAl | 30_15 | 0.65 | 36 | 0.277
US50 | WAl | 30_15 | 0.65 | 37 | 0.279
US50 | WAl | 30_15 | 0.65 | 38 | 0.28
US50 | WAl | 30_15 | 0.65 | 39 | 0.281
US50 | WAl | 30_15 | 0.65 | 40 | 0.282
US50 | WAl | 30_15 | 0.65 | 41 | 0.284
US50 | WAl | 30_15 | 0.65 | 42 | 0.285
US50 | WAl | 30_15 | 0.65 | 43 | 0.286
US50 | WAl | 30_15 | 0.65 | 44 | 0.287
US50 | WAl | 30_15 | 0.65 | 45 | 0.288
US50 | WAl | 30_15 | 0.65 | 46 | 0.289
US50 | WAl | 30_15 | 0.65 | 47 | 0.29
US50 | WAl | 30_15 | 0.65 | 48 | 0.291
US50 | WAl | 30_15 | 0.65 | 49 | 0.292
US50 | WAl | 30_15 | 0.7 | 25 | 0.275
US50 | WAl | 30_15 | 0.7 | 26 | 0.277
US50 | WAl | 30_15 | 0.7 | 27 | 0.279
US50 | WAl | 30_15 | 0.7 | 28 | 0.281
US50 | WAl | 30_15 | 0.7 | 29 | 0.283
US50 | WAl | 30_15 | 0.7 | 30 | 0.284
US50 | WAl | 30_15 | 0.7 | 31 | 0.286
US50 | WAl | 30_15 | 0.7 | 32 | 0.288
US50 | WAl | 30_15 | 0.7 | 33 | 0.289
US50 | WAl | 30_15 | 0.7 | 34 | 0.291
US50 | WAl | 30_15 | 0.7 | 35 | 0.292
US50 | WAl | 30_15 | 0.7 | 36 | 0.293
US50 | WAl | 30_15 | 0.7 | 37 | 0.294
US50 | WAl | 30_15 | 0.7 | 38 | 0.296
US50 | WAl | 30_15 | 0.7 | 39 | 0.297
US50 | WAl | 30_15 | 0.7 | 40 | 0.298
US50 | WAl | 30_15 | 0.7 | 41 | 0.299
US50 | WAl | 30_15 | 0.7 | 42 | 0.3
US50 | WAl | 30_15 | 0.7 | 43 | 0.301
US50 | WAl | 30_15 | 0.7 | 44 | 0.302
US50 | WAl | 30_15 | 0.7 | 45 | 0.303
US50 | WAl | 30_15 | 0.7 | 46 | 0.303
US50 | WAl | 30_15 | 0.7 | 47 | 0.304
US50 | WAl | 30_15 | 0.7 | 48 | 0.305
US50 | WAl | 30_15 | 0.7 | 49 | 0.306
US50 | WAl | 30_15 | 0.75 | 26 | 0.294
US50 | WAl | 30_15 | 0.75 | 27 | 0.295
US50 | WAl | 30_15 | 0.75 | 28 | 0.297
US50 | WAl | 30_15 | 0.75 | 29 | 0.299
US50 | WAl | 30_15 | 0.75 | 30 | 0.3
US50 | WAl | 30_15 | 0.75 | 31 | 0.302
US50 | WAl | 30_15 | 0.75 | 32 | 0.303
US50 | WAl | 30_15 | 0.75 | 33 | 0.305
US50 | WAl | 30_15 | 0.75 | 34 | 0.306
US50 | WAl | 30_15 | 0.75 | 35 | 0.307
US50 | WAl | 30_15 | 0.75 | 36 | 0.308
US50 | WAl | 30_15 | 0.75 | 37 | 0.31
US50 | WAl | 30_15 | 0.75 | 38 | 0.311
US50 | WAl | 30_15 | 0.75 | 39 | 0.312
US50 | WAl | 30_15 | 0.75 | 40 | 0.313
US50 | WAl | 30_15 | 0.75 | 41 | 0.313
US50 | WAl | 30_15 | 0.75 | 42 | 0.314
US50 | WAl | 30_15 | 0.75 | 43 | 0.315
US50 | WAl | 30_15 | 0.75 | 44 | 0.316
US50 | WAl | 30_15 | 0.75 | 45 | 0.317
US50 | WAl | 30_15 | 0.75 | 46 | 0.318
US50 | WAl | 30_15 | 0.75 | 47 | 0.319
US50 | WAl | 30_15 | 0.75 | 48 | 0.319
US50 | WAl | 30_15 | 0.75 | 49 | 0.32
US50 | WAl | 30_15 | 0.8 | 27 | 0.311
US50 | WAl | 30_15 | 0.8 | 28 | 0.313
US50 | WAl | 30_15 | 0.8 | 29 | 0.314
US50 | WAl | 30_15 | 0.8 | 30 | 0.316
US50 | WAl | 30_15 | 0.8 | 31 | 0.317
US50 | WAl | 30_15 | 0.8 | 32 | 0.318
US50 | WAl | 30_15 | 0.8 | 33 | 0.32
US50 | WAl | 30_15 | 0.8 | 34 | 0.321
US50 | WAl | 30_15 | 0.8 | 35 | 0.322
US50 | WAl | 30_15 | 0.8 | 36 | 0.323
US50 | WAl | 30_15 | 0.8 | 37 | 0.324
US50 | WAl | 30_15 | 0.8 | 38 | 0.325
US50 | WAl | 30_15 | 0.8 | 39 | 0.326
US50 | WAl | 30_15 | 0.8 | 40 | 0.327
US50 | WAl | 30_15 | 0.8 | 41 | 0.328
US50 | WAl | 30_15 | 0.8 | 42 | 0.328
US50 | WAl | 30_15 | 0.8 | 43 | 0.329
US50 | WAl | 30_15 | 0.8 | 44 | 0.33
US50 | WAl | 30_15 | 0.8 | 45 | 0.331
US50 | WAl | 30_15 | 0.8 | 46 | 0.331
US50 | WAl | 30_15 | 0.8 | 47 | 0.332
US50 | WAl | 30_15 | 0.8 | 48 | 0.333
US50 | WAl | 30_15 | 0.8 | 49 | 0.334
US50 | WAl | 30_15 | 0.85 | 28 | 0.328
US50 | WAl | 30_15 | 0.85 | 29 | 0.329
US50 | WAl | 30_15 | 0.85 | 30 | 0.331
US50 | WAl | 30_15 | 0.85 | 31 | 0.332
US50 | WAl | 30_15 | 0.85 | 32 | 0.333
US50 | WAl | 30_15 | 0.85 | 33 | 0.334
US50 | WAl | 30_15 | 0.85 | 34 | 0.335
US50 | WAl | 30_15 | 0.85 | 35 | 0.336
US50 | WAl | 30_15 | 0.85 | 36 | 0.337
US50 | WAl | 30_15 | 0.85 | 37 | 0.338
US50 | WAl | 30_15 | 0.85 | 38 | 0.339
US50 | WAl | 30_15 | 0.85 | 39 | 0.34
US50 | WAl | 30_15 | 0.85 | 40 | 0.341
US50 | WAl | 30_15 | 0.85 | 41 | 0.341
US50 | WAl | 30_15 | 0.85 | 42 | 0.342
US50 | WAl | 30_15 | 0.85 | 43 | 0.343
US50 | WAl | 30_15 | 0.85 | 44 | 0.343
US50 | WAl | 30_15 | 0.85 | 45 | 0.344
US50 | WAl | 30_15 | 0.85 | 46 | 0.345
US50 | WAl | 30_15 | 0.85 | 47 | 0.345
US50 | WAl | 30_15 | 0.85 | 48 | 0.346
US50 | WAl | 30_15 | 0.85 | 49 | 0.347
US50 | WAl | 30_15 | 0.9 | 28 | 0.343
US50 | WAl | 30_15 | 0.9 | 29 | 0.344
US50 | WAl | 30_15 | 0.9 | 30 | 0.345
US50 | WAl | 30_15 | 0.9 | 31 | 0.346
US50 | WAl | 30_15 | 0.9 | 32 | 0.347
US50 | WAl | 30_15 | 0.9 | 33 | 0.348
US50 | WAl | 30_15 | 0.9 | 34 | 0.349
US50 | WAl | 30_15 | 0.9 | 35 | 0.35
US50 | WAl | 30_15 | 0.9 | 36 | 0.351
US50 | WAl | 30_15 | 0.9 | 37 | 0.351
US50 | WAl | 30_15 | 0.9 | 38 | 0.352
US50 | WAl | 30_15 | 0.9 | 39 | 0.353
US50 | WAl | 30_15 | 0.9 | 40 | 0.354
US50 | WAl | 30_15 | 0.9 | 41 | 0.354
US50 | WAl | 30_15 | 0.9 | 42 | 0.355
US50 | WAl | 30_15 | 0.9 | 43 | 0.356
US50 | WAl | 30_15 | 0.9 | 44 | 0.356
US50 | WAl | 30_15 | 0.9 | 45 | 0.357
US50 | WAl | 30_15 | 0.9 | 46 | 0.357
US50 | WAl | 30_15 | 0.9 | 47 | 0.358
US50 | WAl | 30_15 | 0.9 | 48 | 0.359
US50 | WAl | 30_15 | 0.9 | 49 | 0.359
US50 | WAl | 30_15 | 0.95 | 29 | 0.357
US50 | WAl | 30_15 | 0.95 | 30 | 0.358
US50 | WAl | 30_15 | 0.95 | 31 | 0.359
US50 | WAl | 30_15 | 0.95 | 32 | 0.36
US50 | WAl | 30_15 | 0.95 | 33 | 0.361
US50 | WAl | 30_15 | 0.95 | 34 | 0.362
US50 | WAl | 30_15 | 0.95 | 35 | 0.363
US50 | WAl | 30_15 | 0.95 | 36 | 0.364
US50 | WAl | 30_15 | 0.95 | 37 | 0.364
US50 | WAl | 30_15 | 0.95 | 38 | 0.365
US50 | WAl | 30_15 | 0.95 | 39 | 0.366
US50 | WAl | 30_15 | 0.95 | 40 | 0.366
US50 | WAl | 30_15 | 0.95 | 41 | 0.367
US50 | WAl | 30_15 | 0.95 | 42 | 0.368
US50 | WAl | 30_15 | 0.95 | 43 | 0.368
US50 | WAl | 30_15 | 0.95 | 44 | 0.369
US50 | WAl | 30_15 | 0.95 | 45 | 0.369
US50 | WAl | 30_15 | 0.95 | 46 | 0.37
US50 | WAl | 30_15 | 0.95 | 47 | 0.37
US50 | WAl | 30_15 | 0.95 | 48 | 0.371
US50 | WAl | 30_15 | 0.95 | 49 | 0.371
US50 | WAl | 30_15 | 1 | 30 | 0.371
US50 | WAl | 30_15 | 1 | 31 | 0.372
US50 | WAl | 30_15 | 1 | 32 | 0.373
US50 | WAl | 30_15 | 1 | 33 | 0.374
US50 | WAl | 30_15 | 1 | 34 | 0.375
US50 | WAl | 30_15 | 1 | 35 | 0.375
US50 | WAl | 30_15 | 1 | 36 | 0.376
US50 | WAl | 30_15 | 1 | 37 | 0.377
US50 | WAl | 30_15 | 1 | 38 | 0.377
US50 | WAl | 30_15 | 1 | 39 | 0.378
US50 | WAl | 30_15 | 1 | 40 | 0.378
US50 | WAl | 30_15 | 1 | 41 | 0.379
US50 | WAl | 30_15 | 1 | 42 | 0.38
US50 | WAl | 30_15 | 1 | 43 | 0.38
US50 | WAl | 30_15 | 1 | 44 | 0.381
US50 | WAl | 30_15 | 1 | 45 | 0.381
US50 | WAl | 30_15 | 1 | 46 | 0.381
US50 | WAl | 30_15 | 1 | 47 | 0.382
US50 | WAl | 30_15 | 1 | 48 | 0.382
US50 | WAl | 30_15 | 1 | 49 | 0.383
US50 | WAl | 30_15 | 1.05 | 31 | 0.385
US50 | WAl | 30_15 | 1.05 | 32 | 0.385
US50 | WAl | 30_15 | 1.05 | 33 | 0.386
US50 | WAl | 30_15 | 1.05 | 34 | 0.387
US50 | WAl | 30_15 | 1.05 | 35 | 0.387
US50 | WAl | 30_15 | 1.05 | 36 | 0.388
US50 | WAl | 30_15 | 1.05 | 37 | 0.389
US50 | WAl | 30_15 | 1.05 | 38 | 0.389
US50 | WAl | 30_15 | 1.05 | 39 | 0.39
US50 | WAl | 30_15 | 1.05 | 40 | 0.39
US50 | WAl | 30_15 | 1.05 | 41 | 0.391
US50 | WAl | 30_15 | 1.05 | 42 | 0.391
US50 | WAl | 30_15 | 1.05 | 43 | 0.392
US50 | WAl | 30_15 | 1.05 | 44 | 0.392
US50 | WAl | 30_15 | 1.05 | 45 | 0.392
US50 | WAl | 30_15 | 1.05 | 46 | 0.393
US50 | WAl | 30_15 | 1.05 | 47 | 0.393
US50 | WAl | 30_15 | 1.05 | 48 | 0.394
US50 | WAl | 30_15 | 1.05 | 49 | 0.394
US50 | WAl | 30_15 | 1.1 | 32 | 0.397
US50 | WAl | 30_15 | 1.1 | 33 | 0.398
US50 | WAl | 30_15 | 1.1 | 34 | 0.398
US50 | WAl | 30_15 | 1.1 | 35 | 0.399
US50 | WAl | 30_15 | 1.1 | 36 | 0.399
US50 | WAl | 30_15 | 1.1 | 37 | 0.4
US50 | WAl | 30_15 | 1.1 | 38 | 0.4
US50 | WAl | 30_15 | 1.1 | 39 | 0.401
US50 | WAl | 30_15 | 1.1 | 40 | 0.401
US50 | WAl | 30_15 | 1.1 | 41 | 0.402
US50 | WAl | 30_15 | 1.1 | 42 | 0.402
US50 | WAl | 30_15 | 1.1 | 43 | 0.403
US50 | WAl | 30_15 | 1.1 | 44 | 0.403
US50 | WAl | 30_15 | 1.1 | 45 | 0.403
US50 | WAl | 30_15 | 1.1 | 46 | 0.404
US50 | WAl | 30_15 | 1.1 | 47 | 0.404
US50 | WAl | 30_15 | 1.1 | 48 | 0.404
US50 | WAl | 30_15 | 1.1 | 49 | 0.405
US50 | WAl | 30_15 | 1.15 | 33 | 0.409
US50 | WAl | 30_15 | 1.15 | 34 | 0.41
US50 | WAl | 30_15 | 1.15 | 35 | 0.41
US50 | WAl | 30_15 | 1.15 | 36 | 0.41
US50 | WAl | 30_15 | 1.15 | 37 | 0.411
US50 | WAl | 30_15 | 1.15 | 38 | 0.411
US50 | WAl | 30_15 | 1.15 | 39 | 0.412
US50 | WAl | 30_15 | 1.15 | 40 | 0.412
US50 | WAl | 30_15 | 1.15 | 41 | 0.413
US50 | WAl | 30_15 | 1.15 | 42 | 0.413
US50 | WAl | 30_15 | 1.15 | 43 | 0.413
US50 | WAl | 30_15 | 1.15 | 44 | 0.414
US50 | WAl | 30_15 | 1.15 | 45 | 0.414
US50 | WAl | 30_15 | 1.15 | 46 | 0.414
US50 | WAl | 30_15 | 1.15 | 47 | 0.414
US50 | WAl | 30_15 | 1.15 | 48 | 0.415
US50 | WAl | 30_15 | 1.15 | 49 | 0.415
US50 | WAl | 30_15 | 1.2 | 34 | 0.42
US50 | WAl | 30_15 | 1.2 | 35 | 0.421
US50 | WAl | 30_15 | 1.2 | 36 | 0.421
US50 | WAl | 30_15 | 1.2 | 37 | 0.421
US50 | WAl | 30_15 | 1.2 | 38 | 0.422
US50 | WAl | 30_15 | 1.2 | 39 | 0.422
US50 | WAl | 30_15 | 1.2 | 40 | 0.422
US50 | WAl | 30_15 | 1.2 | 41 | 0.423
US50 | WAl | 30_15 | 1.2 | 42 | 0.423
US50 | WAl | 30_15 | 1.2 | 43 | 0.423
US50 | WAl | 30_15 | 1.2 | 44 | 0.424
US50 | WAl | 30_15 | 1.2 | 45 | 0.424
US50 | WAl | 30_15 | 1.2 | 46 | 0.424
US50 | WAl | 30_15 | 1.2 | 47 | 0.424
US50 | WAl | 30_15 | 1.2 | 48 | 0.425
US50 | WAl | 30_15 | 1.2 | 49 | 0.425
US50 | WAl | 30_15 | 1.25 | 36 | 0.431
US50 | WAl | 30_15 | 1.25 | 37 | 0.432
US50 | WAl | 30_15 | 1.25 | 38 | 0.432
US50 | WAl | 30_15 | 1.25 | 39 | 0.432
US50 | WAl | 30_15 | 1.25 | 40 | 0.432
US50 | WAl | 30_15 | 1.25 | 41 | 0.433
US50 | WAl | 30_15 | 1.25 | 42 | 0.433
US50 | WAl | 30_15 | 1.25 | 43 | 0.433
US50 | WAl | 30_15 | 1.25 | 44 | 0.433
US50 | WAl | 30_15 | 1.25 | 45 | 0.434
US50 | WAl | 30_15 | 1.25 | 46 | 0.434
US50 | WAl | 30_15 | 1.25 | 47 | 0.434
US50 | WAl | 30_15 | 1.25 | 48 | 0.434
US50 | WAl | 30_15 | 1.25 | 49 | 0.435
US50 | WAl | 30_15 | 1.3 | 41 | 0.442
US50 | WAl | 30_15 | 1.3 | 42 | 0.443
US50 | WAl | 30_15 | 1.3 | 43 | 0.443
US50 | WAl | 30_15 | 1.3 | 44 | 0.443
US50 | WAl | 30_15 | 1.3 | 45 | 0.443
US50 | WAl | 30_15 | 1.3 | 46 | 0.443
US50 | WAl | 30_15 | 1.3 | 47 | 0.443
US50 | WAl | 30_15 | 1.3 | 48 | 0.444
US50 | WAl | 30_15 | 1.3 | 49 | 0.444
US50 | WAl | 30_15 | 1.35 | 45 | 0.452
US50 | WAl | 30_15 | 1.35 | 46 | 0.452
US50 | WAl | 30_15 | 1.35 | 47 | 0.452
US50 | WAl | 30_15 | 1.35 | 48 | 0.453
US50 | WAl | 30_15 | 1.35 | 49 | 0.453
US50 | WAl | 25_50 | 0.2 | 20 | 0.055
US50 | WAl | 25_50 | 0.2 | 21 | 0.058
US50 | WAl | 25_50 | 0.25 | 20 | 0.075
US50 | WAl | 25_50 | 0.25 | 21 | 0.079
US50 | WAl | 25_50 | 0.25 | 22 | 0.083
US50 | WAl | 25_50 | 0.25 | 23 | 0.087
US50 | WAl | 25_50 | 0.25 | 24 | 0.091
US50 | WAl | 25_50 | 0.25 | 25 | 0.094
US50 | WAl | 25_50 | 0.3 | 20 | 0.095
US50 | WAl | 25_50 | 0.3 | 21 | 0.099
US50 | WAl | 25_50 | 0.3 | 22 | 0.104
US50 | WAl | 25_50 | 0.3 | 23 | 0.108
US50 | WAl | 25_50 | 0.3 | 24 | 0.112
US50 | WAl | 25_50 | 0.3 | 25 | 0.116
US50 | WAl | 25_50 | 0.3 | 26 | 0.119
US50 | WAl | 25_50 | 0.3 | 27 | 0.122
US50 | WAl | 25_50 | 0.3 | 28 | 0.126
US50 | WAl | 25_50 | 0.3 | 29 | 0.129
US50 | WAl | 25_50 | 0.3 | 30 | 0.132
US50 | WAl | 25_50 | 0.35 | 20 | 0.115
US50 | WAl | 25_50 | 0.35 | 21 | 0.12
US50 | WAl | 25_50 | 0.35 | 22 | 0.124
US50 | WAl | 25_50 | 0.35 | 23 | 0.129
US50 | WAl | 25_50 | 0.35 | 24 | 0.133
US50 | WAl | 25_50 | 0.35 | 25 | 0.136
US50 | WAl | 25_50 | 0.35 | 26 | 0.14
US50 | WAl | 25_50 | 0.35 | 27 | 0.143
US50 | WAl | 25_50 | 0.35 | 28 | 0.146
US50 | WAl | 25_50 | 0.35 | 29 | 0.15
US50 | WAl | 25_50 | 0.35 | 30 | 0.152
US50 | WAl | 25_50 | 0.35 | 31 | 0.155
US50 | WAl | 25_50 | 0.35 | 32 | 0.158
US50 | WAl | 25_50 | 0.35 | 33 | 0.16
US50 | WAl | 25_50 | 0.35 | 34 | 0.163
US50 | WAl | 25_50 | 0.35 | 35 | 0.165
US50 | WAl | 25_50 | 0.4 | 20 | 0.137
US50 | WAl | 25_50 | 0.4 | 21 | 0.141
US50 | WAl | 25_50 | 0.4 | 22 | 0.145
US50 | WAl | 25_50 | 0.4 | 23 | 0.149
US50 | WAl | 25_50 | 0.4 | 24 | 0.153
US50 | WAl | 25_50 | 0.4 | 25 | 0.157
US50 | WAl | 25_50 | 0.4 | 26 | 0.16
US50 | WAl | 25_50 | 0.4 | 27 | 0.163
US50 | WAl | 25_50 | 0.4 | 28 | 0.167
US50 | WAl | 25_50 | 0.4 | 29 | 0.17
US50 | WAl | 25_50 | 0.4 | 30 | 0.172
US50 | WAl | 25_50 | 0.4 | 31 | 0.175
US50 | WAl | 25_50 | 0.4 | 32 | 0.178
US50 | WAl | 25_50 | 0.4 | 33 | 0.18
US50 | WAl | 25_50 | 0.4 | 34 | 0.183
US50 | WAl | 25_50 | 0.4 | 35 | 0.185
US50 | WAl | 25_50 | 0.4 | 36 | 0.187
US50 | WAl | 25_50 | 0.4 | 37 | 0.189
US50 | WAl | 25_50 | 0.4 | 38 | 0.191
US50 | WAl | 25_50 | 0.4 | 39 | 0.193
US50 | WAl | 25_50 | 0.4 | 40 | 0.195
US50 | WAl | 25_50 | 0.45 | 20 | 0.158
US50 | WAl | 25_50 | 0.45 | 21 | 0.162
US50 | WAl | 25_50 | 0.45 | 22 | 0.166
US50 | WAl | 25_50 | 0.45 | 23 | 0.17
US50 | WAl | 25_50 | 0.45 | 24 | 0.174
US50 | WAl | 25_50 | 0.45 | 25 | 0.177
US50 | WAl | 25_50 | 0.45 | 26 | 0.18
US50 | WAl | 25_50 | 0.45 | 27 | 0.183
US50 | WAl | 25_50 | 0.45 | 28 | 0.186
US50 | WAl | 25_50 | 0.45 | 29 | 0.189
US50 | WAl | 25_50 | 0.45 | 30 | 0.192
US50 | WAl | 25_50 | 0.45 | 31 | 0.194
US50 | WAl | 25_50 | 0.45 | 32 | 0.197
US50 | WAl | 25_50 | 0.45 | 33 | 0.199
US50 | WAl | 25_50 | 0.45 | 34 | 0.201
US50 | WAl | 25_50 | 0.45 | 35 | 0.203
US50 | WAl | 25_50 | 0.45 | 36 | 0.205
US50 | WAl | 25_50 | 0.45 | 37 | 0.207
US50 | WAl | 25_50 | 0.45 | 38 | 0.209
US50 | WAl | 25_50 | 0.45 | 39 | 0.211
US50 | WAl | 25_50 | 0.45 | 40 | 0.213
US50 | WAl | 25_50 | 0.45 | 41 | 0.215
US50 | WAl | 25_50 | 0.45 | 42 | 0.216
US50 | WAl | 25_50 | 0.45 | 43 | 0.218
US50 | WAl | 25_50 | 0.45 | 44 | 0.219
US50 | WAl | 25_50 | 0.5 | 21 | 0.183
US50 | WAl | 25_50 | 0.5 | 22 | 0.187
US50 | WAl | 25_50 | 0.5 | 23 | 0.19
US50 | WAl | 25_50 | 0.5 | 24 | 0.194
US50 | WAl | 25_50 | 0.5 | 25 | 0.197
US50 | WAl | 25_50 | 0.5 | 26 | 0.2
US50 | WAl | 25_50 | 0.5 | 27 | 0.203
US50 | WAl | 25_50 | 0.5 | 28 | 0.205
US50 | WAl | 25_50 | 0.5 | 29 | 0.208
US50 | WAl | 25_50 | 0.5 | 30 | 0.211
US50 | WAl | 25_50 | 0.5 | 31 | 0.213
US50 | WAl | 25_50 | 0.5 | 32 | 0.215
US50 | WAl | 25_50 | 0.5 | 33 | 0.217
US50 | WAl | 25_50 | 0.5 | 34 | 0.22
US50 | WAl | 25_50 | 0.5 | 35 | 0.221
US50 | WAl | 25_50 | 0.5 | 36 | 0.223
US50 | WAl | 25_50 | 0.5 | 37 | 0.225
US50 | WAl | 25_50 | 0.5 | 38 | 0.227
US50 | WAl | 25_50 | 0.5 | 39 | 0.229
US50 | WAl | 25_50 | 0.5 | 40 | 0.23
US50 | WAl | 25_50 | 0.5 | 41 | 0.232
US50 | WAl | 25_50 | 0.5 | 42 | 0.234
US50 | WAl | 25_50 | 0.5 | 43 | 0.235
US50 | WAl | 25_50 | 0.5 | 44 | 0.236
US50 | WAl | 25_50 | 0.5 | 45 | 0.238
US50 | WAl | 25_50 | 0.5 | 46 | 0.239
US50 | WAl | 25_50 | 0.5 | 47 | 0.24
US50 | WAl | 25_50 | 0.5 | 48 | 0.242
US50 | WAl | 25_50 | 0.5 | 49 | 0.243
US50 | WAl | 25_50 | 0.55 | 22 | 0.207
US50 | WAl | 25_50 | 0.55 | 23 | 0.21
US50 | WAl | 25_50 | 0.55 | 24 | 0.213
US50 | WAl | 25_50 | 0.55 | 25 | 0.216
US50 | WAl | 25_50 | 0.55 | 26 | 0.219
US50 | WAl | 25_50 | 0.55 | 27 | 0.222
US50 | WAl | 25_50 | 0.55 | 28 | 0.224
US50 | WAl | 25_50 | 0.55 | 29 | 0.227
US50 | WAl | 25_50 | 0.55 | 30 | 0.229
US50 | WAl | 25_50 | 0.55 | 31 | 0.231
US50 | WAl | 25_50 | 0.55 | 32 | 0.233
US50 | WAl | 25_50 | 0.55 | 33 | 0.235
US50 | WAl | 25_50 | 0.55 | 34 | 0.237
US50 | WAl | 25_50 | 0.55 | 35 | 0.239
US50 | WAl | 25_50 | 0.55 | 36 | 0.241
US50 | WAl | 25_50 | 0.55 | 37 | 0.243
US50 | WAl | 25_50 | 0.55 | 38 | 0.244
US50 | WAl | 25_50 | 0.55 | 39 | 0.246
US50 | WAl | 25_50 | 0.55 | 40 | 0.247
US50 | WAl | 25_50 | 0.55 | 41 | 0.249
US50 | WAl | 25_50 | 0.55 | 42 | 0.25
US50 | WAl | 25_50 | 0.55 | 43 | 0.252
US50 | WAl | 25_50 | 0.55 | 44 | 0.253
US50 | WAl | 25_50 | 0.55 | 45 | 0.254
US50 | WAl | 25_50 | 0.55 | 46 | 0.255
US50 | WAl | 25_50 | 0.55 | 47 | 0.257
US50 | WAl | 25_50 | 0.55 | 48 | 0.258
US50 | WAl | 25_50 | 0.55 | 49 | 0.259
US50 | WAl | 25_50 | 0.6 | 23 | 0.23
US50 | WAl | 25_50 | 0.6 | 24 | 0.233
US50 | WAl | 25_50 | 0.6 | 25 | 0.235
US50 | WAl | 25_50 | 0.6 | 26 | 0.238
US50 | WAl | 25_50 | 0.6 | 27 | 0.24
US50 | WAl | 25_50 | 0.6 | 28 | 0.243
US50 | WAl | 25_50 | 0.6 | 29 | 0.245
US50 | WAl | 25_50 | 0.6 | 30 | 0.247
US50 | WAl | 25_50 | 0.6 | 31 | 0.249
US50 | WAl | 25_50 | 0.6 | 32 | 0.251
US50 | WAl | 25_50 | 0.6 | 33 | 0.253
US50 | WAl | 25_50 | 0.6 | 34 | 0.255
US50 | WAl | 25_50 | 0.6 | 35 | 0.256
US50 | WAl | 25_50 | 0.6 | 36 | 0.258
US50 | WAl | 25_50 | 0.6 | 37 | 0.259
US50 | WAl | 25_50 | 0.6 | 38 | 0.261
US50 | WAl | 25_50 | 0.6 | 39 | 0.262
US50 | WAl | 25_50 | 0.6 | 40 | 0.264
US50 | WAl | 25_50 | 0.6 | 41 | 0.265
US50 | WAl | 25_50 | 0.6 | 42 | 0.266
US50 | WAl | 25_50 | 0.6 | 43 | 0.268
US50 | WAl | 25_50 | 0.6 | 44 | 0.269
US50 | WAl | 25_50 | 0.6 | 45 | 0.27
US50 | WAl | 25_50 | 0.6 | 46 | 0.271
US50 | WAl | 25_50 | 0.6 | 47 | 0.272
US50 | WAl | 25_50 | 0.6 | 48 | 0.273
US50 | WAl | 25_50 | 0.6 | 49 | 0.274
US50 | WAl | 25_50 | 0.65 | 24 | 0.251
US50 | WAl | 25_50 | 0.65 | 25 | 0.253
US50 | WAl | 25_50 | 0.65 | 26 | 0.256
US50 | WAl | 25_50 | 0.65 | 27 | 0.258
US50 | WAl | 25_50 | 0.65 | 28 | 0.26
US50 | WAl | 25_50 | 0.65 | 29 | 0.262
US50 | WAl | 25_50 | 0.65 | 30 | 0.264
US50 | WAl | 25_50 | 0.65 | 31 | 0.266
US50 | WAl | 25_50 | 0.65 | 32 | 0.268
US50 | WAl | 25_50 | 0.65 | 33 | 0.27
US50 | WAl | 25_50 | 0.65 | 34 | 0.271
US50 | WAl | 25_50 | 0.65 | 35 | 0.273
US50 | WAl | 25_50 | 0.65 | 36 | 0.274
US50 | WAl | 25_50 | 0.65 | 37 | 0.276
US50 | WAl | 25_50 | 0.65 | 38 | 0.277
US50 | WAl | 25_50 | 0.65 | 39 | 0.279
US50 | WAl | 25_50 | 0.65 | 40 | 0.28
US50 | WAl | 25_50 | 0.65 | 41 | 0.281
US50 | WAl | 25_50 | 0.65 | 42 | 0.282
US50 | WAl | 25_50 | 0.65 | 43 | 0.283
US50 | WAl | 25_50 | 0.65 | 44 | 0.284
US50 | WAl | 25_50 | 0.65 | 45 | 0.285
US50 | WAl | 25_50 | 0.65 | 46 | 0.287
US50 | WAl | 25_50 | 0.65 | 47 | 0.287
US50 | WAl | 25_50 | 0.65 | 48 | 0.288
US50 | WAl | 25_50 | 0.65 | 49 | 0.289
US50 | WAl | 25_50 | 0.7 | 25 | 0.271
US50 | WAl | 25_50 | 0.7 | 26 | 0.273
US50 | WAl | 25_50 | 0.7 | 27 | 0.275
US50 | WAl | 25_50 | 0.7 | 28 | 0.277
US50 | WAl | 25_50 | 0.7 | 29 | 0.279
US50 | WAl | 25_50 | 0.7 | 30 | 0.281
US50 | WAl | 25_50 | 0.7 | 31 | 0.283
US50 | WAl | 25_50 | 0.7 | 32 | 0.284
US50 | WAl | 25_50 | 0.7 | 33 | 0.286
US50 | WAl | 25_50 | 0.7 | 34 | 0.288
US50 | WAl | 25_50 | 0.7 | 35 | 0.289
US50 | WAl | 25_50 | 0.7 | 36 | 0.29
US50 | WAl | 25_50 | 0.7 | 37 | 0.292
US50 | WAl | 25_50 | 0.7 | 38 | 0.293
US50 | WAl | 25_50 | 0.7 | 39 | 0.294
US50 | WAl | 25_50 | 0.7 | 40 | 0.295
US50 | WAl | 25_50 | 0.7 | 41 | 0.296
US50 | WAl | 25_50 | 0.7 | 42 | 0.297
US50 | WAl | 25_50 | 0.7 | 43 | 0.298
US50 | WAl | 25_50 | 0.7 | 44 | 0.299
US50 | WAl | 25_50 | 0.7 | 45 | 0.3
US50 | WAl | 25_50 | 0.7 | 46 | 0.301
US50 | WAl | 25_50 | 0.7 | 47 | 0.302
US50 | WAl | 25_50 | 0.7 | 48 | 0.303
US50 | WAl | 25_50 | 0.7 | 49 | 0.304
US50 | WAl | 25_50 | 0.75 | 26 | 0.29
US50 | WAl | 25_50 | 0.75 | 27 | 0.292
US50 | WAl | 25_50 | 0.75 | 28 | 0.294
US50 | WAl | 25_50 | 0.75 | 29 | 0.296
US50 | WAl | 25_50 | 0.75 | 30 | 0.297
US50 | WAl | 25_50 | 0.75 | 31 | 0.299
US50 | WAl | 25_50 | 0.75 | 32 | 0.3
US50 | WAl | 25_50 | 0.75 | 33 | 0.302
US50 | WAl | 25_50 | 0.75 | 34 | 0.303
US50 | WAl | 25_50 | 0.75 | 35 | 0.304
US50 | WAl | 25_50 | 0.75 | 36 | 0.306
US50 | WAl | 25_50 | 0.75 | 37 | 0.307
US50 | WAl | 25_50 | 0.75 | 38 | 0.308
US50 | WAl | 25_50 | 0.75 | 39 | 0.309
US50 | WAl | 25_50 | 0.75 | 40 | 0.31
US50 | WAl | 25_50 | 0.75 | 41 | 0.311
US50 | WAl | 25_50 | 0.75 | 42 | 0.312
US50 | WAl | 25_50 | 0.75 | 43 | 0.313
US50 | WAl | 25_50 | 0.75 | 44 | 0.314
US50 | WAl | 25_50 | 0.75 | 45 | 0.315
US50 | WAl | 25_50 | 0.75 | 46 | 0.316
US50 | WAl | 25_50 | 0.75 | 47 | 0.316
US50 | WAl | 25_50 | 0.75 | 48 | 0.317
US50 | WAl | 25_50 | 0.75 | 49 | 0.318
US50 | WAl | 25_50 | 0.8 | 27 | 0.308
US50 | WAl | 25_50 | 0.8 | 28 | 0.31
US50 | WAl | 25_50 | 0.8 | 29 | 0.311
US50 | WAl | 25_50 | 0.8 | 30 | 0.313
US50 | WAl | 25_50 | 0.8 | 31 | 0.314
US50 | WAl | 25_50 | 0.8 | 32 | 0.315
US50 | WAl | 25_50 | 0.8 | 33 | 0.317
US50 | WAl | 25_50 | 0.8 | 34 | 0.318
US50 | WAl | 25_50 | 0.8 | 35 | 0.319
US50 | WAl | 25_50 | 0.8 | 36 | 0.32
US50 | WAl | 25_50 | 0.8 | 37 | 0.321
US50 | WAl | 25_50 | 0.8 | 38 | 0.323
US50 | WAl | 25_50 | 0.8 | 39 | 0.324
US50 | WAl | 25_50 | 0.8 | 40 | 0.324
US50 | WAl | 25_50 | 0.8 | 41 | 0.325
US50 | WAl | 25_50 | 0.8 | 42 | 0.326
US50 | WAl | 25_50 | 0.8 | 43 | 0.327
US50 | WAl | 25_50 | 0.8 | 44 | 0.328
US50 | WAl | 25_50 | 0.8 | 45 | 0.329
US50 | WAl | 25_50 | 0.8 | 46 | 0.329
US50 | WAl | 25_50 | 0.8 | 47 | 0.33
US50 | WAl | 25_50 | 0.8 | 48 | 0.331
US50 | WAl | 25_50 | 0.8 | 49 | 0.332
US50 | WAl | 25_50 | 0.85 | 28 | 0.325
US50 | WAl | 25_50 | 0.85 | 29 | 0.326
US50 | WAl | 25_50 | 0.85 | 30 | 0.328
US50 | WAl | 25_50 | 0.85 | 31 | 0.329
US50 | WAl | 25_50 | 0.85 | 32 | 0.33
US50 | WAl | 25_50 | 0.85 | 33 | 0.331
US50 | WAl | 25_50 | 0.85 | 34 | 0.333
US50 | WAl | 25_50 | 0.85 | 35 | 0.334
US50 | WAl | 25_50 | 0.85 | 36 | 0.335
US50 | WAl | 25_50 | 0.85 | 37 | 0.336
US50 | WAl | 25_50 | 0.85 | 38 | 0.337
US50 | WAl | 25_50 | 0.85 | 39 | 0.337
US50 | WAl | 25_50 | 0.85 | 40 | 0.338
US50 | WAl | 25_50 | 0.85 | 41 | 0.339
US50 | WAl | 25_50 | 0.85 | 42 | 0.34
US50 | WAl | 25_50 | 0.85 | 43 | 0.341
US50 | WAl | 25_50 | 0.85 | 44 | 0.341
US50 | WAl | 25_50 | 0.85 | 45 | 0.342
US50 | WAl | 25_50 | 0.85 | 46 | 0.343
US50 | WAl | 25_50 | 0.85 | 47 | 0.343
US50 | WAl | 25_50 | 0.85 | 48 | 0.344
US50 | WAl | 25_50 | 0.85 | 49 | 0.345
US50 | WAl | 25_50 | 0.9 | 28 | 0.339
US50 | WAl | 25_50 | 0.9 | 29 | 0.341
US50 | WAl | 25_50 | 0.9 | 30 | 0.342
US50 | WAl | 25_50 | 0.9 | 31 | 0.343
US50 | WAl | 25_50 | 0.9 | 32 | 0.344
US50 | WAl | 25_50 | 0.9 | 33 | 0.345
US50 | WAl | 25_50 | 0.9 | 34 | 0.346
US50 | WAl | 25_50 | 0.9 | 35 | 0.347
US50 | WAl | 25_50 | 0.9 | 36 | 0.348
US50 | WAl | 25_50 | 0.9 | 37 | 0.349
US50 | WAl | 25_50 | 0.9 | 38 | 0.35
US50 | WAl | 25_50 | 0.9 | 39 | 0.351
US50 | WAl | 25_50 | 0.9 | 40 | 0.352
US50 | WAl | 25_50 | 0.9 | 41 | 0.352
US50 | WAl | 25_50 | 0.9 | 42 | 0.353
US50 | WAl | 25_50 | 0.9 | 43 | 0.354
US50 | WAl | 25_50 | 0.9 | 44 | 0.354
US50 | WAl | 25_50 | 0.9 | 45 | 0.355
US50 | WAl | 25_50 | 0.9 | 46 | 0.356
US50 | WAl | 25_50 | 0.9 | 47 | 0.356
US50 | WAl | 25_50 | 0.9 | 48 | 0.357
US50 | WAl | 25_50 | 0.9 | 49 | 0.357
US50 | WAl | 25_50 | 0.95 | 29 | 0.355
US50 | WAl | 25_50 | 0.95 | 30 | 0.356
US50 | WAl | 25_50 | 0.95 | 31 | 0.357
US50 | WAl | 25_50 | 0.95 | 32 | 0.358
US50 | WAl | 25_50 | 0.95 | 33 | 0.359
US50 | WAl | 25_50 | 0.95 | 34 | 0.36
US50 | WAl | 25_50 | 0.95 | 35 | 0.361
US50 | WAl | 25_50 | 0.95 | 36 | 0.361
US50 | WAl | 25_50 | 0.95 | 37 | 0.362
US50 | WAl | 25_50 | 0.95 | 38 | 0.363
US50 | WAl | 25_50 | 0.95 | 39 | 0.364
US50 | WAl | 25_50 | 0.95 | 40 | 0.364
US50 | WAl | 25_50 | 0.95 | 41 | 0.365
US50 | WAl | 25_50 | 0.95 | 42 | 0.366
US50 | WAl | 25_50 | 0.95 | 43 | 0.366
US50 | WAl | 25_50 | 0.95 | 44 | 0.367
US50 | WAl | 25_50 | 0.95 | 45 | 0.367
US50 | WAl | 25_50 | 0.95 | 46 | 0.368
US50 | WAl | 25_50 | 0.95 | 47 | 0.369
US50 | WAl | 25_50 | 0.95 | 48 | 0.369
US50 | WAl | 25_50 | 0.95 | 49 | 0.37
US50 | WAl | 25_50 | 1 | 30 | 0.369
US50 | WAl | 25_50 | 1 | 31 | 0.37
US50 | WAl | 25_50 | 1 | 32 | 0.371
US50 | WAl | 25_50 | 1 | 33 | 0.372
US50 | WAl | 25_50 | 1 | 34 | 0.372
US50 | WAl | 25_50 | 1 | 35 | 0.373
US50 | WAl | 25_50 | 1 | 36 | 0.374
US50 | WAl | 25_50 | 1 | 37 | 0.375
US50 | WAl | 25_50 | 1 | 38 | 0.375
US50 | WAl | 25_50 | 1 | 39 | 0.376
US50 | WAl | 25_50 | 1 | 40 | 0.377
US50 | WAl | 25_50 | 1 | 41 | 0.377
US50 | WAl | 25_50 | 1 | 42 | 0.378
US50 | WAl | 25_50 | 1 | 43 | 0.378
US50 | WAl | 25_50 | 1 | 44 | 0.379
US50 | WAl | 25_50 | 1 | 45 | 0.379
US50 | WAl | 25_50 | 1 | 46 | 0.38
US50 | WAl | 25_50 | 1 | 47 | 0.38
US50 | WAl | 25_50 | 1 | 48 | 0.381
US50 | WAl | 25_50 | 1 | 49 | 0.381
US50 | WAl | 25_50 | 1.05 | 31 | 0.382
US50 | WAl | 25_50 | 1.05 | 32 | 0.383
US50 | WAl | 25_50 | 1.05 | 33 | 0.384
US50 | WAl | 25_50 | 1.05 | 34 | 0.385
US50 | WAl | 25_50 | 1.05 | 35 | 0.385
US50 | WAl | 25_50 | 1.05 | 36 | 0.386
US50 | WAl | 25_50 | 1.05 | 37 | 0.387
US50 | WAl | 25_50 | 1.05 | 38 | 0.387
US50 | WAl | 25_50 | 1.05 | 39 | 0.388
US50 | WAl | 25_50 | 1.05 | 40 | 0.388
US50 | WAl | 25_50 | 1.05 | 41 | 0.389
US50 | WAl | 25_50 | 1.05 | 42 | 0.389
US50 | WAl | 25_50 | 1.05 | 43 | 0.39
US50 | WAl | 25_50 | 1.05 | 44 | 0.391
US50 | WAl | 25_50 | 1.05 | 45 | 0.391
US50 | WAl | 25_50 | 1.05 | 46 | 0.391
US50 | WAl | 25_50 | 1.05 | 47 | 0.392
US50 | WAl | 25_50 | 1.05 | 48 | 0.392
US50 | WAl | 25_50 | 1.05 | 49 | 0.393
US50 | WAl | 25_50 | 1.1 | 32 | 0.395
US50 | WAl | 25_50 | 1.1 | 33 | 0.396
US50 | WAl | 25_50 | 1.1 | 34 | 0.396
US50 | WAl | 25_50 | 1.1 | 35 | 0.397
US50 | WAl | 25_50 | 1.1 | 36 | 0.398
US50 | WAl | 25_50 | 1.1 | 37 | 0.398
US50 | WAl | 25_50 | 1.1 | 38 | 0.399
US50 | WAl | 25_50 | 1.1 | 39 | 0.399
US50 | WAl | 25_50 | 1.1 | 40 | 0.4
US50 | WAl | 25_50 | 1.1 | 41 | 0.4
US50 | WAl | 25_50 | 1.1 | 42 | 0.401
US50 | WAl | 25_50 | 1.1 | 43 | 0.401
US50 | WAl | 25_50 | 1.1 | 44 | 0.402
US50 | WAl | 25_50 | 1.1 | 45 | 0.402
US50 | WAl | 25_50 | 1.1 | 46 | 0.402
US50 | WAl | 25_50 | 1.1 | 47 | 0.403
US50 | WAl | 25_50 | 1.1 | 48 | 0.403
US50 | WAl | 25_50 | 1.1 | 49 | 0.404
US50 | WAl | 25_50 | 1.15 | 33 | 0.407
US50 | WAl | 25_50 | 1.15 | 34 | 0.408
US50 | WAl | 25_50 | 1.15 | 35 | 0.408
US50 | WAl | 25_50 | 1.15 | 36 | 0.409
US50 | WAl | 25_50 | 1.15 | 37 | 0.409
US50 | WAl | 25_50 | 1.15 | 38 | 0.41
US50 | WAl | 25_50 | 1.15 | 39 | 0.41
US50 | WAl | 25_50 | 1.15 | 40 | 0.411
US50 | WAl | 25_50 | 1.15 | 41 | 0.411
US50 | WAl | 25_50 | 1.15 | 42 | 0.411
US50 | WAl | 25_50 | 1.15 | 43 | 0.412
US50 | WAl | 25_50 | 1.15 | 44 | 0.412
US50 | WAl | 25_50 | 1.15 | 45 | 0.413
US50 | WAl | 25_50 | 1.15 | 46 | 0.413
US50 | WAl | 25_50 | 1.15 | 47 | 0.413
US50 | WAl | 25_50 | 1.15 | 48 | 0.414
US50 | WAl | 25_50 | 1.15 | 49 | 0.414
US50 | WAl | 25_50 | 1.2 | 34 | 0.419
US50 | WAl | 25_50 | 1.2 | 35 | 0.419
US50 | WAl | 25_50 | 1.2 | 36 | 0.42
US50 | WAl | 25_50 | 1.2 | 37 | 0.42
US50 | WAl | 25_50 | 1.2 | 38 | 0.42
US50 | WAl | 25_50 | 1.2 | 39 | 0.421
US50 | WAl | 25_50 | 1.2 | 40 | 0.421
US50 | WAl | 25_50 | 1.2 | 41 | 0.422
US50 | WAl | 25_50 | 1.2 | 42 | 0.422
US50 | WAl | 25_50 | 1.2 | 43 | 0.422
US50 | WAl | 25_50 | 1.2 | 44 | 0.423
US50 | WAl | 25_50 | 1.2 | 45 | 0.423
US50 | WAl | 25_50 | 1.2 | 46 | 0.423
US50 | WAl | 25_50 | 1.2 | 47 | 0.423
US50 | WAl | 25_50 | 1.2 | 48 | 0.424
US50 | WAl | 25_50 | 1.2 | 49 | 0.424
US50 | WAl | 25_50 | 1.25 | 36 | 0.43
US50 | WAl | 25_50 | 1.25 | 37 | 0.43
US50 | WAl | 25_50 | 1.25 | 38 | 0.431
US50 | WAl | 25_50 | 1.25 | 39 | 0.431
US50 | WAl | 25_50 | 1.25 | 40 | 0.431
US50 | WAl | 25_50 | 1.25 | 41 | 0.432
US50 | WAl | 25_50 | 1.25 | 42 | 0.432
US50 | WAl | 25_50 | 1.25 | 43 | 0.432
US50 | WAl | 25_50 | 1.25 | 44 | 0.433
US50 | WAl | 25_50 | 1.25 | 45 | 0.433
US50 | WAl | 25_50 | 1.25 | 46 | 0.433
US50 | WAl | 25_50 | 1.25 | 47 | 0.433
US50 | WAl | 25_50 | 1.25 | 48 | 0.433
US50 | WAl | 25_50 | 1.25 | 49 | 0.434
US50 | WAl | 25_50 | 1.3 | 41 | 0.441
US50 | WAl | 25_50 | 1.3 | 42 | 0.442
US50 | WAl | 25_50 | 1.3 | 43 | 0.442
US50 | WAl | 25_50 | 1.3 | 44 | 0.442
US50 | WAl | 25_50 | 1.3 | 45 | 0.442
US50 | WAl | 25_50 | 1.3 | 46 | 0.443
US50 | WAl | 25_50 | 1.3 | 47 | 0.443
US50 | WAl | 25_50 | 1.3 | 48 | 0.443
US50 | WAl | 25_50 | 1.3 | 49 | 0.443
US50 | WAl | 25_50 | 1.35 | 45 | 0.451
US50 | WAl | 25_50 | 1.35 | 46 | 0.452
US50 | WAl | 25_50 | 1.35 | 47 | 0.452
US50 | WAl | 25_50 | 1.35 | 48 | 0.452
US50 | WAl | 25_50 | 1.35 | 49 | 0.452
US50 | WRh | 25_50 | 0.2 | 20 | 0.055
US50 | WRh | 25_50 | 0.2 | 21 | 0.059
US50 | WRh | 25_50 | 0.225 | 20 | 0.065
US50 | WRh | 25_50 | 0.225 | 21 | 0.069
US50 | WRh | 25_50 | 0.225 | 22 | 0.073
US50 | WRh | 25_50 | 0.225 | 23 | 0.077
US50 | WRh | 25_50 | 0.225 | 24 | 0.08
US50 | WRh | 25_50 | 0.25 | 20 | 0.075
US50 | WRh | 25_50 | 0.25 | 21 | 0.08
US50 | WRh | 25_50 | 0.25 | 22 | 0.084
US50 | WRh | 25_50 | 0.25 | 23 | 0.088
US50 | WRh | 25_50 | 0.25 | 24 | 0.092
US50 | WRh | 25_50 | 0.25 | 25 | 0.095
US50 | WRh | 25_50 | 0.25 | 26 | 0.098
US50 | WRh | 25_50 | 0.275 | 20 | 0.085
US50 | WRh | 25_50 | 0.275 | 21 | 0.09
US50 | WRh | 25_50 | 0.275 | 22 | 0.094
US50 | WRh | 25_50 | 0.275 | 23 | 0.098
US50 | WRh | 25_50 | 0.275 | 24 | 0.102
US50 | WRh | 25_50 | 0.275 | 25 | 0.106
US50 | WRh | 25_50 | 0.275 | 26 | 0.108
US50 | WRh | 25_50 | 0.275 | 27 | 0.111
US50 | WRh | 25_50 | 0.275 | 28 | 0.113
US50 | WRh | 25_50 | 0.275 | 29 | 0.115
US50 | WRh | 25_50 | 0.275 | 30 | 0.116
US50 | WRh | 25_50 | 0.275 | 31 | 0.118
US50 | WRh | 25_50 | 0.275 | 32 | 0.12
US50 | WRh | 25_50 | 0.3 | 20 | 0.095
US50 | WRh | 25_50 | 0.3 | 21 | 0.1
US50 | WRh | 25_50 | 0.3 | 22 | 0.104
US50 | WRh | 25_50 | 0.3 | 23 | 0.109
US50 | WRh | 25_50 | 0.3 | 24 | 0.113
US50 | WRh | 25_50 | 0.3 | 25 | 0.116
US50 | WRh | 25_50 | 0.3 | 26 | 0.119
US50 | WRh | 25_50 | 0.3 | 27 | 0.121
US50 | WRh | 25_50 | 0.3 | 28 | 0.123
US50 | WRh | 25_50 | 0.3 | 29 | 0.125
US50 | WRh | 25_50 | 0.3 | 30 | 0.127
US50 | WRh | 25_50 | 0.3 | 31 | 0.128
US50 | WRh | 25_50 | 0.3 | 32 | 0.13
US50 | WRh | 25_50 | 0.3 | 33 | 0.132
US50 | WRh | 25_50 | 0.3 | 34 | 0.134
US50 | WRh | 25_50 | 0.3 | 35 | 0.136
US50 | WRh | 25_50 | 0.3 | 36 | 0.137
US50 | WRh | 25_50 | 0.325 | 20 | 0.105
US50 | WRh | 25_50 | 0.325 | 21 | 0.11
US50 | WRh | 25_50 | 0.325 | 22 | 0.115
US50 | WRh | 25_50 | 0.325 | 23 | 0.119
US50 | WRh | 25_50 | 0.325 | 24 | 0.123
US50 | WRh | 25_50 | 0.325 | 25 | 0.127
US50 | WRh | 25_50 | 0.325 | 26 | 0.129
US50 | WRh | 25_50 | 0.325 | 27 | 0.131
US50 | WRh | 25_50 | 0.325 | 28 | 0.133
US50 | WRh | 25_50 | 0.325 | 29 | 0.135
US50 | WRh | 25_50 | 0.325 | 30 | 0.136
US50 | WRh | 25_50 | 0.325 | 31 | 0.138
US50 | WRh | 25_50 | 0.325 | 32 | 0.14
US50 | WRh | 25_50 | 0.325 | 33 | 0.141
US50 | WRh | 25_50 | 0.325 | 34 | 0.143
US50 | WRh | 25_50 | 0.325 | 35 | 0.145
US50 | WRh | 25_50 | 0.325 | 36 | 0.147
US50 | WRh | 25_50 | 0.325 | 37 | 0.149
US50 | WRh | 25_50 | 0.325 | 38 | 0.15
US50 | WRh | 25_50 | 0.325 | 39 | 0.152
US50 | WRh | 25_50 | 0.325 | 40 | 0.154
US50 | WRh | 25_50 | 0.325 | 41 | 0.155
US50 | WRh | 25_50 | 0.325 | 42 | 0.157
US50 | WRh | 25_50 | 0.35 | 20 | 0.116
US50 | WRh | 25_50 | 0.35 | 21 | 0.121
US50 | WRh | 25_50 | 0.35 | 22 | 0.125
US50 | WRh | 25_50 | 0.35 | 23 | 0.129
US50 | WRh | 25_50 | 0.35 | 24 | 0.134
US50 | WRh | 25_50 | 0.35 | 25 | 0.137
US50 | WRh | 25_50 | 0.35 | 26 | 0.139
US50 | WRh | 25_50 | 0.35 | 27 | 0.141
US50 | WRh | 25_50 | 0.35 | 28 | 0.143
US50 | WRh | 25_50 | 0.35 | 29 | 0.144
US50 | WRh | 25_50 | 0.35 | 30 | 0.146
US50 | WRh | 25_50 | 0.35 | 31 | 0.148
US50 | WRh | 25_50 | 0.35 | 32 | 0.149
US50 | WRh | 25_50 | 0.35 | 33 | 0.151
US50 | WRh | 25_50 | 0.35 | 34 | 0.152
US50 | WRh | 25_50 | 0.35 | 35 | 0.154
US50 | WRh | 25_50 | 0.35 | 36 | 0.156
US50 | WRh | 25_50 | 0.35 | 37 | 0.158
US50 | WRh | 25_50 | 0.35 | 38 | 0.159
US50 | WRh | 25_50 | 0.35 | 39 | 0.161
US50 | WRh | 25_50 | 0.35 | 40 | 0.163
US50 | WRh | 25_50 | 0.35 | 41 | 0.164
US50 | WRh | 25_50 | 0.35 | 42 | 0.166
US50 | WRh | 25_50 | 0.35 | 43 | 0.168
US50 | WRh | 25_50 | 0.35 | 44 | 0.169
US50 | WRh | 25_50 | 0.35 | 45 | 0.171
US50 | WRh | 25_50 | 0.35 | 46 | 0.172
US50 | WRh | 25_50 | 0.35 | 47 | 0.174
US50 | WRh | 25_50 | 0.375 | 20 | 0.126
US50 | WRh | 25_50 | 0.375 | 21 | 0.131
US50 | WRh | 25_50 | 0.375 | 22 | 0.135
US50 | WRh | 25_50 | 0.375 | 23 | 0.14
US50 | WRh | 25_50 | 0.375 | 24 | 0.144
US50 | WRh | 25_50 | 0.375 | 25 | 0.147
US50 | WRh | 25_50 | 0.375 | 26 | 0.149
US50 | WRh | 25_50 | 0.375 | 27 | 0.151
US50 | WRh | 25_50 | 0.375 | 28 | 0.153
US50 | WRh | 25_50 | 0.375 | 29 | 0.154
US50 | WRh | 25_50 | 0.375 | 30 | 0.155
US50 | WRh | 25_50 | 0.375 | 31 | 0.157
US50 | WRh | 25_50 | 0.375 | 32 | 0.158
US50 | WRh | 25_50 | 0.375 | 33 | 0.16
US50 | WRh | 25_50 | 0.375 | 34 | 0.161
US50 | WRh | 25_50 | 0.375 | 35 | 0.163
US50 | WRh | 25_50 | 0.375 | 36 | 0.165
US50 | WRh | 25_50 | 0.375 | 37 | 0.166
US50 | WRh | 25_50 | 0.375 | 38 | 0.168
US50 | WRh | 25_50 | 0.375 | 39 | 0.17
US50 | WRh | 25_50 | 0.375 | 40 | 0.171
US50 | WRh | 25_50 | 0.375 | 41 | 0.173
US50 | WRh | 25_50 | 0.375 | 42 | 0.175
US50 | WRh | 25_50 | 0.375 | 43 | 0.176
US50 | WRh | 25_50 | 0.375 | 44 | 0.178
US50 | WRh | 25_50 | 0.375 | 45 | 0.18
US50 | WRh | 25_50 | 0.375 | 46 | 0.181
US50 | WRh | 25_50 | 0.375 | 47 | 0.183
US50 | WRh | 25_50 | 0.375 | 48 | 0.184
US50 | WRh | 25_50 | 0.375 | 49 | 0.186
US50 | WRh | 25_50 | 0.4 | 20 | 0.137
US50 | WRh | 25_50 | 0.4 | 21 | 0.141
US50 | WRh | 25_50 | 0.4 | 22 | 0.146
US50 | WRh | 25_50 | 0.4 | 23 | 0.15
US50 | WRh | 25_50 | 0.4 | 24 | 0.154
US50 | WRh | 25_50 | 0.4 | 25 | 0.157
US50 | WRh | 25_50 | 0.4 | 26 | 0.159
US50 | WRh | 25_50 | 0.4 | 27 | 0.161
US50 | WRh | 25_50 | 0.4 | 28 | 0.162
US50 | WRh | 25_50 | 0.4 | 29 | 0.164
US50 | WRh | 25_50 | 0.4 | 30 | 0.165
US50 | WRh | 25_50 | 0.4 | 31 | 0.166
US50 | WRh | 25_50 | 0.4 | 32 | 0.167
US50 | WRh | 25_50 | 0.4 | 33 | 0.169
US50 | WRh | 25_50 | 0.4 | 34 | 0.17
US50 | WRh | 25_50 | 0.4 | 35 | 0.172
US50 | WRh | 25_50 | 0.4 | 36 | 0.173
US50 | WRh | 25_50 | 0.4 | 37 | 0.175
US50 | WRh | 25_50 | 0.4 | 38 | 0.177
US50 | WRh | 25_50 | 0.4 | 39 | 0.178
US50 | WRh | 25_50 | 0.4 | 40 | 0.18
US50 | WRh | 25_50 | 0.4 | 41 | 0.181
US50 | WRh | 25_50 | 0.4 | 42 | 0.183
US50 | WRh | 25_50 | 0.4 | 43 | 0.185
US50 | WRh | 25_50 | 0.4 | 44 | 0.186
US50 | WRh | 25_50 | 0.4 | 45 | 0.188
US50 | WRh | 25_50 | 0.4 | 46 | 0.189
US50 | WRh | 25_50 | 0.4 | 47 | 0.191
US50 | WRh | 25_50 | 0.4 | 48 | 0.193
US50 | WRh | 25_50 | 0.4 | 49 | 0.194
US50 | WRh | 25_50 | 0.425 | 20 | 0.147
US50 | WRh | 25_50 | 0.425 | 21 | 0.152
US50 | WRh | 25_50 | 0.425 | 22 | 0.156
US50 | WRh | 25_50 | 0.425 | 23 | 0.16
US50 | WRh | 25_50 | 0.425 | 24 | 0.164
US50 | WRh | 25_50 | 0.425 | 25 | 0.167
US50 | WRh | 25_50 | 0.425 | 26 | 0.169
US50 | WRh | 25_50 | 0.425 | 27 | 0.171
US50 | WRh | 25_50 | 0.425 | 28 | 0.172
US50 | WRh | 25_50 | 0.425 | 29 | 0.173
US50 | WRh | 25_50 | 0.425 | 30 | 0.174
US50 | WRh | 25_50 | 0.425 | 31 | 0.175
US50 | WRh | 25_50 | 0.425 | 32 | 0.177
US50 | WRh | 25_50 | 0.425 | 33 | 0.178
US50 | WRh | 25_50 | 0.425 | 34 | 0.179
US50 | WRh | 25_50 | 0.425 | 35 | 0.181
US50 | WRh | 25_50 | 0.425 | 36 | 0.182
US50 | WRh | 25_50 | 0.425 | 37 | 0.184
US50 | WRh | 25_50 | 0.425 | 38 | 0.185
US50 | WRh | 25_50 | 0.425 | 39 | 0.187
US50 | WRh | 25_50 | 0.425 | 40 | 0.188
US50 | WRh | 25_50 | 0.425 | 41 | 0.19
US50 | WRh | 25_50 | 0.425 | 42 | 0.191
US50 | WRh | 25_50 | 0.425 | 43 | 0.193
US50 | WRh | 25_50 | 0.425 | 44 | 0.195
US50 | WRh | 25_50 | 0.425 | 45 | 0.196
US50 | WRh | 25_50 | 0.425 | 46 | 0.198
US50 | WRh | 25_50 | 0.425 | 47 | 0.199
US50 | WRh | 25_50 | 0.425 | 48 | 0.201
US50 | WRh | 25_50 | 0.425 | 49 | 0.202
US50 | WRh | 25_50 | 0.45 | 20 | 0.158
US50 | WRh | 25_50 | 0.45 | 21 | 0.163
US50 | WRh | 25_50 | 0.45 | 22 | 0.167
US50 | WRh | 25_50 | 0.45 | 23 | 0.171
US50 | WRh | 25_50 | 0.45 | 24 | 0.174
US50 | WRh | 25_50 | 0.45 | 25 | 0.177
US50 | WRh | 25_50 | 0.45 | 26 | 0.179
US50 | WRh | 25_50 | 0.45 | 27 | 0.18
US50 | WRh | 25_50 | 0.45 | 28 | 0.182
US50 | WRh | 25_50 | 0.45 | 29 | 0.183
US50 | WRh | 25_50 | 0.45 | 30 | 0.183
US50 | WRh | 25_50 | 0.45 | 31 | 0.184
US50 | WRh | 25_50 | 0.45 | 32 | 0.186
US50 | WRh | 25_50 | 0.45 | 33 | 0.187
US50 | WRh | 25_50 | 0.45 | 34 | 0.188
US50 | WRh | 25_50 | 0.45 | 35 | 0.189
US50 | WRh | 25_50 | 0.45 | 36 | 0.19
US50 | WRh | 25_50 | 0.45 | 37 | 0.192
US50 | WRh | 25_50 | 0.45 | 38 | 0.193
US50 | WRh | 25_50 | 0.45 | 39 | 0.195
US50 | WRh | 25_50 | 0.45 | 40 | 0.196
US50 | WRh | 25_50 | 0.45 | 41 | 0.198
US50 | WRh | 25_50 | 0.45 | 42 | 0.199
US50 | WRh | 25_50 | 0.45 | 43 | 0.201
US50 | WRh | 25_50 | 0.45 | 44 | 0.203
US50 | WRh | 25_50 | 0.45 | 45 | 0.204
US50 | WRh | 25_50 | 0.45 | 46 | 0.206
US50 | WRh | 25_50 | 0.45 | 47 | 0.207
US50 | WRh | 25_50 | 0.45 | 48 | 0.209
US50 | WRh | 25_50 | 0.45 | 49 | 0.21
US50 | WRh | 25_50 | 0.475 | 20 | 0.169
US50 | WRh | 25_50 | 0.475 | 21 | 0.173
US50 | WRh | 25_50 | 0.475 | 22 | 0.177
US50 | WRh | 25_50 | 0.475 | 23 | 0.181
US50 | WRh | 25_50 | 0.475 | 24 | 0.184
US50 | WRh | 25_50 | 0.475 | 25 | 0.187
US50 | WRh | 25_50 | 0.475 | 26 | 0.189
US50 | WRh | 25_50 | 0.475 | 27 | 0.19
US50 | WRh | 25_50 | 0.475 | 28 | 0.191
US50 | WRh | 25_50 | 0.475 | 29 | 0.192
US50 | WRh | 25_50 | 0.475 | 30 | 0.193
US50 | WRh | 25_50 | 0.475 | 31 | 0.194
US50 | WRh | 25_50 | 0.475 | 32 | 0.194
US50 | WRh | 25_50 | 0.475 | 33 | 0.196
US50 | WRh | 25_50 | 0.475 | 34 | 0.197
US50 | WRh | 25_50 | 0.475 | 35 | 0.198
US50 | WRh | 25_50 | 0.475 | 36 | 0.199
US50 | WRh | 25_50 | 0.475 | 37 | 0.2
US50 | WRh | 25_50 | 0.475 | 38 | 0.202
US50 | WRh | 25_50 | 0.475 | 39 | 0.203
US50 | WRh | 25_50 | 0.475 | 40 | 0.205
US50 | WRh | 25_50 | 0.475 | 41 | 0.206
US50 | WRh | 25_50 | 0.475 | 42 | 0.208
US50 | WRh | 25_50 | 0.475 | 43 | 0.209
US50 | WRh | 25_50 | 0.475 | 44 | 0.211
US50 | WRh | 25_50 | 0.475 | 45 | 0.212
US50 | WRh | 25_50 | 0.475 | 46 | 0.214
US50 | WRh | 25_50 | 0.475 | 47 | 0.215
US50 | WRh | 25_50 | 0.475 | 48 | 0.217
US50 | WRh | 25_50 | 0.475 | 49 | 0.218
US50 | WRh | 25_50 | 0.5 | 20 | 0.18
US50 | WRh | 25_50 | 0.5 | 21 | 0.183
US50 | WRh | 25_50 | 0.5 | 22 | 0.187
US50 | WRh | 25_50 | 0.5 | 23 | 0.191
US50 | WRh | 25_50 | 0.5 | 24 | 0.194
US50 | WRh | 25_50 | 0.5 | 25 | 0.197
US50 | WRh | 25_50 | 0.5 | 26 | 0.199
US50 | WRh | 25_50 | 0.5 | 27 | 0.2
US50 | WRh | 25_50 | 0.5 | 28 | 0.2
US50 | WRh | 25_50 | 0.5 | 29 | 0.201
US50 | WRh | 25_50 | 0.5 | 30 | 0.202
US50 | WRh | 25_50 | 0.5 | 31 | 0.203
US50 | WRh | 25_50 | 0.5 | 32 | 0.203
US50 | WRh | 25_50 | 0.5 | 33 | 0.204
US50 | WRh | 25_50 | 0.5 | 34 | 0.205
US50 | WRh | 25_50 | 0.5 | 35 | 0.206
US50 | WRh | 25_50 | 0.5 | 36 | 0.207
US50 | WRh | 25_50 | 0.5 | 37 | 0.209
US50 | WRh | 25_50 | 0.5 | 38 | 0.21
US50 | WRh | 25_50 | 0.5 | 39 | 0.211
US50 | WRh | 25_50 | 0.5 | 40 | 0.213
US50 | WRh | 25_50 | 0.5 | 41 | 0.214
US50 | WRh | 25_50 | 0.5 | 42 | 0.216
US50 | WRh | 25_50 | 0.5 | 43 | 0.217
US50 | WRh | 25_50 | 0.5 | 44 | 0.218
US50 | WRh | 25_50 | 0.5 | 45 | 0.22
US50 | WRh | 25_50 | 0.5 | 46 | 0.221
US50 | WRh | 25_50 | 0.5 | 47 | 0.223
US50 | WRh | 25_50 | 0.5 | 48 | 0.224
US50 | WRh | 25_50 | 0.5 | 49 | 0.226
US50 | WRh | 25_50 | 0.525 | 21 | 0.194
US50 | WRh | 25_50 | 0.525 | 22 | 0.197
US50 | WRh | 25_50 | 0.525 | 23 | 0.201
US50 | WRh | 25_50 | 0.525 | 24 | 0.204
US50 | WRh | 25_50 | 0.525 | 25 | 0.207
US50 | WRh | 25_50 | 0.525 | 26 | 0.208
US50 | WRh | 25_50 | 0.525 | 27 | 0.209
US50 | WRh | 25_50 | 0.525 | 28 | 0.21
US50 | WRh | 25_50 | 0.525 | 29 | 0.211
US50 | WRh | 25_50 | 0.525 | 30 | 0.211
US50 | WRh | 25_50 | 0.525 | 31 | 0.212
US50 | WRh | 25_50 | 0.525 | 32 | 0.212
US50 | WRh | 25_50 | 0.525 | 33 | 0.213
US50 | WRh | 25_50 | 0.525 | 34 | 0.214
US50 | WRh | 25_50 | 0.525 | 35 | 0.215
US50 | WRh | 25_50 | 0.525 | 36 | 0.216
US50 | WRh | 25_50 | 0.525 | 37 | 0.217
US50 | WRh | 25_50 | 0.525 | 38 | 0.218
US50 | WRh | 25_50 | 0.525 | 39 | 0.22
US50 | WRh | 25_50 | 0.525 | 40 | 0.221
US50 | WRh | 25_50 | 0.525 | 41 | 0.222
US50 | WRh | 25_50 | 0.525 | 42 | 0.224
US50 | WRh | 25_50 | 0.525 | 43 | 0.225
US50 | WRh | 25_50 | 0.525 | 44 | 0.226
US50 | WRh | 25_50 | 0.525 | 45 | 0.228
US50 | WRh | 25_50 | 0.525 | 46 | 0.229
US50 | WRh | 25_50 | 0.525 | 47 | 0.231
US50 | WRh | 25_50 | 0.525 | 48 | 0.232
US50 | WRh | 25_50 | 0.525 | 49 | 0.234
US50 | WRh | 25_50 | 0.55 | 21 | 0.204
US50 | WRh | 25_50 | 0.55 | 22 | 0.207
US50 | WRh | 25_50 | 0.55 | 23 | 0.211
US50 | WRh | 25_50 | 0.55 | 24 | 0.214
US50 | WRh | 25_50 | 0.55 | 25 | 0.216
US50 | WRh | 25_50 | 0.55 | 26 | 0.218
US50 | WRh | 25_50 | 0.55 | 27 | 0.218
US50 | WRh | 25_50 | 0.55 | 28 | 0.219
US50 | WRh | 25_50 | 0.55 | 29 | 0.22
US50 | WRh | 25_50 | 0.55 | 30 | 0.22
US50 | WRh | 25_50 | 0.55 | 31 | 0.221
US50 | WRh | 25_50 | 0.55 | 32 | 0.221
US50 | WRh | 25_50 | 0.55 | 33 | 0.222
US50 | WRh | 25_50 | 0.55 | 34 | 0.223
US50 | WRh | 25_50 | 0.55 | 35 | 0.223
US50 | WRh | 25_50 | 0.55 | 36 | 0.224
US50 | WRh | 25_50 | 0.55 | 37 | 0.225
US50 | WRh | 25_50 | 0.55 | 38 | 0.226
US50 | WRh | 25_50 | 0.55 | 39 | 0.228
US50 | WRh | 25_50 | 0.55 | 40 | 0.229
US50 | WRh | 25_50 | 0.55 | 41 | 0.23
US50 | WRh | 25_50 | 0.55 | 42 | 0.232
US50 | WRh | 25_50 | 0.55 | 43 | 0.233
US50 | WRh | 25_50 | 0.55 | 44 | 0.234
US50 | WRh | 25_50 | 0.55 | 45 | 0.236
US50 | WRh | 25_50 | 0.55 | 46 | 0.237
US50 | WRh | 25_50 | 0.55 | 47 | 0.238
US50 | WRh | 25_50 | 0.55 | 48 | 0.24
US50 | WRh | 25_50 | 0.55 | 49 | 0.241
US50 | WRh | 25_50 | 0.575 | 22 | 0.217
US50 | WRh | 25_50 | 0.575 | 23 | 0.22
US50 | WRh | 25_50 | 0.575 | 24 | 0.223
US50 | WRh | 25_50 | 0.575 | 25 | 0.226
US50 | WRh | 25_50 | 0.575 | 26 | 0.227
US50 | WRh | 25_50 | 0.575 | 27 | 0.228
US50 | WRh | 25_50 | 0.575 | 28 | 0.228
US50 | WRh | 25_50 | 0.575 | 29 | 0.229
US50 | WRh | 25_50 | 0.575 | 30 | 0.229
US50 | WRh | 25_50 | 0.575 | 31 | 0.23
US50 | WRh | 25_50 | 0.575 | 32 | 0.23
US50 | WRh | 25_50 | 0.575 | 33 | 0.231
US50 | WRh | 25_50 | 0.575 | 34 | 0.231
US50 | WRh | 25_50 | 0.575 | 35 | 0.232
US50 | WRh | 25_50 | 0.575 | 36 | 0.233
US50 | WRh | 25_50 | 0.575 | 37 | 0.234
US50 | WRh | 25_50 | 0.575 | 38 | 0.235
US50 | WRh | 25_50 | 0.575 | 39 | 0.236
US50 | WRh | 25_50 | 0.575 | 40 | 0.237
US50 | WRh | 25_50 | 0.575 | 41 | 0.238
US50 | WRh | 25_50 | 0.575 | 42 | 0.239
US50 | WRh | 25_50 | 0.575 | 43 | 0.241
US50 | WRh | 25_50 | 0.575 | 44 | 0.242
US50 | WRh | 25_50 | 0.575 | 45 | 0.243
US50 | WRh | 25_50 | 0.575 | 46 | 0.245
US50 | WRh | 25_50 | 0.575 | 47 | 0.246
US50 | WRh | 25_50 | 0.575 | 48 | 0.248
US50 | WRh | 25_50 | 0.575 | 49 | 0.249
US50 | WRh | 25_50 | 0.6 | 22 | 0.227
US50 | WRh | 25_50 | 0.6 | 23 | 0.23
US50 | WRh | 25_50 | 0.6 | 24 | 0.233
US50 | WRh | 25_50 | 0.6 | 25 | 0.235
US50 | WRh | 25_50 | 0.6 | 26 | 0.236
US50 | WRh | 25_50 | 0.6 | 27 | 0.237
US50 | WRh | 25_50 | 0.6 | 28 | 0.237
US50 | WRh | 25_50 | 0.6 | 29 | 0.238
US50 | WRh | 25_50 | 0.6 | 30 | 0.238
US50 | WRh | 25_50 | 0.6 | 31 | 0.239
US50 | WRh | 25_50 | 0.6 | 32 | 0.239
US50 | WRh | 25_50 | 0.6 | 33 | 0.239
US50 | WRh | 25_50 | 0.6 | 34 | 0.24
US50 | WRh | 25_50 | 0.6 | 35 | 0.24
US50 | WRh | 25_50 | 0.6 | 36 | 0.241
US50 | WRh | 25_50 | 0.6 | 37 | 0.242
US50 | WRh | 25_50 | 0.6 | 38 | 0.243
US50 | WRh | 25_50 | 0.6 | 39 | 0.244
US50 | WRh | 25_50 | 0.6 | 40 | 0.245
US50 | WRh | 25_50 | 0.6 | 41 | 0.246
US50 | WRh | 25_50 | 0.6 | 42 | 0.247
US50 | WRh | 25_50 | 0.6 | 43 | 0.248
US50 | WRh | 25_50 | 0.6 | 44 | 0.25
US50 | WRh | 25_50 | 0.6 | 45 | 0.251
US50 | WRh | 25_50 | 0.6 | 46 | 0.252
US50 | WRh | 25_50 | 0.6 | 47 | 0.254
US50 | WRh | 25_50 | 0.6 | 48 | 0.255
US50 | WRh | 25_50 | 0.6 | 49 | 0.256
US50 | WRh | 25_50 | 0.625 | 22 | 0.237
US50 | WRh | 25_50 | 0.625 | 23 | 0.24
US50 | WRh | 25_50 | 0.625 | 24 | 0.242
US50 | WRh | 25_50 | 0.625 | 25 | 0.244
US50 | WRh | 25_50 | 0.625 | 26 | 0.245
US50 | WRh | 25_50 | 0.625 | 27 | 0.246
US50 | WRh | 25_50 | 0.625 | 28 | 0.246
US50 | WRh | 25_50 | 0.625 | 29 | 0.247
US50 | WRh | 25_50 | 0.625 | 30 | 0.247
US50 | WRh | 25_50 | 0.625 | 31 | 0.247
US50 | WRh | 25_50 | 0.625 | 32 | 0.248
US50 | WRh | 25_50 | 0.625 | 33 | 0.248
US50 | WRh | 25_50 | 0.625 | 34 | 0.249
US50 | WRh | 25_50 | 0.625 | 35 | 0.249
US50 | WRh | 25_50 | 0.625 | 36 | 0.25
US50 | WRh | 25_50 | 0.625 | 37 | 0.25
US50 | WRh | 25_50 | 0.625 | 38 | 0.251
US50 | WRh | 25_50 | 0.625 | 39 | 0.252
US50 | WRh | 25_50 | 0.625 | 40 | 0.253
US50 | WRh | 25_50 | 0.625 | 41 | 0.254
US50 | WRh | 25_50 | 0.625 | 42 | 0.255
US50 | WRh | 25_50 | 0.625 | 43 | 0.256
US50 | WRh | 25_50 | 0.625 | 44 | 0.257
US50 | WRh | 25_50 | 0.625 | 45 | 0.259
US50 | WRh | 25_50 | 0.625 | 46 | 0.26
US50 | WRh | 25_50 | 0.625 | 47 | 0.261
US50 | WRh | 25_50 | 0.625 | 48 | 0.263
US50 | WRh | 25_50 | 0.625 | 49 | 0.264
US50 | WRh | 25_50 | 0.65 | 23 | 0.249
US50 | WRh | 25_50 | 0.65 | 24 | 0.251
US50 | WRh | 25_50 | 0.65 | 25 | 0.253
US50 | WRh | 25_50 | 0.65 | 26 | 0.254
US50 | WRh | 25_50 | 0.65 | 27 | 0.255
US50 | WRh | 25_50 | 0.65 | 28 | 0.255
US50 | WRh | 25_50 | 0.65 | 29 | 0.255
US50 | WRh | 25_50 | 0.65 | 30 | 0.256
US50 | WRh | 25_50 | 0.65 | 31 | 0.256
US50 | WRh | 25_50 | 0.65 | 32 | 0.256
US50 | WRh | 25_50 | 0.65 | 33 | 0.257
US50 | WRh | 25_50 | 0.65 | 34 | 0.257
US50 | WRh | 25_50 | 0.65 | 35 | 0.257
US50 | WRh | 25_50 | 0.65 | 36 | 0.258
US50 | WRh | 25_50 | 0.65 | 37 | 0.258
US50 | WRh | 25_50 | 0.65 | 38 | 0.259
US50 | WRh | 25_50 | 0.65 | 39 | 0.26
US50 | WRh | 25_50 | 0.65 | 40 | 0.261
US50 | WRh | 25_50 | 0.65 | 41 | 0.262
US50 | WRh | 25_50 | 0.65 | 42 | 0.263
US50 | WRh | 25_50 | 0.65 | 43 | 0.264
US50 | WRh | 25_50 | 0.65 | 44 | 0.265
US50 | WRh | 25_50 | 0.65 | 45 | 0.266
US50 | WRh | 25_50 | 0.65 | 46 | 0.268
US50 | WRh | 25_50 | 0.65 | 47 | 0.269
US50 | WRh | 25_50 | 0.65 | 48 | 0.27
US50 | WRh | 25_50 | 0.65 | 49 | 0.272
US50 | WRh | 25_50 | 0.675 | 25 | 0.262
US50 | WRh | 25_50 | 0.675 | 26 | 0.263
US50 | WRh | 25_50 | 0.675 | 27 | 0.264
US50 | WRh | 25_50 | 0.675 | 28 | 0.264
US50 | WRh | 25_50 | 0.675 | 29 | 0.264
US50 | WRh | 25_50 | 0.675 | 30 | 0.264
US50 | WRh | 25_50 | 0.675 | 31 | 0.265
US50 | WRh | 25_50 | 0.675 | 32 | 0.265
US50 | WRh | 25_50 | 0.675 | 33 | 0.265
US50 | WRh | 25_50 | 0.675 | 34 | 0.265
US50 | WRh | 25_50 | 0.675 | 35 | 0.266
US50 | WRh | 25_50 | 0.675 | 36 | 0.266
US50 | WRh | 25_50 | 0.675 | 37 | 0.267
US50 | WRh | 25_50 | 0.675 | 38 | 0.267
US50 | WRh | 25_50 | 0.675 | 39 | 0.268
US50 | WRh | 25_50 | 0.675 | 40 | 0.269
US50 | WRh | 25_50 | 0.675 | 41 | 0.27
US50 | WRh | 25_50 | 0.675 | 42 | 0.271
US50 | WRh | 25_50 | 0.675 | 43 | 0.272
US50 | WRh | 25_50 | 0.675 | 44 | 0.273
US50 | WRh | 25_50 | 0.675 | 45 | 0.274
US50 | WRh | 25_50 | 0.675 | 46 | 0.275
US50 | WRh | 25_50 | 0.675 | 47 | 0.276
US50 | WRh | 25_50 | 0.675 | 48 | 0.278
US50 | WRh | 25_50 | 0.675 | 49 | 0.279
US50 | WRh | 25_50 | 0.7 | 27 | 0.272
US50 | WRh | 25_50 | 0.7 | 28 | 0.272
US50 | WRh | 25_50 | 0.7 | 29 | 0.273
US50 | WRh | 25_50 | 0.7 | 30 | 0.273
US50 | WRh | 25_50 | 0.7 | 31 | 0.273
US50 | WRh | 25_50 | 0.7 | 32 | 0.273
US50 | WRh | 25_50 | 0.7 | 33 | 0.273
US50 | WRh | 25_50 | 0.7 | 34 | 0.274
US50 | WRh | 25_50 | 0.7 | 35 | 0.274
US50 | WRh | 25_50 | 0.7 | 36 | 0.274
US50 | WRh | 25_50 | 0.7 | 37 | 0.275
US50 | WRh | 25_50 | 0.7 | 38 | 0.275
US50 | WRh | 25_50 | 0.7 | 39 | 0.276
US50 | WRh | 25_50 | 0.7 | 40 | 0.276
US50 | WRh | 25_50 | 0.7 | 41 | 0.277
US50 | WRh | 25_50 | 0.7 | 42 | 0.278
US50 | WRh | 25_50 | 0.7 | 43 | 0.279
US50 | WRh | 25_50 | 0.7 | 44 | 0.28
US50 | WRh | 25_50 | 0.7 | 45 | 0.281
US50 | WRh | 25_50 | 0.7 | 46 | 0.283
US50 | WRh | 25_50 | 0.7 | 47 | 0.284
US50 | WRh | 25_50 | 0.7 | 48 | 0.285
US50 | WRh | 25_50 | 0.7 | 49 | 0.286
US50 | WRh | 25_50 | 0.725 | 33 | 0.282
US50 | WRh | 25_50 | 0.725 | 34 | 0.282
US50 | WRh | 25_50 | 0.725 | 35 | 0.282
US50 | WRh | 25_50 | 0.725 | 36 | 0.282
US50 | WRh | 25_50 | 0.725 | 37 | 0.283
US50 | WRh | 25_50 | 0.725 | 38 | 0.283
US50 | WRh | 25_50 | 0.725 | 39 | 0.284
US50 | WRh | 25_50 | 0.725 | 40 | 0.284
US50 | WRh | 25_50 | 0.725 | 41 | 0.285
US50 | WRh | 25_50 | 0.725 | 42 | 0.286
US50 | WRh | 25_50 | 0.725 | 43 | 0.287
US50 | WRh | 25_50 | 0.725 | 44 | 0.288
US50 | WRh | 25_50 | 0.725 | 45 | 0.289
US50 | WRh | 25_50 | 0.725 | 46 | 0.29
US50 | WRh | 25_50 | 0.725 | 47 | 0.291
US50 | WRh | 25_50 | 0.725 | 48 | 0.292
US50 | WRh | 25_50 | 0.725 | 49 | 0.294
US50 | WRh | 25_50 | 0.75 | 37 | 0.291
US50 | WRh | 25_50 | 0.75 | 38 | 0.291
US50 | WRh | 25_50 | 0.75 | 39 | 0.292
US50 | WRh | 25_50 | 0.75 | 40 | 0.292
US50 | WRh | 25_50 | 0.75 | 41 | 0.293
US50 | WRh | 25_50 | 0.75 | 42 | 0.293
US50 | WRh | 25_50 | 0.75 | 43 | 0.294
US50 | WRh | 25_50 | 0.75 | 44 | 0.295
US50 | WRh | 25_50 | 0.75 | 45 | 0.296
US50 | WRh | 25_50 | 0.75 | 46 | 0.297
US50 | WRh | 25_50 | 0.75 | 47 | 0.299
US50 | WRh | 25_50 | 0.75 | 48 | 0.3
US50 | WRh | 25_50 | 0.75 | 49 | 0.301
US50 | WRh | 25_50 | 0.775 | 43 | 0.302
US50 | WRh | 25_50 | 0.775 | 44 | 0.303
US50 | WRh | 25_50 | 0.775 | 45 | 0.304
US50 | WRh | 25_50 | 0.775 | 46 | 0.305
US50 | WRh | 25_50 | 0.775 | 47 | 0.306
US50 | WRh | 25_50 | 0.775 | 48 | 0.307
US50 | WRh | 25_50 | 0.775 | 49 | 0.308
US50 | WRh | 25_50 | 0.8 | 48 | 0.314
US50 | WRh | 25_50 | 0.8 | 49 | 0.315

## DBT Z Res
3. DBT Z Resolution | [16 empty] | Removed from QC
|  | Facility Name | 0 | [6 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
|  | Mfr & Model | [7 empty] | Room ID | 0
Procedure |  |  | Equipment: | DBT image from Phantom Image Quality test
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
| Slice # |  | Relative Slice Location | Maximum Speck Signal | [10 empty] | Mean Background Signal
|  |  |  | Center Speck | 12:00 Speck | 2:00 Speck | 5:00 Speck | 7:00 Speck |  | 10:00 Speck |  | Ave Max Signal |  |  |  |  | Half Max Location
|  |  | -3 | [8 empty] | #DIV/0! |  |  |  |  | 0 | 0.5 |  |  | 0
|  |  | -2 | [8 empty] | #DIV/0! |  |  |  |  | 0 | Rising edge |  |  | 0
|  |  | -1 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 0 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 1 | [8 empty] | #DIV/0! |  |  |  |  | 0 | Falling edge
|  |  | 2 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 3 | [8 empty] | #DIV/0!
[9 empty] | Z-Res Diff (Ave Max - Background Mean) |  |  |  | △ Z-Res Diff relative to DBT Slice 0 |  |  |  | For Chart
[17 empty] | Loc | Val | FWHM
[9 empty] | #DIV/0! | [7 empty] | -3 |  | 0 | 0.5
[9 empty] | #DIV/0! | [7 empty] | -2 |  | 0 | 0.5
[9 empty] | #DIV/0! | [7 empty] | -1
[9 empty] | #DIV/0! | [7 empty] | 0
[9 empty] | #DIV/0! | [7 empty] | 1
[9 empty] | #DIV/0! | [7 empty] | 2
[9 empty] | #DIV/0! | [7 empty] | 3
[9 empty] | Current FWHM |  |  |  |  | 0.00 |  | mm
[9 empty] | Baseline FWHM | [6 empty] | mm
[9 empty] | % Change (Current vs Initial) |  |  |  |  | #DIV/0!
Overall Pass/Fail
Action Limits |  |  | Required: | Annual Survey: FWHM must be within ±30% of the baseline value. Initial: No action limit.
|  |  | Timeframe: | Failures must be corrected within 30 days.
Slice -2
Slice -1
Slice 0
Slice 1
Slice 2
ROI Placement:
Mean Background Signal - Next to Center Speck
ROI Placement:
Maximum Signal - Over
Each Speck

## 13. Compression Thickness
13. Compression Thickness Indicator Accuracy | [14 empty] | RW Notes: Worksheet no longer needed.
|  | Facility Name | 0 | [5 empty] | SBBAP ID-Unit# | 0 |  |  | - | 0
|  | Mfr & Model | [6 empty] | Room ID | 0
Procedure
Test results from the survey of the diagnostic unit may be entered in this form, without repeating the test.
| Nominal Thickness (cm) |  |  |  | Mesaured Thickness (cm) | Average | Error | Acceptable Range? (mean within +/- 0.5 cm) |  |  |  | Acceptable repeatibility? (within 0.2 cm)
2
4
6
8
Overall Pass/Fail
Action Limits |  |  | Required:
|  |  | Timeframe: |  | Failures must be corrected within 30 days; for Initial, before clinical use.
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This test is not in the main part of the MP section. It is still in the table of MP tests.