## Instructions
Radiologic Technologist's QC Test Form Workbook
Instructions
• | The forms may be completed on paper or on computer by using the spreadsheets.
• | If completing via computer, be sure to verify all entries so they are correct.
• | Before completing the forms, either on paper or by computer, enter facility information in Facility Info tab first; information auto-populates in other pages.
• | If completing the forms on computer, drop down lists are available within some cells to complete the required information; if the list does not contain the necessary information; complete the cell yourself.
• | Your system may have multiple, clinically-used Image Modes that must be tested (e.g., Stereotactic, Tomosynthesis). See next tab for the required tests.
If completing via paper, write in the Image Mode that is being tested on the top of all applicable forms. Make additional copies of the forms for each image mode tested.
If completing via computer, make a copy of the tab with each form for all applicable image modes tested. Select the Image Mode tested on each form using the drop-down list of Image Modes.
• | Enter data into white cells.
• | Light yellow cells are for pass/fail results of subtests within the main test; depending on the page, you may enter the pass/fail results using the dropdown list.
• | Bright yellow cells are for pass/fail results of the overall test; depending on the page, you may enter the pass/fail results using the dropdown list.
• | Do not enter data into gray cells. Some gray cells might contain formulae that will be lost.
Revisions

## TOC
Quality Control Tests
Technologist's Tests
Note: Complete Facility tab first to populate facility information into the forms
| Test* | Minimum Frequency** | Corrective Action Timeframe
1. | Localization Accuracy | Daily | Before clinical use
2. | Phantom Image Quality and Artifacts | Weekly | Before clinical use
3. | Visual Checklist | Monthly | Critical: Before clinical use; Less critical: Within 30 days
4. | Compression Thickness Indicator | Monthly | Before clinical use
5. | Acquisition Workstation (AW) Monitor QC | Monthly | Within 30 days
6. | Compression Force | Semiannually | Within 30 days
7. | Facility QC Review | Annually | Per relevant QC test(s)
8. | Manufacturer Calibrations (if applicable) | Per Manufacturer | Per Manufacturer
9. | Optional: Physician Feedback | Annually | Facility's choice
10 | Optional: Repeat Analysis | Semiannually | Facility's choice
* All required tests (except Facility QC Review) must be performed after installation before first clinical use.
** This is a minimum frequency; tests may be performed more often if problems are noted. Also, weekly tests do not need to be performed if biopsy procedures are not performed during that week. However, the test must be performed prior to patient biopsies, once biopsy procedures resume. In these cases, be sure to note in the QC charts that biopsy procedures were not performed during this time. Mobile units have additional Minimum Frequencies (see below).
Management Forms
Corrective Action Log
Digital Mammography Unit QC Summary Checklist
Mobile Systems
In addition to meeting the minimum frequencies outlined in the table above, the following tests must be performed according to their procedure after each move of the mobile system to a new location:
• Phantom Image Quality - after each relocation and prior to biopsy procedures.
• Compression Thickness Indicator - after each relocation and prior to biopsy procedures.
• Localization Accuracy (Daily procedure).
• Any QC test(s) that are required for a mobile mammography unit according to its relevant QC manual or manufacturer requirements.

## Facility Info
Facility, Unit and Test Equipment Data
Enter your facility's information in the appropriate cells.
Facility Information |  | Your facility's information will populate the appropriate cells in the workbook's worksheets.
Facility Name: | SurgiCenter |  | Entry is passed to many tabs.
Address: | 706 Weston Way |  | Entry is passed to many tabs.
Cell is not referenced in another worksheet
City, State, Zip: | Arlington, VA
MGBBAP ID# (00000): | 12001 |  | Entry is passed to many tabs.
MGBBAP ID Unit# (00): | 1 |  | Entry is passed to many tabs.
Supervising Physician: | P. Piatt |  | Cell is not referenced in another worksheet
MGBB Quality Control Technologist: | T. Bran |  | Cell is not referenced in another worksheet
Unit Information
Room ID: | Biopsy |  | Entry is passed to many tabs.
X-Ray Unit Manufacturer: | Hologic |  | Entry is passed to many tabs.
X-Ray Unit Model: | Affirm |  | Entry is passed to many tabs.
X-Ray Unit Gantry Serial #: | 10-1205 |  | Cell is not referenced in another worksheet
X-Ray Unit Date of Manufacture: | 1/26/2023 |  | Cell is not referenced in another worksheet
X-Ray Unit Date of Installation: | 2/13/2024 |  | Cell is not referenced in another worksheet
Imaging Mode A: | Stereotactic |  | Entry is passed to Localization Verification tab.
Imaging Mode B: | Tomosynthesis |  | Entry is passed to Localization Verification tab.
Test Equipment Info
Image Quality Phantom Type**: | ACR Original |  | Entry is passed to Image Quality tab.
Phantom Manufacturer: | Gammex |  | Entry is passed to Image Quality tab.
Phantom SN: | 2590-03 |  | Entry is passed to Image Quality tab.
Medical Physicist Info
Medical Physicist Name: | R. Wong |  | Cell is not referenced in another worksheet
Telephone Number: | 555-551-1212 |  | Cell is not referenced in another worksheet
email: | GreatestMP@gmail.com |  | Cell is not referenced in another worksheet
**Image Quality Phantom Type: The acceptable phantoms are 1) the original ACR phantom (4" square), 2) the "Mini Phantom", and 3) the ACR Digital Mammography Phantom (12" x 10").

## 1. (hidden)
1. Localization accuracy | [9 empty] | Daily | Done, Draft 1
11/19/2024: RW changed all SBB to MGBBAP
Facility | SurgiCenter | [5 empty] | Room ID |  | Biopsy
MGBBAP ID-Unit# |  | 12001 |  |  | - | 1
Unit Mfr & Model |  | Hologic Affirm |  |  |  |  | Imaging Mode | Stereotactic
Image receptor size
AEC mode
Target/filter
kVp
|  |  | Coordinate system* | #REF!

## 1. Localization Accuracy (Daily
1. Localization Accuracy | [13 empty] | Daily | Done, Draft 1 | 1/9/2026 RW: Note to self: Finished reformatting all calculated cells to gray background.
[15 empty] | Polar coordinates | Also, changed calculations so difference is now (Calculated Location- Preset).
Facility |  |  | SurgiCenter | [5 empty] | Room ID |  |  | Biopsy
SBBAP ID-Unit# |  |  | 12001 |  | - | 1 |  |  | Unit Mfr & Model |  |  | Hologic Affirm |  |  |  | 4/1 Pass/fail calculations not working
Procedure: Use either the manufacturer's methodolgy or ACR methology to target a preset location.
CC Approach
[5 empty] | Stereo Mode |  |  |  | DBT Mode |  |  |  | Pass/Fail
[5 empty] | x |  | y | z | x | y | z
Date | Preset Location
Calculated Location
| Difference | [13 empty] | =IF(OR(COUNTA(J13:U13)=0,A13=""),"N/A",
| Calculated Location | [13 empty] | IF(AND(
| Difference | [13 empty] | OR(IFERROR(ABS(J14)<=1,FALSE),J14=""),
| Calculated Location | [13 empty] | OR(IFERROR(ABS(L14)<=1,FALSE),L14=""),
| Difference | [13 empty] | OR(IFERROR(ABS(N14)<=1,FALSE),N14=""),
| Calculated Location | [13 empty] | OR(IFERROR(ABS(P14)<=1,FALSE),P14=""),
| Difference | [13 empty] | OR(IFERROR(ABS(R14)<=1,FALSE),R14=""),
| Calculated Location | [13 empty] | OR(IFERROR(ABS(T14)<=1,FALSE),T14="")),
| Difference | [13 empty] | "PASS","FAIL"))
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Analyses: | Measure or calculate the difference between the Calculated Location (of the needle tip/phantom, calculated by the targeting computer), and the Preset Location, for each coordinate.
Action Limits |  | Required: |  | The needle tip should be within 1 mm in each of the cardinal directions.
Action Limits | Criteria: |  | Each calculated coordinate (X, Y, and Z) must be within 1 mm of the coordinates of the preset location, regardless of which procedure was used.
| Timeframe: |  | The cause of the test failure must be corrected before clinical use.

## 1. Loc. Accuracy (Monthly)
1. Localization Accuracy | [21 empty] | Monthly | Done, Draft 1 | 1/9/2026 RW: Note to self: Finished reformatting all calculated cells to gray background.
[23 empty] | Polar coordinates | Also, changed calculations so difference is now (Calculated Location- Preset).
Facility |  |  |  | SurgiCenter | [6 empty] | Room ID |  |  | Biopsy
SBBAP ID-Unit# |  |  |  | 12001 |  | - | 1 |  |  |  | Unit Mfr & Model |  |  | Hologic Affirm | [8 empty] | 4/1 Pass/fail calculations not working
Procedure: Use either the manufacturer's methodolgy or ACR methology to target a preset location.
[6 empty] | Left Lateral Approach | [8 empty] | Right Lateral Approach
[6 empty] | Stereo Mode |  |  |  | DBT Mode |  |  | PASS/FAIL |  | Stereo Mode |  |  | DBT Mode |  |  | PASS/FAIL
[6 empty] | x |  | y | z | x | y | z |  |  | x | y | z | x | y | z
Date | Preset Location |  |  |  |  | 10 |  | 20 | 30 | 10 | 20 | 30
| Calculated Location |  |  |  |  | 11 |  | 23 | 29
| Difference |  |  |  |  | 1 |  | 3 | -1 | [13 empty] | =IF(OR(COUNTA(J13:U13)=0,A13=""),"N/A",
| Calculated Location | [21 empty] | IF(AND(
| Difference | [21 empty] | OR(IFERROR(ABS(J14)<=1,FALSE),J14=""),
| Calculated Location | [21 empty] | OR(IFERROR(ABS(L14)<=1,FALSE),L14=""),
| Difference | [21 empty] | OR(IFERROR(ABS(N14)<=1,FALSE),N14=""),
| Calculated Location | [21 empty] | OR(IFERROR(ABS(P14)<=1,FALSE),P14=""),
| Difference | [21 empty] | OR(IFERROR(ABS(R14)<=1,FALSE),R14=""),
| Calculated Location | [21 empty] | OR(IFERROR(ABS(T14)<=1,FALSE),T14="")),
| Difference | [21 empty] | "PASS","FAIL"))
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Calculated Location
Difference
Analyses:
Measure or calculate the difference between the Calculated Location (of the needle tip/phantom, calculated by the targeting computer), and the Preset Location, for each coordinate.
Action Limits |  | Criteria: |  |  | Each calculated coordinate (X, Y, and Z) must be within 1 mm of the coordinates of the preset location, regardless of which procedure was used.
|  |  | Timeframe: |  | The cause of the test failure must be corrected before clinical use.

## 1. A (hidden)
1. Localization accuracy | [10 empty] | Daily | Done, Draft 1
Cartesian Coordinates
Facility |  | SurgiCenter | [5 empty] | Room ID |  | Biopsy
MGBBAP ID-Unit# |  |  | 12001 |  | - | 1
Unit Mfr & Model |  |  | Hologic Affirm |  |  |  |  | Imaging Mode | 2D w/Add-on DBT
Year
Date
Tech
Preset Location | [6 empty] | #REF!
#REF!
Z
Calculated Location | [6 empty] | #REF!
#REF!
Z
Difference | [6 empty] | #REF!
#REF!
Z
Overall Pass/Fail
Analyses: | Measure or calculate the difference between the Calculated Location (of the needle tip, calcuated by the targetting computer), and the Preset Location, for each coordinate.
Action Limits |  | Required: |  | The needle tip should be within 1 mm in each of the cardinal directions.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 1. A (2) (hidden)
1. Localization accuracy | [10 empty] | Daily | Done, Draft 1
Cartesian Coordinates
Facility |  | SurgiCenter | [5 empty] | Room ID |  | Biopsy
MGBBAP ID-Unit# |  |  | 12001 |  | - | 1
Unit Mfr & Model |  |  | Hologic Affirm |  |  |  |  | Imaging Mode | Tomosynthesis
Year
Date
Tech
Preset Location | [6 empty] | #REF!
#REF!
Z
Calculated Location | [6 empty] | #REF!
#REF!
Z
Difference | [6 empty] | #REF!
#REF!
Z
Overall Pass/Fail
Analyses: | Measure or calculate the difference between the Calculated Location (of the needle tip, calcuated by the targetting computer), and the Preset Location, for each coordinate.
Action Limits |  | Required: |  | The needle tip should be within 1 mm in each of the cardinal directions.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 1. B (hidden)
1. Localization accuracy | [10 empty] | Daily | Done, Draft 1
Polar coordinates
Facility |  | SurgiCenter | [5 empty] | Room ID |  | Biopsy
MGBBAP ID-Unit# |  |  | 12001 |  | - | 1 | [5 empty] | Need WG consensus: Delete this worksheet? Any systems still using polar coordinates? RW
Unit Mfr & Model |  |  | Hologic Affirm |  |  |  |  | Imaging Mode | Stereotactic
Year
Date
Tech
Preset Location
Calculated Location
Difference
Overall Pass/Fail
Analyses: | Measure or calculate the difference between the Calculated Location (of the needle tip, calcuated by the targetting computer), and the Preset Location, for each coordinate.
Action Limits |  | Required: |  | The needle tip should be within 1 mm in each of the cardinal directions.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 1. B (2) (hidden)
1. Localization accuracy | [10 empty] | Daily | Done, Draft 1
Polar coordinates
Facility |  | SurgiCenter | [5 empty] | Room ID |  | Biopsy
MGBBAP ID-Unit# |  |  | 12001 |  | - | 1 | [5 empty] | Need WG consensus: Delete this worksheet? Any systems still using polar coordinates? RW
Unit Mfr & Model |  |  | Hologic Affirm |  |  |  |  | Imaging Mode | Tomosynthesis
Year
Date
Tech
Preset Location
Calculated Location
Difference
Overall Pass/Fail
Analyses: | Measure or calculate the difference between the Calculated Location (of the needle tip, calcuated by the targetting computer), and the Preset Location, for each coordinate.
Action Limits |  | Required: |  | The needle tip should be within 1 mm in each of the cardinal directions.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 2. Phantom IQ - Artifacts (SBB)
2. Phantom Image Quality and Artifacts (SBB) | [15 empty] | Weekly
| Facility |  | SurgiCenter | [6 empty] | Room ID |  |  | Biopsy
| SBBAP ID-Unit # |  | 12001 |  | - | 1 |  |  |  | Phantom |  |  | Gammex 2590-03
| Unit Mfr & Model |  | Hologic Affirm | [7 empty] | Phantom Type: |  | ACR Original
[7 empty] | Imaging Mode | Stereotactic
Year
Date (month & day) | [16 empty] | Stereotactic | DBT
Tech Initials | [16 empty] | Mo/Mo | Mo/Mo
Resulting Techniques |  | View | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Mo/Rh | Mo/Rh
|  | Paddle | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Rh/Ag | Rh/Ag
|  | AEC mode | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Rh/Rh | Rh/Rh
|  | AEC Position | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | W/Ag | W/Al
|  | Exposure Comensation Step | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | W/Al | W/Rh
|  | Slice or slab # (DBT only) | [5 empty] | -- | -- |  |  | -- |  | -- | -- |  | W/Rh
Target/filter
kVp
|  | mAs | [16 empty] | Object
Phantom Score |  | Artifacts P/F | [5 empty] | P | [8 empty] | Min Score | Type | Fibers | Specks | Masses
|  | Fiber score | [5 empty] | 4 | [8 empty] | 4 | ACR DM | 2 | 3 | 2
|  | Speck group score | [5 empty] | 3 | [8 empty] | 3 | ACR Mini | 2 | 2 | 2
|  | Mass score | [5 empty] | 3 | [8 empty] | 3 | ACR Original | 4 | 3 | 3
Overall Pass/Fail | [7 empty] | Pass
Comments:
Analyses: |  | Scoring |  | Original ACR Phantom |  |  |  | Mini Phantom |  | ACR DM Phantom
|  |  | Fibers | 4 |  |  |  | 2 |  | 2
|  |  | Specks | 3 |  |  |  | 2 |  | 3
|  |  | Masses | 3 |  |  |  | 2 |  | 2
Action Limits |  | Required: |  | Phantom image must be free of clinically significant artifacts.
The phantom image must score at least the minimum score (above), appropriate to the phantom used.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 2. Phantom IQ - Artifacts (DBT)
2. Phantom Image Quality and Artifacts (DBT) | [15 empty] | Weekly
| Facility |  | SurgiCenter | [6 empty] | Room ID |  |  | Biopsy
| SBBAP ID-Unit # |  | 12001 |  | - | 1 |  |  |  | Phantom |  |  | Gammex 2590-03
| Unit Mfr & Model |  | Hologic Affirm | [7 empty] | Phantom Type: |  | ACR Original
[7 empty] | Imaging Mode | Tomosynthesis
Year
Date (month & day) | [16 empty] | Stereotactic | DBT
Tech Initials | [16 empty] | Mo/Mo | Mo/Mo
Resulting Techniques |  | View | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Mo/Rh | Mo/Rh
|  | Paddle | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Rh/Ag | Rh/Ag
|  | AEC mode | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | Rh/Rh | Rh/Rh
|  | AEC Position | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | W/Ag | W/Al
|  | Exposure Comensation Step | [6 empty] | 0 |  |  | 0 |  | 0 | 0 |  | W/Al | W/Rh
|  | Slice or slab # (DBT only) | [5 empty] | -- | -- |  |  | -- |  | -- | -- |  | W/Rh
Target/filter
kVp
|  | mAs | [16 empty] | Object
Phantom Score |  | Artifacts P/F | [5 empty] | P | [8 empty] | Min Score | Type | Fibers | Specks | Masses
|  | Fiber score | [5 empty] | 4 | [8 empty] | 4 | ACR DM | 2 | 3 | 2
|  | Speck group score | [5 empty] | 3 | [8 empty] | 3 | ACR Mini | 2 | 2 | 2
|  | Mass score | [5 empty] | 3 | [8 empty] | 3 | ACR Original | 4 | 3 | 3
Overall Pass/Fail | [7 empty] | Pass
Comments:
Analyses: |  | Scoring |  | Original ACR Phantom |  |  |  | Mini Phantom |  | ACR DM Phantom
|  |  | Fibers | 4 |  |  |  | 2 |  | 2
|  |  | Specks | 3 |  |  |  | 2 |  | 3
|  |  | Masses | 3 |  |  |  | 2 |  | 2
Action Limits |  | Required: |  | Phantom image must be free of clinically significant artifacts.
The phantom image must score at least the minimum score (above), appropriate to the phantom used.
|  | Timeframe: |  | Required items must be corrected before clinical use.

## 3. Visual Checklist
3. Visual Checklist | [11 empty] | Monthly
| Facility | SurgiCenter | [6 empty] | Room ID |  | Biopsy |  |  | Visual Checklist from Technical Standard
| SBBAP ID-Unit# | 12001 | - | 1
| Unit Mfr & Model | Hologic Affirm | [6 empty] | Imaging Mode |  | Stereotactic
| Procedure | Inspect the unit and evaluate the functionality according to the checklist below.
Year
Month |  | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Tech Initials
All locks and detents are working properly.*
Table and/or breast support device are immobilized when patient is in compression.*
Lighting is adequate.
All moving parts move smoothly without obstructions to motion.*
All foot switches operate properly.
Biopsy unit is mechanically stable.*
Needle holder and guides are free from excessive vibrations.*
Paddles are free from cracks, sharp edges and other hazards.*
Operator and other personnel are shielded from radiation.*
Patient is visible to the operator during the procedure.*
Technique charts are posted.
Cleaning supplies and disinfectants are available and used after every patient.*
All blood and/or potentially infectious materials have been cleaned from the equipment.*
Other, as recommended by the manufacturer.
[6 empty] | P = Pass |  | F = Fail |  | NA = Not Applicable
Action Limits |  | Jan: Required:; Apr: All items must pass. Critical items are marked with an *.
|  | Timeframe: |  |  | Failures of critical items (*) must be corrected before clinical use; less critical items must be corrected within 30 days.

## 4. Comp Thk Ind.
4. Compression Thickness Indicator | [17 empty] | Monthly | 1/9/2026 RW: Note to self: Finished reformatting all calculated cells to gray background.
The worksheet accounts for the thickness input in mm or cm.
| Facility |  | SurgiCenter | [7 empty] | Room ID |  |  | Biopsy |  |  |  |  | Enter a value in cell C8 for the phantom thickness.
| SBBAP ID-Unit# |  | 12001 |  |  | - | 1 | [11 empty] | Select from drop/down list units of thickness and of pressure.
Enter a value in cell C8 for the phantom thickness.
Enter the measured values in data cells (row 15).
Description of phantom used for compression thickness test | [18 empty] | The units chosen determine the p/f limit, 5 for mm, 0.5 for cm.
Actual thickness of phantom | [8 empty] | Units of thickness |  |  |  | mm | [5 empty] | 5 | P/F limit
Target compression pressure | [8 empty] | Units of pressure |  |  |  | lb
Year
Month |  | Jan |  | Feb | Mar | Apr |  | May |  | Jun | Jul | Aug | Sep |  | Oct | Nov | Dec
Date
Tech Initials
Compression Pressure Used (if available)
Indicated thickness
Difference between indicated and actual thicknesses (Indicated - Actual)
Overall Pass/Fail
Action Limits |  | Required: |  |  | Compression thickness indicator must be accurate to within 0.5 cm (5 mm) of the actual thickness.
|  | Timeframe: |  |  | Failures must be corrected within 30 days.
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Reviewed 1/7/2025 Review of this form has caused a new comment in the Tech section

## 4. Comp Thk Ind. (Mobile)
4. Compression Thickness Indicator (Mobile Units) | [12 empty] | 1/9/2026 RW: Note to self: Finished reformatting all calculated cells to gray background.
[12 empty] | Post Relocation | The worksheet accounts for the thickness input in mm or cm.
| Facility |  | SurgiCenter |  |  |  |  | Room ID |  | Biopsy |  |  | Enter a value in cell C8 for the phantom thickness.
| SBBAP ID-Unit# |  | 12001 |  | - | 1 | [6 empty] | Select from drop/down list units of thickness and of pressure.
Enter a value in cell C8 for the phantom thickness.
Enter the measured values in data cells (row 15).
Description of compression thickness phantom |  | ACR Phantom | [10 empty] | The units chosen determine the p/f limit, 5 for mm, 0.5 for cm.
Actual thickness of phantom |  | 4.2 |  |  |  |  | Units of thickness |  | mm |  |  |  | 5 | P/F limit
Target compression pressure |  | 25 |  |  |  |  | Units of pressure |  | lb
Year
Month
Date
Tech Initials
Compression Pressure Used (if available)
Indicated thickness
Difference between indicated and actual thicknesses (Indicated - Actual)
Overall Pass/Fail
[9 empty] | P = Pass |  | F = Fail
Action Limits |  | Required: |  | Compression thickness indicator must be accurate to within 0.5 cm (5 mm) of the actual thickness.
|  | Timeframe: |  | Failures must be corrected within 30 days.
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Reviewed 1/7/2025 Review of this form has caused a new comment in the Tech section

## 5. AW Monitor
5. Acquisition Workstation Monitor QC | [15 empty] | Monthly
| Facility |  | SurgiCenter | [7 empty] | Room ID |  | Biopsy
| SBBAP ID-Unit# |  | 12001 |  |  | - | 1
| Unit Mfr & Model |  | Hologic Affirm | [8 empty] | Imaging Mode | Stereotactic
Year
Month |  | Jan |  | Feb | Mar |  | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Tech Initials
Monitor Condition P/F (clinically significant findings)
Test Pattern Image Quality (if test object is available, use for any pattern) | 5% square visible in 0% square (Patterns C-D)?
95% square visible in 100% square (Patterns C-D)?
High Contrast bar pattern(s) visualized correctly?
Low Contrast bar pattern(s) visualized correctly?
Grayscale gradient bar visualized correctly?
Mfr Automated Test P/F (if available)
Overall Pass/Fail
[12 empty] | P = Pass |  | F = Fail
Action Limits |  | Required: |  |  | Any screen blemish that could interfere with clinical information must be removed.
Test pattern image quality must pass all visual tests.
Manufacturer's automated tests, if available, must pass manufacturer's specifications (if one or more test fails, indicate F for test(s) and Overall).
|  | Timeframe: |  |  | Clinically significant monitor defects must be corrected before clinical use; all other required tests must be corrected within 30 days.

## 6. Comp. Force
6. Compression Force | [9 empty] | Semiannually
| Facility |  | SurgiCenter |  |  | Room ID |  | Biopsy
| SBBAP ID-Unit# |  | 12001 - 01 |  |  | Unit Mfr & Model |  | Hologic Affirm
Procedure |  |  | Required Equipment: Bathroom scale; towels
Place towel on detector, place bathroom scale on towel with dial or read-out positioned for reading.
Place another towel on top of scale.
Using initial power-drive mode, activate compression until it stops automatically.
Read and record the initial compression force.
Continue to compress using power-drive mode and record maximum compression force achieved.
Using manual fine-adjustment mode, activate compression until maximum compression is reached.
Read and record the manual maximum compression force.
Set the compression force to a clincally used force, approximately 10 lbs
Check that force is maintained within 1 lb for approximately 2 minutes.
Note: Ensure that the scale that is used produces accurate readings as pressure is increased.
Year
Date (month & day)
Tech Initials
|  |  | Measured Force |  | Units |  | Measured Force |  | Units
Initial power-drive compression force | [8 empty] | 0 |  | RW: Resticted selections to "lb" and "daN"
Maximum power-drive compression force |  |  |  |  | 0 |  |  |  | 0
Power-drive force is at least 10 lbs (4.4 daN) but no more than 45 lbs (20.0 daN) Pass/Fail |  |  | Fail |  |  |  | Fail |  |  |  | RW: New formula. Improve if you can.
Maximum manual fine-adjustment compression force |  |  |  |  | 0 |  |  |  | 0
Manual compression force is at least 10 lb (4.4 daN) |  |  | Fail |  |  |  | Fail |  |  |  | RW: New formula. Improve if you can.
Compression Force maintains at least 10 lbs (4.4 daN) for at least 2 minutes Pass/Fail | [10 empty] | RW: Created Drop-down for these cells, for Tech input.
Overall Pass/Fail |  |  | Fail |  |  |  | Fail |  |  |  | RW: New formula. Improve if you can.
|  |  | Enter number where appropriate. | [5 empty] | P = Pass | F = Fail
|  |  | Legend: | lb = pound
daN = decanewton
Action Limit |  | Required: | The compression force using the power mode must be at least 10 pounds, sufficient to stabilize the breast before final positioning, and must not exceed 45 pounds. The unit must be able to hold a compression of at least 10 pounds for at least 2 minutes and should not vary by more than one pound. The manually applied compression force must be at least 10 pounds but may exceed 45 pounds.
|  | Timeframe: | Failures must be corrected before clinical use.

## 7. Facility Review
7. Facility QC Review | [12 empty] | Annually
| Facility | SurgiCenter | [9 empty] | Date of QC Review
1. Review of Medical Physics Surveys and Results | [11 empty] | Reviewed
[5 empty] | Room 1 |  |  | Room 2 |  | Room 3 |  | Room 4
Room ID
Date of last Medical Physicist (MP) survey
MP QC Test Summary reviewed by radiologist?
All MP corrective actions completed?
ACR Phantom Average Glandular Dose (mGy)
Fiber Score
Speck Score
Mass Score
Significant Artifacts?
Comments:
Summary Comments from Past Year
2. Review of Technologist's QC Tests | [11 empty] | Reviewed
| Test |  | Frequency
| 1. Localization Accuracy |  | Daily
| 2. Phantom Image Quality and Artifacts |  | Weekly
|  |  |  | Room 1 |  |  | Room 2 |  | Room 3 |  | Room 4
| Scores of most recent phantom image: |  | Date
Fiber score
Speck group score
Mass score
Significant Artifacts?
| 3. Visual Checklist |  | Monthly
| 4. Compression Thickness Indicator |  | Monthly
| 5. AW Monitor QC |  | Monthly
| 6. Compression Force |  | Semiannually
| 7. Repeat Analysis (Optional) |  | Semiannually |  |  |  | % Repeats
| 9. Manufacturer calibrations |  | Per manufacturer
| 10. Radiologist Feedback (Optional) |  | Annually
3. Confirm procedures for Corrective Action when/if necessary
4. Technique Chart review for each room (see MP report for recommendations)
5. Infection Control procedures followed
6. Past and future service or service upgrades discussed (if app)
7. Past and future ACR Accreditation issues discussed (if app)
8. Review and verify completion of all Corrective Action
9. Procedures for continuity of ACR QC standards
10. Procedure for Supervising Physician oversight of QC program and corrective actions
continued on next page, 7. Facility QC Review (cont)
Note on B26: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Test numbering verified against Table o Contents

## 7. Facility Review (2)
7. Facility QC Review (cont) | [10 empty] | Annually
|  | Facility | SurgiCenter |  |  |  |  | Date of QC review
Follow-up Confirmed (If App.)
10. | Notable findings during QC review
11. | Items for quality improvement from QC review
12. | Other QC Notes
Overall Pass/Fail
Supervising Physician | [5 empty] | Supervisor (If App) | QC Technologist
signature | [5 empty] | signature | signature
Action Limit: |  |  |  | Required: | Supervising physician and QC technologist must review QC annually.
|  |  |  | Timeframe: | Not applicable.

## 8. Man. Calibrations
8. Manufacturer Calibrations | [5 empty] | Per Manufacturer's QC Manual
Facility | SurgiCenter |  |  | Room ID | Biopsy
SBBAP ID-Unit# | 12001 | - | 1
Unit Mfr & Model | Hologic Affirm |  |  | Imaging Mode
Manufacturer Procedure | Frequency:
Frequency: Per manufacturer recommendations (if applicable).
Note: Consult medical physicist and manufacturer (or service engineer) for instructions
Year
Date (month & day)
Tech Initials
Name of Calibration
[5 empty] | P = Pass | F = Fail
Action Limits | Required: | Unit should pass manufacturer's prescribed periodic tests and calibrations.
| Timeframe: | Failures should be corrected per the manufacturer's guidelines.

## 9. PhysicianFeedback - Case
9. Optional: Radiologist/Physician Feedback | [6 empty] | As Needed
Case report
Radiologist/Physician:
Date:
Biopsy Room:
Patient Identifier:
Technologists:
Procedure |  | This report should be completed by the radiologist/physician performing a procedure
The form should be used to provide feedback on the quality of care.
The radiologist/physician should complete this form as needed for each case.
A system should be in place for analyzing feedback and taking measures for improvement as necessary.
| Objective: | For the Radiologist/Physician to provide routine feedback to the technologists and manager on the quality of the images and of the procedure as a whole.
| Assessment |  | Excellent | Good | Needs improve-ment | Corrective Actions needed
Pre-Procedure set up
Patient set up
Patient / Breast positioning
Localization images/Image Quality
Targeting
Pre-fire Images
Post-Fire Images
Core Sampling
Core Imaging
Patient Recovery
Imaging/Targeting Equipment
Core Sampling Equipment
Artifacts
Incorrect ID
| Influencing Factors |  | Yes |  | No
Patient habitus
Patient mobility
Patient demeanor
Breast size
Multiple sites
|  |  | Good Case |  | Needs improvement
Overall Assessment
Comments
Radiologist/Physician Signature:
Caution: The Date Span of the patients used in this summary should be long enough to have a sufficient number of patients to achieve reasonable statistical relevance. Date spans of consecutive assessments may overlap in order to have a sufficient number of patients and good statistics.

## 9. PhysicianFeedback - Summary
9. Optional: Radiologist/Physician Feedback | [7 empty] | As Needed
Summary report
Facility:
Date of summary:
Biopsy Room:
Procedure |  | This report should be compiled at least annually to summarize the quality of the program.
A system should be in place for analyzing feedback and taking measures for improvement as necessary.
Ther results should be reviewed and discussed at Facility QC meeting.
Enter into the appropriate box, the percentage of each assessment.
| Objective: | To periodically assess the quality of the breast biopsy program.
| Date span: |  | to
Number of Patients:
| Assessment |  |  | Excellent | Good | Needs Improve- ment | Corrective Actions needed
Pre-Procedure set up
Patient set up
Patient / Breast positioning
Localization images/Image Quality
Targeting
Pre-fire Images
Post-Fire Images
Core Sampling
Core Imaging
Patient Recovery
Imaging/Targeting Equipment
Core Sampling Equipment
Artifacts
Incorrect ID
| Influencing Factors |  |  | Yes |  | No
Patient habitus
Patient mobility
Patient demeanor
Breast size
Multiple sites
|  |  |  | Good Case |  | Needs improvement
Overall Assessment
Comments
Radiologist/Physician Signature:
Caution: The Date Span of the patients used in this summary should be long enough to have a sufficient number of patients to achieve reasonable statistical relevance. Date spans of consecutive assessments may overlap in order to have a sufficient number of patients and good statistics.

## 10. Repeat Analysis - Summary
10. Optional: Repeat Analysis | [7 empty] | Semiannually
Summary Form
Date Start (Month/Day/Yr)
Date End (Month/Day/Yr)
Test Procedure | Required Equipment: All repeated biopsy exposures and means to count and sort them
Record the total number of biopsy exposures for the collection period (month or semi-annual).
Record the total number of biopsy repeat exposures for that time period.
Calculate by hand, or use formulas in spreadsheet, to calculate repeat rate.
Note: Some units may automatically calculate % Repeats. If so, enter this number into "% Repeats".
| Monthly Analysis |  |  |  | Semi-Annual Analysis
| Total # of Biopsy Exposures | # of Repeat Biopsy Exposures | % Repeats | Pass or Fail | Total # of Biopsy Exposures | # of Repeat Biopsy Exposures | % Repeats | Pass or Fail
Jan |  |  | #DIV/0! |  | 0 | 0 | #DIV/0!
Feb |  |  | #DIV/0!
Mar |  |  | #DIV/0!
Apr |  |  | #DIV/0!
May |  |  | #DIV/0!
Jun |  |  | #DIV/0!
Jul |  |  | #DIV/0! |  | 0 | 0 | #DIV/0!
Aug |  |  | #DIV/0!
Sep |  |  | #DIV/0!
Oct |  |  | #DIV/0!
Nov |  |  | #DIV/0!
Dec |  |  | #DIV/0!
| % Repeats = (# of Biopsy Repeat Exposures / Total # of Biopsy Exposures) x 100 | [5 empty] | P = Pass | F = Fail
Action Limits | Recommended: |  | The overall repeat rate should be lower than 20%. If the overall repeat rate exceeds 20% then the repeat rate in each category can be used to identify areas for improvement.
| Timeframe: |  | A corrective action plan should be implemented within 30 days and assessed at the next six month interval.

## 10. Repeat Analysis - Tally
10. Optional: Repeat Analysis
Tally Sheet
Date Start (Month/Day/Yr)
Date End (Month/Day/Yr)
Test Procedure |  | Record all biopsy repeat exposures on the form below.
Transfer the Total Number of Biopsy Repeats from below to the "Repeat Analysis -Summary Form"
for final calculation of Repeat Analysis.
Total # of biopsy exposures for time period
Reason |  | Comments/Notes | Total # of Biopsy Repeat Exposures | % Repeats
Patient-Related Repeats:
Positioning |  |  |  | #DIV/0!
Patient motion |  |  |  | #DIV/0!
Patient-caused artifacts |  |  |  | #DIV/0!
Other - Patient related |  |  |  | #DIV/0!
Technical Repeats:
Exposure too low (excessive noise) |  |  |  | #DIV/0!
Exposure too high (image saturation) |  |  |  | #DIV/0!
Equipment-caused artifacts |  |  |  | #DIV/0!
X-ray equipment failure |  |  |  | #DIV/0!
Software failure |  |  |  | #DIV/0!
Aborted AEC exposure
Inadequate Tissue Samplings
Failed Marker Deployment |  |  |  | #DIV/0!
Miscellaneous Repeats:
Blank images |  |  |  | #DIV/0!
Good images (no apparent reason) |  |  |  | #DIV/0!
Procedural-Induced Target Shift
Incorrect Targeting
Other - miscellaneous |  |  |  | #DIV/0!
Do Not Count as Repeats:
All localizations (eg, wire, I-125 seed, etc. |  |  |  | Not Included in Repeat Analysis
Quality control
Total: |  |  | 0 | #DIV/0!

## 10. Repeat Analysis - Daily
10. Optional: Repeat Analysis
Daily Counting Sheet
Date Start (Month/Day/Yr)
Date End (Month/Day/Yr)
Test Procedure |  | Record these counts daily.
For each patient, record all biopsy exposures including repeat biopsy exposures.
Transfer the Total Number of Biopsy Repeats from below to the "Repeat Analysis - Tally Sheet" for final calculation of Repeat Analysis.
Biopsy Procedure Date |  | Total Number of Biopsy Exposures | Number of Repeat Biopsy Exposures | Positioning | Patient motion | Patient-caused artifacts | Other - Patient Related | Exposure too low | Exposure too high | Equipment-caused artifacts | X-ray equipment failure | Software failure | Aborted AEC exposure | Blank images | Good images (no apparent reason) | Other - miscellaneous
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
Total |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Corrective Action
Corrective Action Log
| Facility | SurgiCenter | [7 empty] | SBBAP # | 12001 | - | 1
Room or Equipment ID |  |  | Biopsy | [5 empty] | Date
QC Test Name and # (if app):
Description:
Relevant Personnel Notified: |  |  |  | Personnel Name: | [5 empty] | Date/Time of Call/Notification:
(Radiologist, MP, tech, manager, service engineer)
Describe Actions Taken:
To Be Monitored
Confirmation of Resolution: | [5 empty] | Yes |  | NA
[5 empty] | Event resolved? | [5 empty] | Tech Signature
[5 empty] | Documentation from service engineer obtained? | [5 empty] | Date

## Unit Checklist Jan - Jun
QC Summary Checklist
Facility |  |  | SurgiCenter | [5 empty] | Room ID |  |  | Biopsy
SBBAP ID# |  | 12001 |  | - | 1 | [5 empty] | Unit Mfr & Model | Hologic Affirm
Imaging Mode |  | Stereotactic
Year
| Month | Jan |  |  |  | Feb |  | Mar |  | Apr |  |  | May | June
Phantom Image Quality (weekly) | date
initials
Localization Accuracy (Add'l Approach/Mode),if app (monthly) | date
initials
Compression Thickness Indicator (monthly) | date
initials
Visual Checklist (monthly) | date
initials
AW Monitor QC (monthly) | date
initials
Compression (semiannual) | date
initials
Mfr Detector Calibration, if app
Annual Facility QC Review
Repeat Analysis (optional - as needed)
Detector Calibration Freq: | [6 empty] | Date and initial each test: | date | [6 empty] | Cross out boxes where mfr calibration test is not required: | X
[8 empty] | initial | [7 empty] | X

## Unit Checklist Jul - Dec
QC Summary Checklist
Facility |  |  | SurgiCenter | [5 empty] | Room ID |  |  | Biopsy
SBBAP ID# |  | 12001 |  | - | 1 | [5 empty] | Unit Mfr & Model | Hologic Affirm
Imaging Mode |  | Stereotactic
Year
| Month | July |  |  |  | Aug |  | Sept |  | Oct |  |  | Nov | Dec
Phantom Image Quality (weekly) | date
initials
Localization Accuracy (Add'l Approach/Mode),if app (monthly) | date
initials
Compression Thickness Indicator (monthly) | date
initials
Visual Checklist (monthly) | date
initials
AW Monitor QC (monthly) | date
initials
Compression (semiannual) | date
initials
Mfr Detector Calibration, if app
Annual Facility QC Review
Repeat Analysis (optional - as needed)
Detector Calibration Freq: | [6 empty] | Date and initial each test: | date | [6 empty] | Cross out boxes where mfr calibration test is not required: | X
[8 empty] | initial | [7 empty] | X

## General Notes
Corrective Action Log
| Facility | SurgiCenter |  |  |  | SBBAP # | 12001 | - | 1
Room or Equipment ID |  |  |  | Biopsy | Date
Notes/Comments
Date |  |  | Comment:

## DD
PF/PFNA | Pass/Fail
P | Pass
F | Fail
NA | NA
TF | Paddle
Mo/Mo | reg
Mo/Rh | flex
Rh/Al | other
Rh/Ag
Rh/Rh
W/Ag | None
W/Al | None
W/Rh
Target | Phantom Type
Molybdenum | ACR Original
Rhodium | ACR Mini
Tungsten | ACR DM
YN | Test Pattern
Yes | SMPTE
No | TG-18 QC
TG-270 QC
NAYN
NA
Yes
No
Check

Y/N/NA
Yes
No
NA
Timeframe
Immediately
W/in 30 days
NA
Req
Required
Recommended
NA
Display Device
AW
Viewbox
DM-DBT
Stereotactic
DBT
All