## Exercise Instructions
Hands-on Exercise 1: Restructure a POCT Quality Spreadsheet
Scenario: You received a messy quality-issues export compiled from multiple fictional hospital sites. The dataset is entirely synthetic and intentionally inconsistent. Use a GPT to reorganize it into a usable quality action tracker.
Suggested task | Reformat the Raw POCT Export into a standardized quality action tracker.
Preserve | Every original record, observation, and note.
Create | Site, Device/Test, Issue Category, Original Finding, Date Identified, Assigned Owner, Priority, Proposed Next Action, Status.
Do not invent | Owners, dates, patient outcomes, regulatory citations, or missing facts.
AI-derived fields | Priority, standardized category, and proposed next action may be generated, but should be clearly treated as recommendations.
Bonus | Create a summary table showing findings by site and issue category.
Teaching note: All hospital names, sites, findings, names, ticket numbers, and scenarios in this workbook are fictional. The 'Raw POCT Export' tab is intentionally messy; inconsistent site names, device names, dates, statuses, owners, categories, abbreviations, and free-text notes are part of the exercise. Material is part of University of Pittsburgh - Computational Pathology and AI Center of Excellence (CPACE) educational content.

## Raw POCT Export
Site / Location | Device / Test | Finding Type | Issue / Observation | Date | Reported By | Owner | Current Status | Follow-up / Notes | Ticket / CAPA #
Riverview Medical Center | Accu-Chek Inform II | QC | Level 2 QC failed twice on meter G-214 | 46211 | JH |  | open | Meter still on unit when reviewed; charge RN says repeat later
Riverview Medical Center | Glucose | qc issue | Controls not documented for night shift | 7/09/26 | POCT | M. Carter | In Progress | Need to verify if testing occurred during gap | POC-1842
Riverview Medical Center ED | iSTAT | Competency | Operator locked out during trauma; competency expired | 07-11-2026 | ED educator |  | OPEN | Staff member completed annual module but hands-on record not found
Grand Oaks University Hospital | i-STAT Alinity | operator | Two users sharing operator ID noted during observation | 46217 | NT | S. Patel | Pending | Education initiated | CAPA 26-071
Grand Oaks University Hospital | Glucose meter | Documentation | Critical glucose 31 mg/dL repeated but notification field blank | 7/15/2026 | Audit |  | Open | Result appears in EHR; no call documentation visible
Grand Oaks University Hospital ED | AccuChek | critical result | HI result documented as numeric 601 in note | Jul 16 | POCT | A. Nguyen | reviewing | Need to determine source of manually entered value | INC-5521
Lakeview Regional Hospital | Hemochron Signature Elite | QC | Electronic QC passed but liquid QC sticker on kit expired | 46221 | ML |  | OPEN | No patient impact identified yet
Lakeview Regional Hospital | ACT | Supplies | ACT cuvettes found in open box without opened date | 7/18/26 | Inspection | K. Young | Closed? | Unit says box opened yesterday; no documentation
Evergreen Women's Hospital | AmniSure | storage | Test kits stored above labeled temperature range per room log | 46222 | POCT office |  | Hold | 4 kits quarantined; lot 26A17 | Q-119
Evergreen Women's Hospital | Glucose | Procedure | Old quick reference guide still posted beside meter | 07/20/26 | Round | L. Brooks | In progress | Guide references discontinued strip lot entry workflow
Summit Children's Hospital | i-STAT | QC Documentation | Simulator check completed but not recorded in local binder | 7/21/2026 | POCT |  | open | Instrument record may be retrievable from middleware
Summit Children's Hospital | iSTAT CHEM8+ | Specimen | Sample labeled after testing observed in PICU | 46225 | Observer | PICU educator | Open | Operator stated patient was verbally confirmed before draw | PS-883
Summit Children's Hospital | Glucose | Operator competency | New hire tested patient before final competency signoff | 7-22-26 | Audit |  | open | Training complete; evaluator signature missing
Harborview Medical Center | TEG 6s | QC | Normal QC cartridge not run on scheduled monthly date | 46226 | POCT | R. Evans | Pending review | Completed 2 days late after reminder
Harborview Medical Center | TEG6S | Documentation | QC review form missing second reviewer initials | 07/24/2026 | QA |  | open | All QC results acceptable | TEG-24
Harborview Medical Center ED | Accu-Chek Inform II | Supplies/Storage | Strip vial left uncapped at bedside | 7/25/26 | Rounds | C. Ford | Closed | Vial discarded; staff coached
Pinecrest Community Hospital | CoaguChek XS Plus | Patient ID | Temporary patient ID used; result did not cross to chart | 46228 | POCT |  | OPEN | Manual chart correction completed | IT-904
Pinecrest Community Hospital | INR | Connectivity | Three results queued in meter for >12 hours | 7/26 | Helpdesk | IT | resolved | Wi-Fi issue on unit; results later transmitted | HD-44380
North Valley Medical Center | Glucose | QC | QC range entered manually from old insert | 46230 | Tech specialist | D. Green | Open | New lot introduced 7/24; range mismatch suspected
North Valley Medical Center | Accu Check | Lot management | New control lot in use before lot was entered into middleware | 07/27/26 | POCT |  | in progress | Affected meters being identified | POC-1870
Meadowlands Regional Hospital | Abbott i-STAT 1 | Maintenance | Analyzer visibly soiled around cartridge port | 46231 | Inspection |  | Open | Unit cleaned device during visit
Meadowlands Regional Hospital | iSTAT | Procedure | Printed procedure at bench was version 5; current version is 7 | 7/28/26 | NT | C. Diaz | Closed | Old copy removed
Cedar Ridge Hospital | Glucose | Training | Agency nurse unable to locate critical result procedure | 46232 | Rounds |  | OPEN | Nurse had active operator access
Cedar Ridge Hospital | Accu-chek | QC | QC not performed after new strip lot loaded | 07-30-2026 | POCT | TBD | Open | Patient tests performed after lot change | QMS-26-144
Stonebridge Community Hospital | hCG urine | Kit issue | Cardinal hCG kit box opened; no open date written | 7/30/26 | Inspection |  | open | Expiration date valid
Stonebridge Community Hospital | Urine pregnancy | Procedure | Staff still referring to test as Alere hCG after Cardinal-labeled kits introduced | 46234 | POCT | E. Sims | Monitor | No testing error noted; update education materials
Clearwater Medical Center | Glucose | Documentation | Meter comment 'MD notified' used but no provider name documented | 08/01/26 | Audit |  | OPEN | Need to confirm local documentation expectation
Clearwater Medical Center | AccuChek Inform | Competency | Competency file contains quiz only; direct observation form absent | 46235 | QA | Unit Educator | In Progress | 12 operators sampled; 3 files incomplete | CAPA-26-081
Willow Springs Behavioral Health | Glucose | Patient ID | Barcode unreadable; result entered under downtime ID | 8/2/26 | Staff report |  | closed | Result reconciled later same shift
Willow Springs Behavioral Health | Accu-Chek | supplies | Control bottle found without opened date | 46236 | Rounds | Nurse manager | open | Bottle discarded
Blue Ridge Community Hospital | i-STAT | Temperature | Cartridge refrigerator min temp recorded 1.2 C overnight | 08/03/2026 | POCT |  | Hold | Cartridges quarantined pending review | TEMP-771
Blue Ridge Community Hospital | CG8+ | Specimen | Air bubble noted in syringe before cartridge fill | 46237 | Observation |  | Open | Operator repeated test using new sample
Mountainview Regional Medical Center | Glucose | QC | QC failure followed by patient test before successful repeat QC | 8/4/26 | Audit | P. Hall | OPEN | One patient result involved; chart review requested | PSR-1049
Mountainview Regional Medical Center | Accu-Chek Inform II | Connectivity | Meter dock not communicating with middleware | 46238 | Help Desk | Biomed? | In progress | Approx 19 results held on meter | HD-44903
Riverbend Community Hospital | Hemoccult | Competency | Operator roster includes 6 inactive employees | 08-05-26 | POCT |  | open | Access removal requested
Riverbend Community Hospital | Gastric occult blood | Procedure | Procedure printout has handwritten changes | 46239 | Inspection | B. White | Pending | Need controlled copy replacement
Riverview Medical Center | Glucose | Storage | Test strips stored on windowsill in direct sunlight | 8/6/2026 | Rounds |  | OPEN | Two vials removed from use | Q-128
Grand Oaks University Hospital ED | iSTAT ACT-K | QC | QC result transcribed to paper log with wrong lot number | 46240 | Audit | S. Patel | open | Electronic instrument record has correct lot
Lakeview Regional Hospital ICU | Glucose Meter | Critical value | Critical low repeated on same finger without documented assessment | 08/07/26 | Chart review |  | Review | Need to compare with approved repeat-testing procedure
Evergreen Women's Hospital L&D | Amnisure | Operator | Operator tested after access should have expired | 46241 | Middleware audit | L. Brooks | OPEN | Possible interface delay in lockout file | IT-922
Summit Children's Hospital ED | i-STAT | Specimen type | Capillary sample used for cartridge normally run on venous specimens | 8/8/26 | POCT |  | Pending review | No patient identifiers included in audit export
Harborview Medical Center OR | TEG6s | Training | Travel nurse unfamiliar with cartridge equilibration instructions | 46242 | Rounds | OR educator | closed | Just-in-time training completed
North Valley Medical Center ICU | AccuChek | QC | Control result entered under wrong level | 08/08/2026 | Audit |  | open | Result acceptable for other level; documentation correction needed
Meadowlands Regional Hospital ED | i-STAT Alinity |  | Cartridge lot expired in analyzer inventory but physical box not expired | 46243 | POCT | C. Diaz | Investigating | Possible lot-entry date issue
Stonebridge Community Hospital ED | hCG combo | documentation | Positive urine hCG result repeated but repeat reason not documented | 8/9/26 | Audit |  | OPEN

## Suggested Prompt
Suggested Prompt for the Exercise
Reformat the attached POCT quality spreadsheet into a standardized quality action tracker. Preserve every original record and retain the original observation and notes. Create columns for Site, Device/Test, Issue Category, Original Finding, Date Identified, Assigned Owner, Priority, Proposed Next Action, and Status. Standardize inconsistent site/device/category names. Do not invent owners, dates, patient outcomes, regulatory citations, or other missing facts; use 'Not provided' where appropriate. You may recommend a priority and next action, but clearly identify those as AI-generated recommendations. Then create a summary table showing the number of findings by site and issue category.