## Source_Data
Source Data - Preserve Unchanged
Intentionally messy fictional event log. Values are reproduced as entered; spelling, date formats, blanks, contradictions, and duplicate identifiers are part of the exercise.
Source_Record_ID | Date_Reported | Facility_as_Entered | Unit_as_Entered | Test_or_Device_as_Entered | Source_Type | Finding_Narrative_as_Entered | Immediate_Action_as_Entered | Owner_as_Entered | Due_Date_as_Entered | Status_as_Entered | Severity_as_Entered | Patient_Outcome_as_Entered | Source_Reference | Reporter_Role_as_Entered
POC-2401 | 01/05/26 | Northstar Main | ED | GlucoSure GX | Monthly audit | Three operators used the generic TRAINING login for patient testing during night shift. Charge RN says badge reader was down. | Generic login disabled after shift. | ED educator | 1/12/26 | closed | High | No harm noted | Audit tab 3, rows 18-20 | POCT specialist
POC-2402 | 2026-01-08 | Northstar-Main | Emergency Dept | GlucoSure | Staff email | Meter showed patient result but it did not cross to EHR. Nurse manually entered 46 mg/dL; interface later sent the result too. | Duplicate removed per ticket. | IT / POC | ASAP | Resolved | moderate | none | Help desk INC-8821 | RN
POC-2403 | 1/9/2026 | Riverside | ICU | GlucoSure GX meter | QC review | Level 2 QC failed twice. Patient testing continued for about 40 minutes because another meter was not nearby. Three patient results were produced. | Meter pulled from use after supervisor arrived. |  |  | Open | HIGH | No outcome documented | QC exception report | Lab tech
POC-2404 | Jan 10 2026 | Riverside Hospital | Intensive Care | GlucoSure GX | Manager note | Likely same QC event as 2403; manager says only two patients were tested, not three. | under review | ICU mgr | next Friday | open? | High |  | Email chain “meter QC” | Nurse manager
POC-2405 | 2026/01/14 | East Pavilion | Med/Surg 4 | CoaguCheck One | Observation | Test strips found open on cart; vial cap was loose. Open date written as 1/4 but staff member said it was opened yesterday. | Vial discarded. Staff reminded. | Unit education | 2/1 | In progress | low | N/A | Rounding sheet 1/14 | POCT coordinator
POC-2406 | 01-17-2026 | Children’s Center | PICU | Blood gas - ABL Mini | Temperature log | Reagent refrigerator log blank for 3 days. Min/max display was 9 C when checked. It is unclear how long temperature was high. | Reagents moved to lab refrigerator; vendor called. | J. Lee |  | Pending | Med | unknown | Temp log B | RT supervisor
POC-2407 | 1/18/26 | Childrens Ctr | PICU | ABL mini | Follow-up email | Email states affected cartridges were “probably okay” and patient correlation looked fine, but no patient list or comparison data attached. | Testing resumed. | Resp Therapy | 1/18/26 | done | Medium | no issues seen | Email 1/18 | RT lead
POC-2408 | 2026-01-22 | Northstar Main Campus | OR | ACT Pro | Competency report | Four active operators show competency expired 12/31/25. Two performed testing in January. | Access removed for four operators. | OR educator | 2026-01-25 | Closed | major | No known impact | Competency dashboard export | POCT analyst
POC-2409 | 01/24/2026 | West Clinic | Urgent Care | RespiraNow Flu A/B Ag | Complaint | Staff read cassettes whenever convenient; observed one cassette read around 22 minutes. Result documented negative. | Provider notified and molecular specimen collected. | Clinic supervisor | 1/24/26 | Complete | High | PCR later positive | Complaint form 26-04 | MA
POC-2410 | 2026-01-25 | West Ambulatory | UC | Flu A B rapid | Huddle note | Team says timers are often missing and phones are used. Nobody could confirm which patient cassette was read late last week. | New timer requested. |  |  | ongoing | ? |  | Safety huddle notes | Clinic RN
POC-2411 | 01/28/26 | Riverside | Labor & Delivery | GlucoSure GX | Critical result audit | Critical glucose 31 documented in EHR. Notification field says “MD aware” but has no name, time, or read-back. | Record referred to unit manager. | L&D Manager | 2/10/26 | open | High | Infant transferred to NICU - relationship unknown | Critical report audit line 6 | Quality analyst
POC-2412 | 2026-01-31 | Northstar Main | NICU | GlucoSure GX | Event report | Heel specimen was collected soon after alcohol cleaning; operator recalls site may still have been wet. Meter result 28, repeat central lab 52. | Treatment already started; POCT notified. | NICU educator |  | Reviewing | critical | Dextrose given | Safety report SR-551 | RN
POC-2413 | 02/02/26 | North Star Main | ED | i-Stat CG8 | Supply check | Cartridges stored in unit refrigerator door. Several pouches had no room-temperature removal date. | Undated pouches discarded. | ED |  | Closed | minor | not applicable | Inventory sheet | Lab assistant
POC-2414 | 2026-02-03 | Riverside | ED | iSTAT CG8+ | Device log | Analyzer lockout date had passed but device still allowed testing after date/time reset during battery change. | Clock corrected; analyzer docked. | Biomed/POCT | TBD | Open | High | none reported | Device log 02-03 | POCT specialist
POC-2415 | 2/4/26 | East Pav. | M/S 4 | Coag check | QC log | No external QC documented for current strip lot. Staff state QC was done but entered under old lot number. | Current lot QC completed and passed. | unit educator | 2/4 | closed | Moderate | N/A | QC middleware comment | POCT coordinator
POC-2416 | 2026-02-06 | East Pavilion | Med Surg 4 | CoaguCheck One | Audit | Possible duplicate of POC-2415. Report lists missing QC for lot C118; middleware shows QC under C181. | Lot numbers being verified. | POCT |  | In progress | medium |  | Audit worksheet row 33 | Auditor
POC-2417 | 02/08/26 | Northstar Main | ED | NovaFlu AMP | Instrument alert | Analyzer displayed repeated E17 temperature errors after being placed beside warming cabinet. Four cartridges aborted. | Analyzer moved; tests completed on second device. | ED charge | today | resolved | Low | Testing delayed ~35 min | Analyzer event log | RN
POC-2418 | 2/8/2026 | NS Main | ER | Nova Flu molecular | Help desk ticket | Ticket describes six failed cartridges, not four. No patient accession list included. | Vendor case opened. | Vendor? |  | Closed? | moderate | delay only | INC-8950 | Help desk
POC-2419 | 2026-02-11 | Riverside Hospital | ICU | GlucoSure GX | Cleaning observation | Meter returned to dock with visible blood near strip port. Staff member wiped it quickly but did not keep surface wet for listed contact time. | Meter recleaned by observer. | ICU education | 2/28/26 | Open | medium | N/A | Rounding note 2/11 | Infection prevention
POC-2420 | 02/12/2026 | Riverside | ICU | GlucoSure | Staff statement | Several approved wipes are “too wet,” so staff sometimes squeeze liquid out before use. No specific meter or date identified. | None documented. |  |  | new | Low |  | Interview notes | Educator
POC-2421 | 2026-02-14 | Northstar Main | OR 12 | ACT Pro | Patient ID audit | Two results were associated with temporary ID TRAUMA1. One was later matched to MRN; the other patient is uncertain in the audit file. | Health information management contacted. | POCT and HIM |  | OPEN | critical | unknown | Interface audit 2/14 | POCT analyst
POC-2422 | 2/15/26 | West Clinic | Urgent Care | RespiraNow Ag | QC sheet | Monthly positive control result entered as “pass”; no line pattern, kit lot, expiration, or operator initials recorded. | Control repeated and acceptable. | Clinic lead | March | in progress | medium | N/A | Paper QC log page 2 | POCT coordinator
POC-2423 | 02/18/26 | Childrens Center | PICU | ABL Mini | Maintenance report | Weekly inlet cleaning overdue by 9 days. RT says device prompt never appeared. | Cleaning completed. | RT | 2/18/26 | done | Med | No known impact | Maintenance export | RT
POC-2424 | 2026-02-20 | Northstar Main | ED | GlucoSure GX | Proficiency test review | PT sample tested twice before result submission; first run was not submitted. Staff thought repeating was allowed because first result “looked wrong.” | Education provided to operator. | POCT education | 3/1/26 | Closed | High | N/A | PT attestation | POCT director
POC-2425 | 02/21/2026 | East Pavilion | Pre-op | hCG cassette | Direct observation | Urine sample and cassette were not labeled at bedside. Operator carried both to desk shared by two rooms. | Testing stopped; sample discarded. | Pre-op manager | Immediately | complete | HIGH | Procedure delayed | Observation 2/21 | Lab inspector
POC-2426 | 2026-02-23 | East Pavilion | Preop | QuickCheck hCG | Email | Manager disputes observation: says cassette had room number on wrapper, though not two patient identifiers. | No additional action listed. | Manager |  | closed | Medium | No harm | Email reply 2/23 | Nurse manager
POC-2427 | 2/25/26 | Riverside | L&D | GlucoSure GX | Result audit | Meter result 39 was repeated 18 minutes later rather than immediately. Notification documentation is blank. | Case sent for review. | L&D |  | Open | high |  | Critical audit item 12 | Quality specialist
POC-2428 |  | Northstar Main | ED | i-STAT | Anonymous report | Report says staff sometimes use venous blood from unlabeled syringes during codes. No date, patient, analyzer, cartridge, or operator identified. | Manager discussed at huddle. | ED leadership |  | monitoring | severe | unknown | Anonymous hotline item | unknown
POC-2429 | 2026-03-01 | West Clinic | UC | RespiraNow Flu | Inventory review | Two test kits expired 2/28 remained in active drawer on 3/1. No evidence they were used. | Expired kits removed. | Clinic manager | 3/1/26 | Closed | low | N/A | Inventory photo | POCT tech
POC-2430 | 03/02/26 | West Clinic | Urgent Care | RespiraNow Duo Ag | Result review | One positive result has no collection time or read time. Operator remembers it was read “about ten minutes.” | No action documented. |  |  | Open | medium | patient treated | Chart review 555 | Auditor
POC-2431 | 2026-03-04 | Northstar Main | NICU | GlucoSure GX | Training file review | Annual competency form signed by educator but direct-observation section is blank for 6 operators. All remain active. | Files sent back to educator. | NICU educator | end of month | In progress | High | N/A | Competency binder | POCT educator
POC-2432 | 3/5/26 | Riverside | ED | NovaFlu AMP A/B | Downtime log | Three molecular results were manually entered during interface downtime. Device IDs were omitted; after interface recovery, two duplicate results appeared. | Duplicates marked in EHR. | IT | 3/6 | resolved | Moderate | No reported harm | Downtime log 3/5 | ED RN
POC-2433 | 2026-03-05 | Riverside ER | Emergency | NovaFlu AMP | Interface report | Lists four manually entered results and three duplicates for same downtime window. One accession overlaps POC-2432; remaining accessions not listed. | Reconciliation requested. | POCT/IT |  | open | med | none documented | Interface queue report | Interface analyst
POC-2434 | 03/08/2026 | Children’s Center | PICU | ABL Mini | QC review | Both QC levels passed, but operator entered comment “repeat after fail.” Original failed run cannot be found in device history. | POCT asked vendor to retrieve logs. | POCT |  | Pending | medium | N/A | QC dashboard | POCT specialist
POC-2435 | 3/09/26 | Northstar Main | OR | ACT Pro | Supply room inspection | Cartridge box stored below sink; outer carton damp. Individual pouches appeared intact. | Entire box quarantined. | OR supply | 3/10/26 | Quarantined | Medium | N/A | Inspection photo set | POCT coordinator
POC-2436 | 2026-03-11 | East Pavilion | Med/Surg 4 | CoaguCheck One | Result audit | INR 6.2 documented. No provider notification found. A note 45 minutes later says “team aware.” | Manager asked to verify notification. | Nurse manager |  | Open | High | No adverse event in note | High result report | Quality analyst
POC-2437 | 03/13/26 | East Pav | M/S 4 | CoaguCheck | Interview | Operator states high INR values are usually repeated, but could not locate procedure. This may be a training gap. | None. |  |  | New | low |  | Interview 3/13 | POCT coordinator
POC-2438 | 2026-03-15 | West Clinic | Urgent Care | QuickCheck hCG | Incident note | Control window absent. Operator entered “negative” before noticing invalid cassette. Patient left before repeat. | Provider notified; patient called. | Clinic RN | same day | Closed | High | Follow-up unknown | Incident WC-18 | MA
POC-2439 | 3/16/26 | Northstar Main | ED | GlucoSure GX | Middleware audit | Operator ID field contains badge number assigned to employee who left organization in December. Twelve results linked to ID. | ID deactivated. Access review opened. | POCT security | 3/20 | In progress | High | not assessed | Operator audit | Systems analyst
POC-2440 | 2026-03-18 | Northstar Main | ED | GlucoSure GX | Manager email | Former employee badge was kept in locked drawer and used only for QC, according to manager. Middleware shows 12 patient records, but names are redacted in export. | Awaiting full report. | ED manager / Security |  | Open | high | Unknown | Email + redacted export | POCT manager
POC-2441 | 03/20/26 | Riverside | ICU | iSTAT CG8+ | Specimen review | Arterial syringe sat on counter for approximately 25 minutes before testing. Collection time was entered after result was produced. | Result compared with blood gas lab; described as close. | ICU educator |  | closed | medium | No documented harm | Review worksheet | RT
POC-2442 | 2026-03-22 | Children’s Center | NICU | GlucoSure GX | Lot review | New strip lot placed on two units. QC documented on one meter only. Staff believed one meter was enough for the lot. | QC run on remaining meters. | POCT | 3/22/26 | Resolved | Medium | N/A | Lot deployment log | Lab tech
POC-2443 | 3/24/26 | West Clinic | UC | RespiraNow Ag | Temperature log | Kit storage room reached 32 C. Log has one reading at 16:00; duration above limit unknown. Staff used 7 tests that day. | Remaining kits held. Manufacturer guidance requested. | Clinic manager / POCT | TBD | Hold | High | Outcomes not reviewed | Temp log + inventory | POCT coordinator
POC-2444 | 2026-03-25 | West Ambulatory | Urgent Care | RespiraNow Duo | Vendor call note | Vendor allegedly said kits were acceptable up to 35 C for “brief exposure.” Written confirmation not attached. | Kits returned to use. | POCT director? |  | closed? | Medium |  | Call note 3/25 | Clinic manager
POC-2445 | 03/27/2026 | East Pavilion | Pre-op | QuickCheck hCG | Competency report | Eleven operators completed online module; demonstration and direct observation fields are marked N/A. Local educator says online course alone is sufficient. | No change yet. | POCT Education |  | Open | major | N/A | LMS export + forms | POCT educator
POC-2446 | 3/29/26 | Riverside | ED | NovaFlu AMP | Maintenance audit | 30-day cartridge-bay cleaning documented on paper but not in analyzer. Date is 3/31, two days in the future. | Educator asked operator to clarify. | ED educator | 4/2 | Open | low | N/A | Maintenance sheet | POCT tech
POC-2447 | 2026-04-01 | Northstar Main | OR | ACT Pro | QC trend review | QC values show gradual upward shift over six days but remain within displayed limits. Reviewer wrote “calibration drifting” without service data. | Technical review requested. | POCT technical |  | Monitoring | Moderate | No impact identified | QC chart note | MLS
POC-2448 | 04/02/26 | Children’s Center | PICU | ABL Mini | Supply audit | Two reagent packs have handwritten expiration one month beyond printed date. Staff said vendor extended dating, but extension letter was not available. | Packs quarantined. | RT supervisor | 4/3/26 | Pending | High | unknown if used | Inventory audit | POCT coordinator
POC-2449 | 2026-04-04 | Northstar Main | ED | i-STAT CG8+ | Patient complaint | Patient reports blood was collected twice because first cartridge failed. Complaint alleges delay caused worsening symptoms; chart review not complete. | Complaint forwarded to patient relations. | Patient relations |  | Under Review | severe | alleged harm - unverified | Complaint PR-804 | Patient relations
POC-2450 | 4/05/26 | Northstar-Main | Emergency Dept | iSTAT | Device history | Device log shows cartridge error at 10:42 and successful test at 10:49. No clinical notes reviewed. May relate to complaint POC-2449. | No action documented. |  |  | new | Low |  | Analyzer 1A log | Biomed
POC-2450 | 04/05/2026 | Northstar Main | ED | i-STAT CG8 | Event export | Duplicate ID POC-2450 generated during spreadsheet merge. Narrative says one cartridge error followed by successful repeat. | None. |  |  | Open | minor | not stated | Merged event export row 91 | Data analyst
POC-2452 | 2026-04-08 | Riverside | L and D | GlucoSure GX | Critical result review | Result 27 appears in middleware but not EHR. Downtime form found with value 72 and same timestamp; unclear whether transposition or different patient. | Records requested. | L&D + IT |  | open | Critical | Unknown | Middleware query + paper form | Quality manager
POC-2453 | 4/10/26 | East Pavilion | Med/Surg 4 | CoaguCheck One | Corrective action follow-up | Education from POC-2415 reportedly completed for all staff. Attendance sheet lists 18 names; unit roster has 24. No explanation for six missing names. | Manager asked to reconcile roster. | Unit manager | 4/20/26 | In progress | medium | N/A | Attendance sheet | POCT educator
POC-2454 | 2026-04-12 | West Clinic | Urgent Care | RespiraNow Flu A/B | Observation | Operator used correct timer and read at 10 minutes. Tester did not record kit lot because field was not available in EHR. | Paper lot log started temporarily. | POCT IT | unknown | Open | Low | No harm | Observation 4/12 | POCT specialist
POC-2455 | 04/14/26 | Childrens Center | PICU | ABL Mini | Training note | New hire observed one test and was then given active access. Quiz and simulated error assessment are not in file. Educator says documents may be in another binder. | Access suspended pending records. | PICU educator | Friday | Pending | High | N/A | Training file review | POCT educator
POC-2456 | 2026-04-15 | Northstar Main | Hospital-wide | All POCT | Committee minutes | Committee agreed “documentation remains our biggest problem,” but minutes do not define which devices, units, or measures support the statement. | Education campaign proposed. | POCT Committee | Q2 | Planned | medium | N/A | POCT Committee minutes | Recorder